## Slide 1: Raster Scanning Magnet System CDR-3
Validation of Pre-Series
Heine D. Thomsen, Søren Pape Møller
Aarhus University (AU), Denmark
www.europeanspallationsource.se
January 13, 2017

## Slide 2
2016
2017
2018
2019
2020
2021
2015
Outdated schedule; for indication only!
First Protons on Target, 570 Me V (12 Aug. 2019)
Phase 1
Phase 2
@Danfysik
Phase 3
@Aarhus
Phase 4
@Danfysik
Phase 5
@ESS
Design
Pre-Series
Prod. + Assembly + Inst.
Pre-Series Test
Pre-Series Extended Test
Delivery to Aarhus
FAT I
CDR 3
SAT I
Prod.-Series
Production + Assembly
FAT II – SAR 1
Delivery to ESS
SAT II – SAR 2
TRR
Testing & Commissioning
RFI (10 Apr. 2019)
Preliminary Design
CDR 1
CDR 2
Notes: Flanges are identical size
External power supply expert at CDR-3
Thermal breakers on RSMs
6. Ceramic chamber spares
7+8: FDU and EVR
PDR

## Slide 3: RSM Past, Present & Future
2015.10.08, CDR-1: to start procurement
2016.06.01, Contract signed with vendor (Danfysik, Taastrup, DK)
2017.01.13, CDR-2: Validation of DDR & FAT-I
2017.11: FAT-I
2017.12-2018.04: SAT-I
2018.05.15, CDR-3: approving the performance of the pre-series
2018.06.11: Purchase period II. Vendor initiates procurement of remaining materials and parts
  Spares?
2018.11: delivery of RSMS to ESS (following FAT-2)
Notes: Danfysik: Gantt-chart

## Slide 4: ESS A2T DC Beam Optics: Similar to LANL MTS (B. Blind, LINAC’06, MOP055)
ACH
AP
A2T
QP1-4
QP5,6
2.0 GeV: ΣBxL = ±10 mT.m, ΣByL = ±6 mT.m
PBW
Target
Centroid
10 x RMS beamlet
Raster
CO
Apertures
p’s
n’s
Neutron Shield
(Ø40 x 2000) mm
R12 = R34 = 0
Notes: Apertures

## Slide 5: Lissajous Pattern
5/4 pattern
113/83 pattern (200 kHz filter)
Beamlet RMS
Hor / Ver wfm
(ax, ay, σx, σy) = (60, 20, 12, 4.5) mm
Closed pattern (fx/fy, ϕxy, ax, ay) within a beam pulse:
T0 = 2.86 ms = (350 Hz)-1 fx = nx / T0, fy = ny / T0
Triangle waveforms => No lingering near edges, crosshatch
fx / fy ~ 1 => A single (magnet + supply) design
fx = 39.6 kHz
fy = 29.1 kHz
Notes: produce and monitor frequencies
The simplicity makes the system easy to comprehend for any operator.
Suppresses beamlet profiles
pure Gaussian vs end-to-end beam

## Slide 6: Functional Description
Provides
Triangular beam deflections in H and V at distinct frequencies
Status & monitoring signals
Vacuum along beam axis
Mechanical support for its components
Requires
Configuration (frequency, amplitude)
Trigger, polarity
Electrical power
Notes: Maintenance

## Slide 7: System Requirements
Shall raster the 0.2-2.0 GeV beam in both horizontal (H) and vertical (V) direction across the target
  maintain distinct H&V amplitude setpoints
    5 mT.m / magnet (nominally ByL =1.5 mT.m and BxL =2.6 mT.m)
  maintain distinct H&V frequency setpoints
    configurable up to 40 kHz (10-40 kHz)
Shall be synchronized to (and fully cover) the proton beam pulse (14 Hz x 3.57 ms = 5% duty cycle)
Shall be robust towards localized component and configuration failures:
  4-fold redundancy in both planes, individually powered by dedicated (identical supplies)
…
Notes: 5 mT.m = 100%
Danfysik spreadsheet

## Slide 8: An Early Block Diagram
Notes: Bdot
Interlocks

## Slide 9: Hardware: Magnets + Support
Profile monitor (wires + non-invasive)
AU design sketch
Notes: Cooling water
Bellows, pumps, gauges
Girder

## Slide 10: RSM Supply (feasibility study level)
(MPS: Magnet Power Supply)
35 m
Regulation
Controller
Input Converter
Output Converter

## Slide 11: BDS Optics Design
(Same state as CDR-2)

## Slide 12: A2T Beam Optics
ACH
By By Bx Bx | Bx Bx By By
AP
A2T
QP1-4
QP5,6
2.0 GeV: BxL = ±2.6 mT.m, ByL = ±1.5 mT.m, ±5 mT.m
PBW
BEW
Centroid
10 x RMS beamlet
8 x RSMs
CO
Notes: RSM physical apertures

## Slide 13: Multiparticle Studies
IPAC’14, WEPRO074
1000 HEBTs x 106 particles
RSM aperture (s = 207 m) appears adequate
NSW aperture (s = 220 m) appears adequate

## Slide 14: A2T Quadrupole Magnet Harmonics?
A2T quad harmonics were scanned in log. steps, 1 u to 100 u (un = ∫∆Bn/B0 x 104 = 1 u)
500 k particles tracked to target
7.7 unit contamination for each un

## Slide 15: RSM Systematic Sextupole?
u3 = ∫∆B3/B0(t) x 104
Beam along RSMs:
  rms size < 1 mm
  dx, dy < 3 mm
RSMs
  GFR = ± 15 mm
  Gap = ± 50 mm
  Aperture = 78.0±2.5 mm
RSM u3 < 103 (or 10%)
Systematic sextupole (u3) in H RSM set

## Slide 16: RSM Alignment Errors?
Insensitive to displacements:
  dx, dy < 0.5 mm
  dz < 0.3 mm
Roll errors:
  distorts pattern outline (shearing)
  depends on average roll error in H/V set
  pitch, yaw, roll < 1 mrad
Roll error distribution < 100 mrad

## Slide 17: Magnets

## Slide 18: Raster Scanning Magnet (RSM)
Field quality: Δ(∫Bdl)/∫Bdl < 10% @ GFR = ± 15 mm
Window frame magnet (CMD5005 NiZn ferrite)
Notes: Out of vacuum magnet
40 kHz, Ceramic chamber
  Alignment & field quality
Alignment positions

## Slide 19: RSM Connections
Bus bars at top and bottom of coil ends
Bdot loop: 1 / RSM
  a single loop around return yoke
  ± 51 V at full amp. Divided to ± 10 V

## Slide 20: RSM Pair
Coil busbars
Thermal circuit breakers
Coil ends
RSM housing
Ceramic chamber
alignment
Notes: mechanical properties, G11 Etronax pieces are added to the vulnerable areas and the spaces between each loop will be filled with G11 Etronax pieces.
Not prototype: machining, paint job, equipment labels, etc.

## Slide 21: Coil Impregnation Design?
Cu
G11
epoxy
epoxy + alumina
v2
v1
v3
~60% glass / alumina fill factor in epoxy matrix
Radiation resistance of glass and alumina assumed to be similar
v3: a simpler coil, i.e. expected to be cheaper
Notes: - coils will be vacuum resin impregnated with radiation resistant CTD101K (25 MGy) inside a mould
CTD403 / CTD422

## Slide 22: Magnetic Field Studies – 2D & 3D
Using the transient module in Opera 2D, eddy current and proximity effects have been evaluated
  current distributions are affected but the design is not compromised at 40 kHz (skin depth ~0.3 mm in Cu)
Only minor surprises found:
  The inductance is found to be larger than the analytical calculation.
Integrated field homogeneity @ GFR = 15 mm (spec. <10%)
  DDR simulations: 1%
  FAT: <= 0.26% (calibrated strip line probe)
Thermal calculation: ~15 W (300 W x 5% duty cycle) is dissipated in each magnet

## Slide 23: Field Measured with 1 cm2 Pickup Coil
Notes: Comment on FAT2

## Slide 24: Vacuum Chambers & Support Structures

## Slide 25: Vacuum System
Vacuum system consists of
Ceramic (Al2O2) chambers at the RSMs
Quick-Conflat flanges
Hydro-formed bellows, ID Ø120 mm
No pumps, gauges, etc. are relevant
No metallization layer is provided (but prepared)
One chamber design, QCF100+QCF100
Notes: - - “design and manufacturing shall follow … ESS vacuum standard handbook” and UHV standards in general. Oil free…
- brazed transition pieces made from NiFe alloy
- The ceramic pieces will be prepared for a possible metallization of its interior by having the metallization for brazing extended on the inner surface by 4-5mm.

## Slide 26: Ceramic Chambers
Notes: Metallization plans + thickness

## Slide 27: Ceramic Chamber Coating?
Chamber inner surface is to be metallized at ESS coating facility
Thickness is to be
  Small; enough to be transparent to raster freqs.
  Thick; shielding bunch freqs.
  Uniform? Verified at ESS?
  Could be ~< 1 µm. Capability of coating facility?
Need to align schedules…
Verification? Measure before and after coating and compare.
For a given current amp, note:
  Supply capacitor voltage
  Signal from probe coil + voltage integrator
Probe coil could be
  As wide as chamber ID
  Narrow (to allow scans)
Notes: Chambers coated? Pre-series offer testing.

## Slide 28: RSM Support (1 of 2)
Notes: fixed but removable metallic covers for HV connections
alignment procedure
(dx, dy) < 0.5 mm, dz < 0.3 mm
(ϕx, ϕy, ϕz,) < 1 mrad

## Slide 29: Support
Girder-based: most elements are fixed to precision-machined surfaces on a girder that is aligned as a rigid body relative to two support stand columns.
  magnet-to-magnet alignment will be checked within a girder section during assembly
3 screws + 3 turnbuckles provide six degree of freedom for alignment of the girder
A total of ±25 mm horizontal range and ±35 vertical range
Alignment – iterate:
  reach nominal beam height of 1500 mm
  angular alignment to ±0.1° (pitch/yaw/roll)
  transverse adjustments
RSM housing top plate features
  a set of alignment fiducial holes
  room for a precision level gauge

## Slide 30: AU Test Lab
Termination box
RSM2 RSM1
RSM2 RSM1
PLC
Power cables
Girder
Local PC
PicoScope
EVR
Bdot, Interlock

## Slide 31: RSM Power Supply (RSMPS)

## Slide 32: RSMS-PS Block Diagram: 4 crates/magnet
Notes: Remote line
ISOx

## Slide 33: RSMPS Block Diagram

## Slide 34
Iñigo Alonso
CC
IPC
OC
CCPS
IOC+EVR
Notes: Enable+Interlock

## Slide 35: Control Crate, 19” 3U + IPC

## Slide 36: Control Crate, 19” 3U + IPC
Bechoff CP6706 IPC
Communication to local and remote control
Interlock supervision
ON/OFF control
Analogue measurements
Motherboard for regulation and DCCT module
Signal interface module (including light guides)
  Timing EVR
Remote line, RJ-45
  SCPI protocol
DCCT interface
Auxiliary power supplies for all electronics

## Slide 37: Output Converter, 19” 4U
Cap. bank (and bleeder circuit)
2 x half-bridge modules: SiC MOSFET, 1.2 kV, 3.6 mΩ (CAS325M12HM2)
Gate drivers
Output current transducer (for regulation)
Heat sink with NC thermal switches (supply interlock)
Short to ground on output: soft earth connection
Notes: Integrator

## Slide 38: Output Converter, 19” 4U
DCCT
From
CCPS
Load
cables

## Slide 39: Regulation
Intra-Burst (“fast”)
Pos. and neg. current peak levels are compared for every “pulse”, e.g. every 1/40 kHz
A waveform DC offset can be neutralized in time domain: by introducing a timing skew in the H-bridge
Inter-Burst (“slow”)
Pos. and neg. peak levels are averaged over a burst
Levels are compared to amplitude setpoints
  Predicted DC Link value is reconfigured in the post-processing phase.
SPICE sim. of current
@Magnet
@Supply output
350A / 340A = 1.03
Notes: Current output peaks are detected and sampled (ADC)
Simulations show that offset regulation can be very efficient
Cable distortion and filter effects + slight overshoot (350/340 = 1.03)

## Slide 40: Between CDR-2 to assembled hardware…
Ext. master oscillator discontinued
Studies:
  Cable parameters
  Early testing of SiC MOSFETs
    No snubbers: only 60V overshoot at max. 600V (VDS = 1200V)
  Regulation module, proof of principle
System 8700 (Beckhoff IPC)
Coil impregnation testing
LLI delays:
  Ceramic chamber
  Ferrites
Early current waveforms: 30 A @ 40 kHz
Notes: Snubbers were found unnecessary:

## Slide 41: Termination Box
RSM1
RSM2
Matching filter
RSM1
Matching filter
RSM2
Power cables, 2 / RSM
Bdot and thermal breakers (from RSMs)
Bdot and thermal breakers (to supplies)
Termination Box

## Slide 42: Verification of Pre-Series

## Slide 43: FAT-I: November 2017
Programme (CDR-2)
Performance validation at several setpoint combinations
Stability tests (~8 h)
Safety: high-voltage isolation
System self-protection: short to ground, excessive parameter setpoints, frequency
Visit to Danfysik, Nov 29-30
Inspection of hardware (limited assembly)
Review of FAT results
Training
Inspection and confirmation of interfaces
AU: Søren Pape Møller, Heine
ESS: Iñigo Alonso, Kent Wigren, Javier Cereijo Garcia

## Slide 45: FAT: Parameters > Acceptance Criteria
Absolute amplitude accuracy, TS 2.10. 3.0%-3.2% (1%)
  Power cable discharge: the power cables will continue to discharge energy in the magnet a bit longer than the supply actually does
  Peak detection: the inter-burst regulation loop relies on peak detection in the current waveform (slight frequency dependence)
  Decision: Accepted. Can be remedied with a calibration (beam response)
Stability, TS 2.12. 1.1%-1.2% (±0.5%)
  Specification is only very marginally exceeded, and is to be compared with the relative beam size (~20%)
  Decision: Accepted
Offset, TS 2.13. 1.3% of full current (<1%)
  The measured offset concerns a specific, single period of the burst. On average the offset is expected to be smaller due to regulation loop.
  1% DC offset would correspond to a beam displacement of ~0.6 mm at the target
  Decision: Accepted, tolerance exceedance has only minuscule effect
Notes: Amplitude accuracy: no regulation at load, remote sense wires
Beam position stability at target

## Slide 46: Site Acceptance Test Results

## Slide 47: SAT Test Steps
System installation
System IO
Bdot characterization
EVR from ESS
  Synchronization
  Polarity
Temperature stability
Long-term testing: stability, reliability
(Stress test of supply): not typical application
Notes: Focus on tests not already conducted at vendor

## Slide 48: System Installation in Aarhus Test Lab
2017.12.13: System received
RSMs mounted on girder
Cables coiled on pallet
Supply crates in boxes
Chamber: delayed!
Support stand columns on pallet
Installation experiences:
Straightforward
Power cables are indeed stiff, 2 man.days / 2 RSMs
OCs are heavy (cap. banks): rack trays
2017.12.22:
System installed
Wfm.gen for pre-triggers
2018.01.22: EVR operational
2 x Pre-triggers
2 x Polarity (constant)
2018.03.14
Ceramic chamber
RSM1 coil replacement (v2 -> v3)

## Slide 49: SAT Setup: Block Diagram
Notes: PicoScope 4262, 2ch, 16 bit, 5 MHz, 10 MS/s, will sample both waveforms.

## Slide 51: Bdot Waveform Shape: Scope vs. SPICE (DDR)

## Slide 52: Bdot Linearity -> Reconstructed Current Amp.
Normalized Bdot signal = Bdot amplitude / frequency
Norm. Bdot linear function of current amplitude setpoint (NB: log-log).

## Slide 53: Bdot Signal Crosstalk - ~50:1
Notes: Only measured for a RSM pair!

## Slide 54: SAT: Thermal Testing

## Slide 55: Thermal Testing
Thermal imaging: pinpoint hotspots due to inadequate heat dissipation (thermal contact) or excessive resistance (electrical contact).
PT1000 sensors are placed accordingly: termination box backplate, coils, cables, etc.

## Slide 56: Thermal Testing
At 230 A the matching filter resistors run hot and close to thermal breaker temperature (60 deg C).
Bad thermal connection detected and remedied.

## Slide 57: Active Cooling of Termination Box
Not expected: active cooling required to operate > 200 A (340 A maximum).
The production-series will feature water-cooling
As delivered
Heatsink + fans
Notes: Bushings
Does not meet requirements / expectations

## Slide 58: Thermal Testing @ 340 A + WC
Notice termination box!
Coils are now hottest: 43 deg C.
  30 – 40 deg C predicted in DDR

## Slide 59: Thermal Testing @ 340 A + WC
RSM coils
Termination box backplate (edge)
Ambient temperature
Run started at 13:44: > 6 h to reach steady coil temperatures

## Slide 60: Thermal Testing: Coil Impregnation?
No difference in thermal performance can be observed between the two coil designs
Recommendation: accept coil design v.3 for the production-series
v2
v3

## Slide 61: Pre-Series Performance Testing @ AU
Objectives
To test what Danfysik cannot…
Operation similar to application at ESS
Long-term stability (months).
  Nature and frequency of faults?
  Provide knowledge of how to improve and monitor system performance. Mitigate faults?
Tests
Typical operational modes
Bdot waveform “quality” will be used as measure of performance
Both supplies at 230 A, 40.000 kHz, 14 Hz
Random and wear-out failures will take a large accumulated system operation time to probe
Notes:   Unfamiliar design: Assess system availability on actual hardware
  Bottom-up MTBF estimation can be inaccurate
  Failure types
Interruptions

## Slide 62: Long-Term Testing (LTT): Motivation
Failures types
  Early (already probed at Danfysik?)
  Random
  Wear-out
LTT Motivation:
  Unfamiliar design: Assess system availability on actual hardware
  Bottom-up MTBF estimation can be inaccurate
Both supplies at 230 A, 40.000 kHz, 14 Hz
Random and wear-out failures will take a large accumulated system operation time to probe
Assumptions:
  Random failures (will occur with constant failure rate), i.e. the testing can be interrupted
  Failure rates are expected to depend on steady-state component temperatures
  Uninterrupted system operation durations should thus be of a considerable length compared to warm-up times, but interruptions are tolerable
Notes: Test what vendor cannot:
Stability & Failures

## Slide 63: LabVIEW Code
5 Figures of Merit (FoMs) devised:
Reconstructed peak current: cf. calibration
Raster frequency: robust
RMS of filtered Bdot waveform: rectifier + a 5 kHz low-pass filter (“raster speed”)
Relative burst DC offset
Phase: relative to ext. trigger (TTL from EVR)
Notes: Robust
LabVIEW code – signal generator
Droop = 0.3% / 3 ms

## Slide 64: LabVIEW Code: Burst & Supply Errors
Bdot burst data acquisition, reduction, validation
  FIFO structure containing events of 2-ch raw burst data
  Reduce to FoMs
  Event classification (accepted / not accepted):
Supply error status: regulation loop error, rejecting pre-trigger, interlocks (overheating)
Notes: Errors are logged and counted

## Slide 65: Long-Term Testing (LTT)
Notes: 50 x 24 hours

## Slide 66: LTT: Run Durations

## Slide 67: System Error Candidate Events
2018.03.20. 12:05: RMS of filtered Bdot > 50% of mean value. Tolerance was empirically set, and Bdot waveform contains no significant artifacts.
  Tolerance is increased to ±100%
2018.04.03. 9:27: Bdot amp. 0.72% below expected value. Tolerance: <±0.5% (ESS spec: <±1%).
2018.04.06. 9:24: Bdot amp. 0.99% below expected value. Tolerance: <±0.5% (ESS spec: <±1%).
  Detected amplitude error is small (within spec)
  Both events occurred during PC login: bad BNC connector + desk movement

## Slide 68: Bdot Signal Integrity?
The Bdot signal has caveats:
  Inherently derivative only
  Detector + cable response?
  Temperature-dependent, ~1% / 10 K?
Bdot response should ideally be evaluated with a reference signal
Depending on Bdot amp. acceptance tolerance, pre-heating system may be relevant
Bdot signal has been used to examine the RSMS to the level of system performance requirements, i.e. very closely
Depending on the application of the Bdot signal at ESS, acceptance tolerances could be much larger.
Additional (redundant) sensors? DCCT, Idot, …
Notes: Specifically amplitude, waveform shape

## Slide 69: SAT Evaluation
All SAT test steps are passed with expected (or better) performance
Valuable experience is gained
Untested items
  Alternating polarity (not implemented in EVR)
  Crosstalk between H-V RSMs
  Testing with beam
  Actual alignment
Notes: Close to

## Slide 70: Integrated Control System (ICS)

## Slide 71: RSM Supply Algorithm
Notes: Predictive formula for VDCL
Capacitor charging time constant

## Slide 73: ICS
Apart from mechanical interfaces, RSMS has interfaces with
  Timing system EVR: Pre-Trigger, polarity (light guides)
  IOC: configuration, status readback (Ethernet)
  FDU: Bdot (BNC), ISO11-16, etc.
Several hardware signals are readily available for monitoring the system state on a fast time scale
Control Crate

## Slide 74: Control Crate ISOx
ISO Name | ISO Label | RX / TX | Type |  | Description | Interfacing equip.
2 | Pre-Trig | RX | HFBR-2412TZ | RX 1652 | Initiates raster burst | EVR
3 | Polarity | RX | R-2521Z | 640 | H: positive, L: negative + warning (ISO15), freq > 10 kHz: negative | EVR
11 | Beam Run Permit | TX | HFBR-1414TZ | TX 1644 | Current wfm is within specs: beam may be issued | FDU?
12 | Trig Permit | TX | HFBR-1414TZ | TX 1644 | PS ready to start a pulse burst | FDU?
13 | Status | TX | T-1528Z | 1652 | Status available on serial line | FDU?
14 | I-Ready | TX | T-1528Z | 1652 | Output current of previous burst was within 1% of set point | FDU?
15 | Pol Req Missing | TX | T-1528Z | 1652 | Input for ISO3 not accepted | FDU?
16 | Pre Trig Error | TX | T-1528Z | 1652 | Input for ISO2 not accepted | FDU?

## Slide 75: ICS
Timing EVR has been provided to trigger and control polarity of power supplies
  Allowed an early direct verification of potentially most important interface
  Alternating polarity feature is currently not implemented in EVR
Fault-Detection Unit (FDU)
    Multiple hardware signals are made available at the RSMPS to the FDU
    Dec 2017: Visit by Stephane Gabourin (ESS Machine Interlocks Engineer)
    A FDU mockup is being made: aim is to test with pre-series hardware
Notes: RSMPS + FPGA DevKit ~ FDU

## Slide 76: Synchronization

## Slide 77: Synchronization: Burst Timing
Pre-Trig will be used to synchronize the respective wfms by the assumption that
  the 8 x Pre-Trig signals are received synchronous
  the specified usable 3.57 ms raster burst is deterministic in terms of start and end relative to the Pre-Trig signals, independent of raster wfm frequency or other parameters.
  THE(!) potential source of common mode failures
Internal oscillators (only):
  Free-running, not phase-locked
    may result in phase errors ~1/f.IClk, 200 ns is tolerable on combined wfm
  The 8x oscillators are bought to have very similar absolute frequencies.

## Slide 78: Synchronization: Burst Polarity
The wfm polarity is to alternate with burst (beam pulse) number (neutron quality issue)
Critical to sync. polarity within e.g. 4x H RSMs
  Could lead to wfm cancellation
  Implementation of polarity sync. could determine impact of faults

## Slide 79: Polarity Sync. of Waveforms: Implementation?
Parallel: all power supplies are similar
  External polarity master (EVR)
  High signal = Start positive
  Frequency above 10kHz = Start negative
  Low/No signal = Start negative, set the warning signal
Design Change Request (CR)?
  ISO15 of CC currently states polarity input (ISO3) = OK/NOK
  Even better implementation: ISO15 indicates the interpreted polarity requested (Stephane Gabourin)
  Would require consultancy hours (VHDL prog.), 37.5 kDKK, paid by AU
IN
OUT
IN
OUT
IN
OUT
IN
OUT
Pol?

## Slide 80: Safety, Quality & RAMI

## Slide 81: Conventional Hazards
High voltage cables: ±340 A, ±650 V (pulsed)
  Clear labeling and distinction between HV and signal cables
  Enclosing covers, only removable by tools
  Power supply Enable / Interlock (NC) to prevent operation during access
  CR: reduce cover openings significantly!
Mechanical hazards:
  Topple dangers should be avoided if possible
  Girder should be secured in place once aligned (turnbuckles)
Operation and service manual available
  Includes evaluation of installation and maintenance risks and mitigating precautions.

## Slide 82: RSM Failure?
No rastering
fx = fy
Nominal
Figure of merit:
Peak current density (Jmax) on target
Beam outside nominal footprint regions
Loss of 1 H RSM
50% smaller beam size
50% larger beam size

## Slide 83: RAMI
Incident | Type | Occurence | Impact | Time to recover | Preventive measures
1 x power supply failure | Random | < 1 / y | Not harmful for short periods | < 10 min. by adjusting the remaining 3 RSMs. | Monitor internal operating temperatures
8 x power supply failure | Random, common mode (central system failure, configuration) | ? | Single pulse: not harmful | < 10 min. depending on cause | High degree of parallel redundancy, monitor beam (PBI) and magnets (FDU) closely
 |  |  | Multiple pulses: PBW / target failure | 1-6 months? | 
Cable, coil insulation deterioration | Lifetime | 1 / 5-10 y? | Not harmful, assuming a localized sudden event | Tunnel access + 1h – 1day | Regular inspection for embrittlement
Vacuum leak | Random (following replacements, alignment) | ? | Could prevent operation (beam loss on res. gas) | 1-2 days depending on cool-down time and preparation | Monitor pressure near raster section
Water-cooling | Random, wear-out | 1 / 3 - 5 y | Could prevent operation | Tunnel access + 1h – 1day | Rad.hard hoses, regular inspections
Redundant system of raster magnets with strength contingency allows for a quick recovery
The design aims at using rad-resistant materials, where possible. Additionally components, like cables, are placed far away from the beam center and prepared for swift exchange
Notes: Degraded mode
Cooling water!

## Slide 84: Suggested Spare Components
Part | Quantity | Comment
Cable between connection box and bus bars | 32 | Cables may be exposed to radiation causing degradation. The cables are easily exchangeable causing minimum down time.
Coils | 2 | May be exchanged in case of long term degradation of insulation
Ferrites, horizontal | 2 | Ferrites do not degrade. Ferrites are fragile and could potentially be damaged during exchange of a coil
Ferrites, vertical | 2 | As above
Vacuum chambers | 2 | Ceramic chambers are brittle

## Slide 85: Verification
FAT I (pre-series), Supplier
Mechanical, electrical, vacuum
Magnetics: Performance (B(t) waveforms, field quality, synchronization) of magnets + cable + supplies
Validation program suggested by supplier – to be approved by Aarhus University (and ESS)
SAT I (pre-series), Aarhus University
Visual inspection for potential shipping damage
System installation + nominal performance testing
Operation will be assessed in detail and long-term stability tests will be made (<4+6 months)
Typical operation modes + stress tests. Temperature readings (thermocouples + thermography)
FAT II (production series), Supplier
As FAT I + additional points based on findings from pre-series SAT?
SAT II (production series), ESS + Aarhus University
System installation + nominal performance testing
FAT II description pending
Notes: Danfysik ISO9000, provides monthly status reports reports that are shared with ESS

## Slide 86: Top risks @ CDR-2
Top risks
  Late changes in design parameters: max f (neutronics), etc.
  Unexpected effects and parameters not included in simulations:
    cable, stray fields inducing eddy currents, etc.
    Beam losses, activation due to backstreaming radiation from target
  System not performing well under long-term tests (true & false trips)
Planned risk treatment activities:
  Perform cable tests before design phase is finalized
  Pre-series: testing and maturing design
  MPS discussions: Include designer as early as possible

## Slide 87: Conclusions
The pre-series equipment has passed the FAT and SAT test verification procedures (approved at CDR-2)
  Only negligible deviations from expectations
The pre-series has offered an opportunity to verify and mature the design significantly.
  Production procedures refined
  Experience gained from installation, operation, monitoring, etc.
  Lab visits by system stakeholders
Experience:
  Output signal shape
  Supply behavior
Notes: Caught early

## Slide 88: Pre-Series: DDR vs. Hardware?
Feature / Parameter | DDR | Pre-series
External ref. oscillator | ESS Master Oscillator | No ext. reference
Ceramic chamber ID / mm | 80 | 78.0 ± 2.5
PSU Controller | Danfysik System 8500 | Danfysik System 8700
Control Crate optical IO | As described in DDR | As described in Danfysik Interface Specification
Notes: Coil design

## Slide 89: Pre-Series Endgame?
Maximize benefit of Pre-Series test lab at AU
  Additional verification and interface testing?
  Verify system performance after ceramic chamber test coating
  FDU mockup
Decommissioning proposal
  All equipment is sent to Danfysik, where the system will be refurbished to include the production series amendments
  ESS receives 1 pre-series line (1 RSM, 1 power supply, 1 ceramic vacuum chamber, cables, etc.) to be used as complete spare units
  Danfysik gets to reuse remaining parts in production-series and holds all refurbishing costs

## Slide 90: Thank you for your preparation, attention, and valuable comments!

## Slide 91: Extra Slides

## Slide 92: Bdot Waveforms
Bdot wfms: Square voltage wfms, 10-40 kHz
  14 Hz x ~3.5 ms bursts, >400 kS/s
PicoScope 4262, 2ch, 16 bit, 5 MHz will sample both waveforms.
  (2ch + time) x 16 bit
  Using Pre-Trig as Ext Trig would enable detection of timing jitter
Data reduction on local PC (LabVIEW):
  FFT: fundamental freq. and its amplitude
  Statistical values: mean, (RMS), …
  Polarity (sign of first edge of waveform)?
Buffer containing history of previous N wfms (shift register)
  In case of fault, write buffer to disk
  Fault = wfms pars outside operational window

## Slide 93: Raster Pulse: Alternating Pattern Parity
Shea et al, ICANS’15
Notes: Amplitude spectra of neutron intensity fluctuations induced by raster scanning the
proton beam (0.8 Å neutron pulses from the thermal moderator)