## Slide 1: Ion Source and LEBT Beam Diagnostics
Tom Shea
Beam Diagnostics Section Leader
Beam Diagnostics (WP7) Work Package Manager
SRR 1 – Ion Source & LEBT
2018-07-17--18

## Slide 2: Beam Diagnostics
Plasma chamber
Extraction system
Pumping box
Solenoid
2 x steerer
Gate valve
Iris
Diagnostics tank
Gate valve
Solenoid
2 x steerer
Chopper
EMU H+V
Faraday cup
NPM H+V
Doppler shift
2 x TMP
Repeller
Repeller
ACCT
2 x TMP
Commissioning tank (RFQ)
2 x TMP
HV power supply
Isolation transformer
Collimator
Magnetron
H2 bottle
Ion source
LEBT 1
LEBT 2
3 x solenoid
Waveguide
Primary pump
Primary pump
High voltage (HV) platform
Ion source
Low energy beam transport (LEBT)
PSS0 controlled area
and radiation shielding
Front End and Magnets
Cooling water
Vacuum
PSS
Beam Diagnostics
Control system
ACCT
Voltage divider
EMU H+V
Faraday cup
NPM H+V
Doppler shift
HV controls
Ground potential controls
Beam stop
Alternative location of beam instruments for commissioning
N2 bottle
ACCT: AC current transformer
EMU: Emittance measurement unit
NPM: Non-invasive profile monitor
TMP: Turbomolecular pump
Beam Diagnostics
Beam Diagnostics
Beam Diagnostics

## Slide 3: Table of Content
System Requirements
Safety Aspects
Operating procedures and training
Test & Verification
Remaining issues

## Slide 4: System Requirements
Measurement Requirements allocated to 5 system types
Beam Accounting
  Beam Current Monitors (BCM)
  Faraday Cup (FC)
  Doppler (DPL)
Beam Centroid
  Non-invasive profile monitor (NPM)
Beam Distribution
  Non-invasive profile monitor (NPM)
  Emittance measurement unit (EMU)
1 mA accuracy on current
1% uncertainty on ion species fraction
100 µm accuracy on position
10% beam size precision
10% emittance accuracy
Examples of key requirements

## Slide 5: System Components
Most electronics rack-mounted, power via Schuko sockets
Non-invasive Profile Monitor Camera system
Doppler Spectrometer
Beam Current Monitor for Ion Source Supply Current
Emittance Measurement Unit
Actuator on LEBT
Faraday Cup on LEBT

## Slide 6: Safety Aspects
The Beam Diagnostic systems are not classified as safety systems
During IS and LEBT commissioning, they are not yet relied upon to report beam properties that are limited by licensing conditions.
They do contribute to a safe operations culture by providing situational awareness.
  Trustworthy operation is assured by completing calibration prior to beam operations
  The beam diagnostics section members participate in commissioning activities to verify performance with beam
They do present some hazards that are mitigated by by engineering controls. As a primary example, finger-safe enclosures encapsulate:
  voltage/current
  mechanical pinch points

## Slide 7: Operating procedures and training
One user manual per system. Training will be provided prior to first protons.
One high-level operator screen per system
Key beam data consolidated on overview screens
Faraday Cup screen

## Slide 8: Test & Verification
One Verification Plan/report per system
Data in HDF5 linked to each asset (field-replaceable units and systems)
System | FAT | SAT | Installation | Integration | Verification | Verification with beam
BCM |  |  |  |  |  | 
FC |  |  |  |  |  | 
EMU |  |  |  |  |  | 
Doppler |  |  |  |  |  | 
NPM |  |  |  |  |  | 
completed
in progress

## Slide 9: Remaining Issues
week | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | Oct
First rack powered
ICS infrastructure
Verify BCM
Verify EMU (single plane)
Verify FC
Verify NPM and DPL
Beam verification, EMU, NPM, DPL
Beam verification, BCM, FC
Install repaired EMU (second plane)
~4 weeks to integrate and verify all systems
~3 weeks for issue mitigation and training
Beam