## Slide 1: ISRC and LEBT Vacuum System Design
Simone Maria Scolari
Vacuum System Engineer
SRR 1 – Ion Source & LEBT
2018-07-17--18

## Slide 2: System - Vacuum
Plasma chamber
Extraction system
Pumping box
Solenoid
2 x steerer
Gate valve
Iris
Diagnostics tank
Gate valve
Solenoid
2 x steerer
Chopper
EMU H+V
Faraday cup
NPM H+V
Doppler shift
2 x TMP
Repeller
Repeller
ACCT
2 x TMP
Commissioning tank (RFQ)
2 x TMP
HV power supply
Isolation transformer
Collimator
Magnetron
H2 bottle
Ion source
LEBT 1
LEBT 2
3 x solenoid
Waveguide
Primary pump
Primary pump
High voltage (HV) platform
Ion source
Low energy beam transport (LEBT)
PSS0 controlled area
and radiation shielding
Front End and Magnets
Cooling water
Vacuum
PSS
Beam Diagnostics
Control system
ACCT
Voltage divider
EMU H+V
Faraday cup
NPM H+V
Doppler shift
HV controls
Ground potential controls
Beam stop
Alternative location of beam instruments for commissioning
N2 bottle
ACCT: AC current transformer
EMU: Emittance measurement unit
NPM: Non-invasive profile monitor
TMP: Turbomolecular pump

## Slide 3: Table of content
System Requirements (high level)
Specific safety systems (engineered control)
Operating & safety procedures (administrative control)
Test & Verification
  Allowed testing Envelope
Open issues

## Slide 4: System Requirements (high level)
Leak Tightness
2x10-10 mbar l/s for metal connections;
2E-8 mbar l/s for connections where polymeric gaskets are used
Operational pressure
The gas throughput from the LEBT to the RFQ at the LEBT-RFQ mechanical interface shall not exceed 2x10-3 mbar l/s (P = 6E-5 mbar)

## Slide 5: Specific safety systems
No ”on-line” leak monitoring system is envisaged in the design.
Leak Check of the delivery line will be executed at each bottle exchange according to ESS-0012897 – ”ESS Vacuum Handbook Part 4 – Helium Leak Test”.
Bottle exchange will be carried only by trained operator.
A self-venting regulator has been selected since the accidental release of the entire content of hydrogen is not considered hazardous. A 10 l, 100 bar bottle will be used.

## Slide 6: Operation procedures
Primary pump
Turbomolecular pump
Pneumatic angle valve
Sector gate valve
Gauge
Mass Flow controller

## Slide 7: Test & Verification
Vacuum Control System (ICS scope)
Test and Verification is on going according to:
ESS-0087031 - LEBT Vacuum Control System - System Requirement Specification
ESS-0087038 - LEBT Vacuum Control System - Verification and Validation Plan
Leak Check
Leak Check according to ESS-0012897 – ”ESS Vacuum Handbook Part 4 – Helium Leak Test”.
All the existing connection have been leck checked. The beam stop cooling system presents a water leak.
Leak check on gas delivery systems to be executed.

## Slide 8: Outstanding issues
# | System/Comp. | Status | Description | Due date | Remark
1 | Iris cooling line | OK | Polymeric gaskets on cooling lines | N/A | Risk to run the IRIS with polymeric gasket when RFQ will be in place.
2 | Turbomolecular pump | Failure | The pump is not working | N/A | Not a stop for starting operations
3 | H2 and N2 delivery system | Not ready | Gas delivery systems have not been installed yet | End of July | All part have been already procured
4 | Beam stop | Leaking | A water leak has been found | On going | FC can be used; no issue for starting of operation. Beam stop will be substituted with a water cooled flange.

## Slide 9: Questions
First ESS proton is in this bottle!