## Slide 1: J-PARC test of Aperture and Grid prototype
Dr Cyrille Thomas
Beam Diagnostics Physicist
European Spallation Source ERIC
22-11-2018

## Slide 2: Outline
Beam on Target Diagnostics as a suite of diagnostics supporting beam tuning and neutron production
GRID and APerTure Monitor (APTM) prototype:
  From design to installation at J-PARC
  Preliminary results and lessons learned
Concluding remarks

## Slide 3: Beam on Target Diagnostics
Suite of instruments to measure beam properties on target:
    Size, position (IMG)
    Current density distribution (IMG, GRID)
    Beam outside defined aperture (APTM)
  Scale time: Intra-pulse, pulse, cumulative pulses
Objectives:
  Support tuning beam on Target/Dump:
    establish nominal operation: from non-destructive probe pulses to full 125MW (peak) pulse
    protection
  Support beam on target operation (production):
    monitor
    protection

## Slide 4: Beam on Dump
BPM: phase, position, trajectory
  degraded performance with long bunch
BCM: peak current along pulse
APTM: charge loss into the defined aperture
  Measure all charges from MHz sample to minutes
  Can be biased by secondary electrons
IMG: beam transverse 2D profile, peak current density distribution
  Detects small beam size
  Limitation in dynamic range: too large beam size may not be seen
440 MW/cm2
15 MW/cm2
s = 3mm
s = 16mm

## Slide 5: Beam on Dump
Tuning
for every pulse:
Verify beam size predicted by the model
MPS on: check beam size larger than sth (mode dependant)
Check beam current on aperture (protection)
Check trajectory and divergence (protection)
APTM
Screen 1
Screen 2
Screen 3
Screen 4
X1, Y1, sx1, sy1
DX, DY, Dsx, Dsy
X2, Y2, sx2, sy2
Dump
Image 1
Image 2
IMG
APTM (active collaboration)
Current / sectors (MS/s)
Temperature / Thermocouples (sub Hz)
1µA
314K
0.8µA
3µA
4µA
350K
360K
310K
* future collaboration
Screen 3

## Slide 6: Beam on Target
Accelerator region
Target region
Heine Thomsen (ISA, Aarhus University)
TW
PBW
BEW
PBIP
GRID
APTM
IMG (x2): current density
CO
A2T and Raster Magnets
BPM
BCM
NPM
APTM: beam aperture (protection)
BPM: phase, position
BCM: peak current
NPM: beam profile, size, positon
WS: (low power) beam profile, size
APTM
440 MW/cm2
1.5MW/cm2

## Slide 7: Beam on Target
Beam parameters:
Power: 5MW
Peak Current: 62.5mA
Avg Current: 2.5mA
Pulse: 2.86 ms
PBW
BEW
H. GRID
V. GRID
FFT peaks
V: 29kHz
H: 39kHz
APTM
1µA
0.8µA
3µA
4µA
350k
360K
310K
314K
1µA
0.9µA
2µA
4µA
350K
360K
310K
314K
PBW
PBIP
IMG
Max peak density: 50%
Max peak density: 30%

## Slide 8: APTM and Grid Prototype
Build and install a prototype for the APTM and GRID in an environment similar to ESS target.
Collaboration with J-PARC:
  3NBT Dump:
    downstream Al window, in rough vacuum, (not so) close to the Dump
  Benefit from local experience, which has similar GRID and APTM
Collaboration with IMP:
  Experience mechanical engineers
  Synergy with IMP Diagnostics Group

## Slide 9: Time line of the project
09/2017
11/2017
01/2018
03/2018
06/2018
09/2018
10/2018
FAT at Keye
Construction starts
Initial discussion with J-PARC
Proposal Submission
Proposal Approval
ESS sketch design sent to IMP / J-PARC
Discussion with IMP
Final Design Mechanical Assembly
Sub-contract Keye Industry for manufacturing
Delivery
Installation
Purchase Electronics
Start Control system verification test
Tests with beam
Starts Control system integration

## Slide 10
From design to installation

## Slide 11
Mechanical constraints:
  fit APTM and GRID through DN40 port
    miniature design
  Long range travel >500mm
    High rigidity
34mm
630mm
Mechanical Assembly installed on 3NBT Dump

## Slide 12
MTCA hardware: high performance and robustness;
AMC-Pico-8 digitizer: 8 current input channels per board, 1MSPS @ 20-bit; ±1 mA and ±1 µA range;
EPICS controls with ESS in-house developments;
Full scale integration and deployment;
MRF EPICS integrated timing system
Beckhoff EtherCAT hardware for Motion Control; Temperature controller
ESS open-source framework: ECMC;
Gain experience with control system in real environment
Data Acquisition/Control System
MTCA Hardware
3x AMC-Pico-8
MRF Event Receiver
areaDetector dev. support
Vendor driver
EPICS IOC
for data acquisition
Motion Control Hardware
ECMC
EtherCAT driver
EPICS IOC for motion control
High level control applications

## Slide 13
Pico8 AMC integrated in µTCA crate
AreaDetector: EPICS integration
ECAT: Motion Controller
Open-Hardware EPICS integration
Full system verification prior installation at J-PARC
ECAT: Temperature Controller

## Slide 14
Tests:
W and SiC wire signal
Signal from Ni Blades
Signal from Thermocouples
Parasitic signal and noise: electrons, EMC
Cross-talk between GRID and APTM
Motion control in radiation environment
Prototype experience return
signal intensity
comparison with actual model
Real environment perturbation
Direct comparison with ESS
Tests of fully EPICS integrated system

## Slide 15: Scan APTM
300mm
570mm
Scan APTM
Scan beam
APTM from in to out
Grid out then in

## Slide 16
APTM Blade Signal
V Grid: signal
H Grid: signal wire 1 (SiC)
V grid: proton signal + 26kHz
H Grid: only wire 1 showing signal
  Only electrons signal
Blades: strong 26kHz

## Slide 17: Grid in Position
Bias on APTM: -30V
Grid on
APTM 24mm out
Beam: 50mA, 500µs, width = ~2mm
Retrieval of profile

## Slide 18: MC simulation

## Slide 19: MC simulation
Energy deposition
ANSYS thermo-mechanical:
with full beam power
Melt!

## Slide 20: MC simulation
Protons
Electrons
…and many more particles
Total electrons >2keV on wire:
10-4 /proton (Air, other wires, plate, etc.)
Total electrons 25eV out of wire:
SEY = 10-1 /proton
-> peak current in W wire,
at 2mm beam size: 160µA
-> measured: 400µA

## Slide 21: Concluding Remarks
Successful design and manufacturing of APTM and GRID prototype, delivered in time for the installation
Successful Acquisition Electronics system design, built, and integrated into EPICS ESS control system
Successful deployment of full instrument at J-PARC
Successful test in real environment similar to ESS

## Slide 22: Acknowledgment
Thanks to all the people involved in this challenging project:
  ESS ICS group, Julen Etxeberria Malkorra, Joao Paulo, Simone Farina, Han Lee, Peter van Velze, et al., who help and support the integration of the Electronics and the implementation of the control system
  IMP collaborators, HaiHua, Zhangbin, He, who design the assembly and have it manufactured and tested, and their colleagues from diagnostics group who supported us in the design and factory acceptance tests
  ESS BI group, Hinko Kocevar, Edvard Bergman, Viatcheslav Grishin, Clement Derez, Thomas Shea, who help all along the project in all aspects, and Elena Donegani who performed all MC simulations and Thomas Grandsaert for mechanical engineering support (3D model, ANSYS)
  LTH collaborators: Markus Tormanen, Anders J. Johansson

## Slide 23: Last shot quiz: can you explain what you see?