## Slide 1: Implementation of Multi-Grid Thermal Neutron Detector: test at SEQUOIA, SNS
Anton Khaplanov
On behalf of ESS Detector Group and Collaborators
www.europeanspallationsource.se
IEEE, 2018 11 14, Sydney
WP 4.3: Large-Area Detectors
Horizon 2020 grant agreement 676548

## Slide 2: Aim, Outline
MG.SEQ definition and design
MG.SEQ assembly and tests, installation, beam time
MG.SEQ first measurement results
Outlook: further results, Instrument-focused design
Current Large-Area detectors, He3
Multi-Grid B10 detector at ESS
n

## Slide 3: Previous test – at CNCS
MG
He-tubes
1.1m
A.Khaplanov et al. “Multi-Grid Detector for Neutron Spectroscopy: Results Obtained on Time-of-Flight Spectrometer CNCS" https://arxiv.org/abs/1703.03626 2017 JINST 12 P04030

## Slide 4: MG.SEQ Test Background
Energy ranges of the first ESS ToF spectrometers
CSPEC,
0.76 to 15 meV, tested at CNCS
T-REX, 2 to 160 meV,
Tested at SEQUOIA
 | CSPEC | T-REX
Typical initial λ, Å (meV) | 2 to 15 Å (20 to 0.36 meV) | 0.7 to 6.4 Å (160 to 2 meV)

## Slide 5: SEQUOIA at SNS, ORNL
Leading high-resolution thermal/epithermal spectrometer
Vacuum flight chamber, 5.5m sample detector
Space still available for extra detectors

## Slide 6: MG.SEQ Characteristics
Characteristic | Values
Detector area | 0.8 m2
# of grids | 360*
# of wires | 720
# readout channels | 1080*
# voxels | 28800
# of 10B4C layers | 40
10B4C on radial blades | Yes + No
# of detectors | 3 (+1 spare)
Detector weight | 110kg x3
* Assuming 3 “ESS” detectors
1 m
3 detectors (+ 1 spare)
3 modules per detector
40 grids per module
20 cell layers per grid

## Slide 7: MG.SEQ assembly
Assembly in Lund:
Grids
Modules
Wiring
Vessel assembly
testing

## Slide 8: MG.SEQ – ESS and ILL versions
n
n
3 detectors:
Normal blades coated
Also parallel blades coated
Wide grids
Standard Grid (ESS)
Wide Grid (ILL)

## Slide 9: MG.SEQ vacuum and other tests
Vacuum / pressure tested at EWCON
Tested in SNS vacuum tank:
Atmosphere in detector,
Vacuum in tank
Reached 1 to 7 x10-5
Beam test at ILL, CT2
Tested to 1.1 bar overpressure,
10+ cycles
Leak rate measurements at 1 bar overpressure
Vacuum tank test at SNS

## Slide 10: MG.SEQ installed

## Slide 11: Readout
Detector signal connected with flex PCBs
Analogue + Digital Front-end: Mesytec MMR
Back-end:
Mesytec VMMR, Struck SIS3153
360 channels / detector
3x 128 ch digitizers
3x optic links / detector

## Slide 12: Samples and settings used
Sample | Setting | Energies, meV | Characteristics
Vanadium | High Resolution | 2, 3, 4, 5, 6, 7, 8, 9, 10, 12, 14, 16, 18, 20, 25, 30, 35, 40, 50, 60, 70, 80, 90, 100, 120, 140, 160, 180, 200, 225, 250, 275, 300 | Elastic peak, resolution, efficiency
Vanadium | High Flux | 15, 20, 25, 32, 33, 40, 48, 60, 70, 80, 90, 100, 120, 140, 160, 180, 200, 225, 250, 275, 300, 350, 400, 450, 500, 600, 700, 900, 1000, 2000, 2500, 3000, 3500 | Elastic peak, resolution, efficiency
Vanadium | RRM | 120 (20), 300 (31.7), 400 (34), 500 (36.3), 600 (38), 700 (39.6) , 800 (40.8), 900 (41.8), 1000 (42.8), 2000 (48) | 2 Elastic peaks
Vanadium | Choppers stopped open | all | High rate
C4H2I2S | High Flux | 21, 35, 50, 70, 100, 200, 300, 500 | Non-elastic peaks
US | High Flux | 70, 120 | Fast neutrons
US | White beam | Center at 250 | Bragg peaks
H2O | HF/HR | 4, 8, 16, 32, 50, 80, 160, 250, 600, 800 | Quasielasic scattering
Si powder | HF | 20, 36, 60, 250, 600 | Rings
Si crystal | HF/white | 33 / center at 250 | High peak rate
3-sample changer cryostat

## Slide 13: Energy reconstruction from ToF
T0 chopper, 120-600Hz
Fermi chopper, 120-600Hz
Moderator
Monitor 2
Monitor 1
Sample
Multi-Grid
20 m
5.5 m
25-30°
Ei, vi
Ef, vf
Multi-Grid
He3
Transfer sample to neutron
Transfer neutron to sample
Elastic peak
PRELIMINARY

## Slide 14: Repetition Rate Multiplication
162m <336 Hz
108m <252 Hz
40 m
32 m
+Provision of T0 at 19 m
Many ESS instrument will use RRM to best utilize the long pulse
Time – distance structure on T-REX
RRM on SEQUOIA – 2 energies, here 120 and 20 meV

## Slide 15: Elastic line resolution
FWHM
Resolution matches He3
Deviation at 2 to 12 meV
HR
HF
PRELIMINARY

## Slide 16: Detection Efficiency
Counts Relative to He3 counts
Agreement between HR and HF
Drop from 2 to 12meV, due to different chopper settings.
Agreement with expected for 14 to 3000 meV
Efficiency for 0.76 to 80 meV measured previously on CNCS
PRELIMINARY

## Slide 17: Count rate
Rate on a ToF instrument not constant
Large peaks in position and space
Bragg reflections  highest localized rate
Powder line  large rate over large area
n
PRELIMINARY

## Slide 18: Background
PRELIMINARY

## Slide 19: Outlook – further results
Verification of efficiency
Rate analysis – high instantaneous rate, high local rate
Prompt peak, fast neutron events
Water – quasielastic scattering
Detailed peak shape – effects of scattering
Comparison of the 3 detectors
 scattering, background, efficiency

## Slide 20: CSPEC MG Detector
Neutrons
Detector height 3.5m 
  140 grids/module
2 modules per detector
Sample-detector – 3.5 m
26 detectors to cover
0+ deg to 140 deg

## Slide 21: T-REX MG Detector
Detector height 2.2 m 
  88 grids/module
6 module/detector
Sample-detector – 3m
Angle coverage:
  -35 to +145 deg
Neutrons

## Slide 22: Thank you!

## Slide 23: Multi-Grid Efficiency
Optimized grid
1.5μm
0.5μm
1.0μm
n
Layer thicknesses (16 blades):
7 blades 0.5μm,
7 blades 1.0μm,
3 blades 1.5μm
Layer thicknesses (20 blades):
4 blades 1.0μm,
10 blades 1.25μm,
6 blades 2.0μm.
Efficiency as a function of wavelength for MG optimized for cold vs. thermal spectrum
Cold – optimization centered on 4Å
Thermal – optimization centered on 1Å
1 Å optimization (T-REX)
1 Å optimization
4 Å optimization
4 Å optimization (CSEPC)

## Slide 24: MG.SEQ 10B4C layers
In total 40 deposition runs during Dec 2017 and April-June 2018 => 40 days with 1 deposition run per day
In total 5840 blades of 5 different blade types (IN5-double, Inner, Outer, Grande, Petite)
Total coated area: 60 m2 (Total substrate area: 30 m2)

## Slide 25: High statistics / Background
PRELIMINARY
1 detector – pure aluminium grids
2 detectors – standard aluminium