## Slide 1: The Multi-Blade 10B-based detector
the cutting edge of
neutron reflectometry at ESS
Francesco.Messi@nuclear.lu.se

## Slide 2: Overview
Multi-Blade
  Scope and requirements
  Design
  Characteristics
Characterisation
  Gamma and fast-neutron sensitivity (@STF, Sweden)
  Reflectometry (@ISIS, UK)
  Count-rate and uniformity (@PSI, Switzerland)
Next
  Final design: MB-300
  IFE, Norway: MB-IFE

## Slide 3: Multi-Blade: scope and requirements
ESS will be the world’s best source of neutrons for the study of materials
ESS will be 30 times brighter than ILL, the worlds’ best research reactor
ESS will be 10 times more intense than SNS, the world’s most powerful spallation source
IEEE2012
High demands on detectors
ESS aerial photo
Oct 2018

## Slide 4: Multi-Blade: scope and requirements
detector
sample
l1
l2
l3
q
q

## Slide 5: Multi-Blade: design
Multi-Blade
sample
neutrons
MWPC
cathode strips
anode wires
holding frame
10B4C substrate
modular detector
  Inclined geometry  benefit on count-rate and conversion power
two separate, individually tuneable, mediums
benefit on gamma rejection

## Slide 6: Multi-Blade: time story
June 2013
Multi-Blade MB2
November 2015
MB15
March 2017
September 2018
MB16S
MB18
MB18
MB16T
MB16S
MB15
MB1
MB2
MB0
2012
2013
2015
2017
2018

## Slide 7: Characterisation: gamma & fast-neutron
Fast neutron ∼ 10-5
* first characterization for a thermal n detector
Gamma ∼10-8
✓ Agreement with previous work
Thermal neutron ∼0.6
@STF, Lund University – May 2016 and Mar 2017
F. Piscitelli, F. Messi et al, “The Multi-Blade Boron- 10-based Neutron Detector for high intensity Neutron Reflectometry at ESS” JINST (2017)
G. Mauri, F. Messi et al, “Fast neutron sensitivity of neutron detectors based on boron-10 converter layers” JINST (2018)
F. Messi, “Gamma- and Fast Neutron- sensitivity of 10B-based neutron detectors at ESS” IEEE2017 talk
100 keV

## Slide 8: Characterisation: reflectometry
@CRISP, ISIS – Oct 2017
Ir sample
Fe/Ir supermirror
Off-specular scattering
COLLIMATED: 120 min measurement
DIVERGENT: 14 min measurement
q-dynamic range ~5 OfM
DIVERGENT
COLLIMATED
0.5mm resolution MB detector
2mm resolution detector
conventional 0D detector
F. Piscitelli, F. Messi et al, “Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS for neutron reflectometry at ESS” JINST (2018)
G. Mauri, F. Messi et al, “Neutron reflectometry with the Multi-Blade 10B-based detector” PRS (2018)
T. Ilves, Master thesis, “2D Imaging with 10B-Based Multi-Blade Neutron Detector”, Lund University (2018)
A. Backis Master thesis, “ Neutron Reflectometry performed on an Iridium Sample using the Multi-Blade Detector ” Lund University (2018)

## Slide 9: Characterisation: count-rate & uniformity
@AMOR, PSI – on-going now
shadow
No shadow
Multi-Blade
sample
neutrons
MWPC
DAQ saturated @1.3 103 n / mm2 / s
Not an easy measure: geometry problem
The detector must sit at the designed distance from the sample !!!

## Slide 10: Characterisation: summary
Under
measurement
Under
measurement

## Slide 11: Next: MB-300 and MB-IFE
Active area: 300 x 508 mm2
Number of blades: 50
Number of channels: 4800

## Slide 12: Summary
A long prototyping development (almost) completed
(Almost) all requirements fulfilled
Officially accepted technology for reflectometry at ESS
Request to provide a MB detector @IFE
Bibliography
Under
measurement
Under
measurement

## Slide 13: Bibliography
F. Messi, G. Mauri and F. Piscitelli, “The Multi-Blade: The 10B-based neutron detector for reflectometry at ESS”. In: Nuclear Instruments and Methods in Physics Research Section A (2018).
doi: 10.1016/j.nima.2018.10.058. url: https://www.sciencedirect.com/science/article/pii/S016890021831372X?via%3Dihub.
G. Mauri, F. Messi, et al. “Neutron reflectometry with the Multi-Blade 10B-based detector”. In: Proceedings of the Royal Society of London A: Mathematical, Physical and Engineering Sciences 474.2216 (2018).
doi: 10.1098/rspa.2018.0266. url: http://rspa.royalsocietypublishing.org/content/474/2216/20180266.
G. Mauri, F. Messi, et al. “Fast neutron sensitivity of neutron detectors based on boron-10 converter layers”. In: Journal of Instrumentation 13.03 (2018), P03004.
doi: 10.1088/1748-0221/13/03/P03004. url: http://stacks.iop.org/1748-0221/13/i=03/a=P03004.
F. Piscitelli, G. Mauri, F. Messi, et al. “Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS for neutron reflectometry at ESS”. In: Journal of Instrumentation 13.05 (2018), P05009.
doi: 10.1088/1748-0221/13/05/P05009. url: http://stacks.iop.org/1748-0221/13/i=05/a=P05009.
F. Piscitelli, F. Messi, et al. “The Multi-Blade Boron- 10-based Neutron Detector for high intensity Neutron Reflectometry at ESS”. In: Journal of Instrumentation 12.03 (2017), P03013.
doi: 10.1088/1748-0221/12/03/P03013. url: http://stacks.iop.org/1748-0221/12/i=03/a=P03013.
F. Piscitelli, et al. “Study of a high spatial resolution 10B-based thermal neutron detector for application in neutron reflectometry: the Multi-Blade prototype”.
In: Journal of Instrumentation 9, P03007 (2014);
doi:10.1088/1748-0221/9/03/P03007. url: arXiv:1312.2473v1.

## Slide 14: Happy testing @AMOR, PSI
Thank you for your attention!

## Slide 15: Backup slides

## Slide 16: Characterisation: gamma & fast-neutron
Fast neutron ∼ 10-5
* first characterization for a thermal n detector
Gamma ∼10-8
✓ Agreement with previous work
Thermal neutron ∼0.6
Sensitivity: Cumulative sum of counts
100 keV threshold
@STF, Lund University – May 2016 and Mar 2017

## Slide 17
MB-15
MB-16
MB-18
MB-300

## Slide 18
detector
FEE
BUF
network
CAEN VME digitisers
AREA
(CAEN)
Strips
Wire
Multiplicity 2
disk
Remote access
start/stop acq,
data transfer
and
data visualization
576ch
384ch
Wires
Strip
Multiplicity 2
threshold
T = time stamp
400 1 11 2000
401 1 35 1200
400 1 36 800
…
…
…
…
1268 1 13 1200
1269 1 14 400
1270 1 63 1600
…
…
time stamp
cassette
ch
area
Cluster: 1 wire 2 strips
Cluster: 2 wires 1 strip
(X,Y,T)
(11, 35.4, 400)
(13.2, 63, 1268)
Matlab script
for clustering and visualization
(X,Y)
(Y,T)
MAX AMPLITUDE: (11, 35, 400)
CoG: (11, 35.4, 400)

## Slide 19
ToF (Time-Of-Flight) reflectometer
Reflectometry is a technique to study
SURFACES & INTERFACES
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
l3
l1
l2
q =(4p/l) sin(q)
Incidence angle
Neutron wavelength
Log R =
= reflected
/ incoming
q
l∝1/v = ToF / Distance
l1
l2
l3
detector
sample
l1
l2
l3
q
q
0.5 2 6Å
0.5 6
Line at constant q
q ∝sin(q) / l
Momentum transfer

## Slide 20
sample
q
q
Ir sample
Fe/Si sample
Si sample
DIVERGENT
Multi-Blade
detector
Working mode at CRISP
conventional
l - q encoding
sample
q
q
COLLIMATED
Multi-Blade
detector
high intensity
LOWER RATE
HIGHER RATE

## Slide 21
sample
q
q
Ir sample
Fe/Si sample
Si sample
DIVERGENT
Multi-Blade
detector
sample
q
q
COLLIMATED
Multi-Blade
detector
The spatial resolution and the counting rate capability of the detector is needed to measure in divergent mode
COLLIMATED: 120 min measurement
DIVERGENT: 14 min measurement
DIVERGENT
COLLIMATED
Results Si
LOWER RATE
HIGHER RATE

## Slide 22
sample
q
q
DIVERGENT
Multi-Blade
detector
5 orders of magnitude with a poor shielding for background
q - Dynamic range
Results Si
HIGHER RATE

## Slide 23
Ir sample
Fe/Si sample
Si sample
sample
Results Fe/Si
q
q
ToF (or l)
Specular
Specular
Off-Specular
Off-Specular
0.5 2 6Å
Multi-Blade
detector
The spatial resolution and uniformity is needed to identify the feature of off-specular reflectivity
FeSi
Parallel component of the scattering vector Q.
Investigation of in-plane structures.

## Slide 24
Results Fe/Si
Specular Reflectivity
Supermirror edge
(m=3.8)
Si edge
Correlated roughness
domains from the layers
Spin-flip scattering
signal from the layers
Sample Horizon
Transmitted beam
through the sample
No Transmitted beam,
total reflection
Off-specular scattering from Fe/Si neutron supermirror
All the features of the samples have been clearly identified
MB design is ready to do science