## Slide 1
Online alignment survey for long components hidden under shielding
Sébastien GAUTROT LLB
February the 12th 2019

## Slide 2
WHY?
Instruments specs at ESS:
- 20 m to 150 m of guide shielding
- 60 cm thickness concrete shielding
- 4m foreseen guide housings
- neighboring instruments won’t be built in the same time
If there is a loss of neutrons, where shall we open?
BECAUSE TIME IS MONEY!!!!!

## Slide 3
translation of 5mm OR rotation of 0,1°  Loss of flux ≈ 15%
SOME NUMBERS
30 mm
30 mm
Solution : Online alignment survey

## Slide 4: HOW WITHOUT CONTACT?
Laser + PSD / CCDsensor / camera
Distance photoelectric Sensors
Ultrasound detectors
Magnetic sensors
Inclinometers

## Slide 5
HOW WITH CONTACTS ?
LVDT probe sensors
Linear scales

## Slide 6
Work in a dark and closed environment
Uncertainty (100-500 µm)
Repeatability
Low cost
Easy to install
Radiation resistant
Compact (400X400mm² with guide vessel)
MEASUREMENT CONSTRAINTS

## Slide 7
PRINCIPLE
laser rotation of 0.1° = translation of beam on target
7mm at 4m
90mm at 50m
Absolute or relative measurement
target
laser
Laser + PSD / CCD / camera

## Slide 8: Laser + CCD
For short instruments (to 40m/50m)
0,1° for a 40 m instrument = 70 mm translation
Precision ≤ 100µm and 2.10-3 °
1 CCD 150X80 mm target on the bunker
Two lasers per housing to have translation and rotation
Y
ABSOLUTE REFERENCE
Laser multiplexing
X
Δα
Yo = original adjustment
Ym1 = shifting of the laser 1
Ym2 = shifting of the laser 2
If translation Ym2-Yo2=Ym1-Yo1
If rotation Ym2-Ym1=(X2-X1)sinΔα

## Slide 9: Laser + camera
ABSOLUTE REFERENCE
For instruments up to 80m
1 camera imaging a Ø300 mm target on a wall of the bunker
Y
Two lasers per housing to have translation and rotation

## Slide 10: Laser + CCD or camera
Advantage:
low cost ≈6000€ total for a CCD ≈8000€ total for a camera
Disadvantage :
limited up to 80m instruments
measurement 1 by 1
fine adjustement for lasers far from the sensor.
Impact on shielding?
ABSOLUTE REFERENCE

## Slide 11: Laser + PSD
A reference on the bunker wall and a laser / PSD on each housing.
If we add an inclinometer, we can know if it is a translation or a rotation.
RELATIVE REFERENCE
1 PSD on a wall of the bunker as reference
Chart: Relative/absolute position of the housings
| psd1 | psd2 | psd3 | psd4 | psd5 | psd6 | psd7 | psd8 | psd9 | psd10 | psd11
relative value | 0.1 | -0.1 | 1 | -0.9 | -0.2 | 0.1 | -0.3 | 0.1 | 0.1 | 0.1 | 0.1
absolute value | 0.1 | 0 | 1 | 0.1 | -0.1 | 0 | -0.3 | -0.2 | -0.1 | 0 | 0.1
laser
inclinometer
PSD

## Slide 12: Laser + PSD
Advantage :
easy to position
can be used on all the instruments.
Disadvantage :
price 950€ per housing in the industry.

## Slide 13: Inclinometers
Add to laser solutions
Single?
Only the rotation in 2 axis If rotation or translation = effect on each housing? High precision inclinometer ≈ 0.01°
Need to study more

## Slide 14: Inclinometers
Advantage :
easy to position and to install
can be used on all the instruments
very low price ≈ 200€ per housing
Disadvantage :
can not measure directly a translation
(but may be indirectly)

## Slide 15
Bench Test
Camera C11440-36U Hamamatsu
1920X1200 pixels with 1 pixel = 5,86µm
1pixel correspond to 167µm on target
rotation
Laser 1
Laser 2
Target Surface Ø200mm²
Objective X30 at 500mm

## Slide 16
Some tests
2 Lasers + CCD / camera
Origin X1=312
X2=316
1° X1=512
X2=516
3° X1=896 X2=900
origin
3°
Problem :
no difference between the 2 lasers positions

## Slide 17
The 2 lasers are one the same axis (carter), so no difference!
The ΔY is function of the rotation point and not of the laser position.
If pilar 1 move, so the rotation point is pilar 2

## Slide 18
Laser + CCD or camera +inclinometer
Add a inclinometer and keep only one laser The same as a PSD.
1 2
Y
L1
L2
Problem at long distance to discriminate translation and rotation
At 12m/ 0,01°= 2mm
Need a 0,001° inclinometer (expensive).

## Slide 19
The rotation depend on which pilar moves.
For rotation < 1° tanα = α
Control principle
1
2
Y
L1
L2
1
2
Y
angle - 1 down more than 2
rotation = Y1-Y0 = L2 α
1
2
Y
angle + 2 down more than 1
rotation = Y1-Y0 = L1 α
translation = Y1-Y0
Rotation and translation
YRotation = L1 α or L2 α (sign) translation = ΔY - ΔYrotation
1
2
Y

## Slide 20
Inclinometer
angle ° | U (mV) | Angle manufacturer | Fic caméra | Y | angle measuered
0,000 | 5 | 0 | 56 | 685,5 | 0,000
0,774 | 5,259 | 0,777 | 59 | 483,4 | 0,796
1,548 | 5,52 | 1,56 | 60 | 285,4 | 1,575
1,935 | 5,65 | 1,95 | 61 | 191,0 | 1,947
1,161 | 5,404 | 1,212 | 62 | 373,5 | 1,228
Chart: 
| 5 | 5.2590000000000003 | 5.52 | 5.65 | 5.4039999999999999
angle calc | 0 | 0.79564 | 1.5753 | 1.9472 | 1.2284
Calibration of a Pepperl 0,01° inclinometer
The incertitude of the inclinometer value will induce a incertitude on the translation on the target.
We can calculate it as : Tα = Tα ()
For example : 0,5° at 1,3 m = 0,01°

## Slide 21
Measure of Linearity
Chart: Linearity of laser measurement
| 3.5 | 3 | 2.5 | 2 | 1.5 | 1 | 0.5 | 0
 | 541.403 | 512.397 | 481.974 | 454.083 | 425.284 | 394.714 | 364.4 | 336.269
Angle ° | Y | delta Y pixel | delta Y mm | delta pixel / 0,2° | delta mm / 0,2°
0 | 541,40 | 0 | 0 |  | 
0,2 | 512,40 | 29,00 | 4,84 | 29,01 | 4,84
0,4 | 481,97 | 59,43 | 9,92 | 30,42 | 5,08
0,6 | 454,08 | 87,32 | 14,58 | 27,89 | 4,66
0,8 | 425,28 | 116,12 | 19,39 | 28,80 | 4,81
1 | 394,71 | 146,69 | 24,40 | 30,57 | 5,11
1,2 | 364,4 | 177,0 | 29,56 | 30,31 | 5,06
1,4 | 336,27 | 205,13 | 34,26 | 28,13 | 4,70
 |  |  | Standard deviation | 1,04 | 0,17
Good linearity, with a precision of rotation of ± 0,01°
It is possible to measure small angle deviation < 0,01°
Rotation at 2 m of the target
Laser at 2,7 m of the target

## Slide 22
Online alignment survey for long components hidden under shielding
Stability
 | X | Y
4046 | 658.888 | 205,29
4046 | 658.729 | 205,75
4046 | 658.325 | 205,27
4046 | 658.422 | 205,10
4046 | 658.175 | 204,94
4046 | 658.197 | 205,15
4046 | 657.811 | 205,06
4046 | 658.009 | 205,49
4046 | 657.833 | 205,35
4046 | 657.767 | 205,48
4046 | 657.850 | 205,62
4046 | 657.944 | 205,72
4046 | 658.005 | 205,41
4046 | 658.061 | 205,68
4046 | 657.885 | 205,41
4046 | 657.927 | 205,50
moy (pix) | 658,10 | 205,39
sigma (pixel) | 0,31 | 0,24
sigma (mm) | 0,05 | 0,04
Measure every 10 mn in the same position
Very good stability 40µm on 3 hours

## Slide 23
Precision of the process with a camera
Angle° | Trans mm | INCLINO | angle inclino | difference in mm angle | laserY | difference mm | translation
 |  | 4414 | 0 | 0 | 255,9 | 0,000 | 0
 | 2 | 4414 | 0 | 0 | 243,938 | -1,999 | 1,999
0.1 | 2 | 4376 | 0,114 | 2,53 | 258,582 | 0,449 | 2,078
0.2 | 2 | 4348 | 0,198 | 4,39 | 270,626 | 2,461 | 1,928
0.3 | 2 | 4314 | 0,3 | 6,65 | 284,452 | 4,768 | 1,881
0.4 | 2 | 4278 | 0,408 | 9,04 | 298,11 | 7,049 | 1,995
0.5 | 2 | 4245 | 0,507 | 11,24 | 311,4 | 9,269 | 1,969
0.5 |  | 4245 | 0,507 | 11,24 | 323,371 | 11,268 | -0,030
0.5 | 5 | 4245 | 0,507 | 11,24 | 294,133 | 6,385 | 4,853
 |  |  |  |  |  | standard deviation in mm | 0,0633
Precision of the translation Tx = (+)
For example : 0,5° at 1,3 m = 0,05mm
Good accordance between uncertainty calculation and measure
Precision of the process should be better with a PSD

## Slide 24: conclusion
Several solutions from 200€ to 1500€ per housing of 4m
Mixing of solutions for long instruments
Can be used only in some strategic point
such setup can also be used for manual realignment during shutdown when/if laser tracker systems/ressources not available.
Exemple of Cost
for 50m instrument minimum ≈ 12000€
for 150m instrument maximum ≈ 40000€
The good questions are :
What is the price?
What is the benefit?

## Slide 25: THANK YOU Open to any comment, discussion or information sebastien.gautrot@cea.fr

## Slide 27: Focus on: Laser + PSD / CCD / camera Inclinometers
Incompatible solutions
Magnetic sensor very short distance between sensor and ref (10 mm)
Ultrasound sensor bad uncertainty for 1m between sensor and ref
LVDT probe sensors the sensor musn’t be in contact with the ref
Linear scales Limited distance (30 cm) for target uncertainty
photoelectric Sensors To expensive > 1800€ per housing
TECHNICAL SOLUTIONS