## Slide 1: APTM and GRID Beamline Elements
Dr. Cyrille Thomas
Beam Instrumentation
European Spallation Source ERIC
Date

## Slide 2: Outline
Introduction: Presentation of the systems concept and overview
Concepts for the APTM and for the GRID
BLE and Electronics Projects
Project planning
Concluding remarks

## Slide 3: Concept and Overview
Suite of instruments to measure beam properties on target:
    Size, position (IMG)
    Current density distribution (IMG, GRID)
    Beam outside defined aperture (APTM)
  Scale time: Intra-pulse, pulse, cumulative pulses
Objectives:
  Support beam on target tuning:
    establish nominal operation
    protection
  Support beam on target operation (production):
    monitor
    protection

## Slide 4: Beam on Dump
BPM: phase, position, trajectory
  Blind to long bunch
BCM: peak current along pulse
APTM: charge loss into the defined aperture
  Measure all charges from µs to minutes
  Can be biased by secondary electrons
IMG: beam transverse 2D profile, peak current density distribution
  Detects small beam size
  Limitation in dynamic range: too large beam size may not be seen

## Slide 5: Beam on Dump
Tuning
for every pulse:
Verify beam size predicted by the model
MPS on: check beam size larger than sth (mode dependant)
Check beam current on aperture (protection)
Check trajectory and divergence (protection)
APTM
Screen 1
Screen 2
Screen 4
X1, Y1, sx1, sy1
DX, DY, Dsx, Dsy
X2, Y2, sx2, sy2
Dump
Image 1
Image 2
IMG
APTM
Current / sectors (MS /s)
Temperature / Thermocouples
1µA
0.8µA
3µA
4µA
350k
360k
310k
314k

## Slide 6: Beam on Target
Accelerator region
Target region
Heine Thomsen (ISA, Aarhus University)
TW
PBW
BEW
PBIP
GRID
APTM
IMG (x2): current density
CO
A2T and Raster Magnets
BPM
BCM
NPM
APTM: beam aperture (protection)
BPM: phase, position
BCM: peak current
NPM: beam profile, size, positon
WS: (low power) beam profile, size
APTM
NSW

## Slide 7: Beam on Target
Beam parameters:
Power: 2MW
Peak Current: 62.5mA
Avg Current: 2.6mA
Pulse: 2.86 ms
PBW
BEW
H. GRID
V. GRID
FFT peaks
V: 29kHz
H: 39kHz
APTM
1µA
0.8µA
3µA
4µA
350k
360k
310k
314k
1µA
0.9µA
2µA
4µA
350k
360k
310k
314k
PBW
PBIP
IMG
Max peak density: 50%
Max peak density: 30%

## Slide 8: Concepts for the APTM
APTM in the Target Monolith:
  Beam centering in the target monolith.
  Determination of the total beam current outside of the desired footprint.
  The system must function with full production beam.
  Provided measurements at reduced beam power.
  Baseline: two locations in the monolith
  The opportunity for maintenance and/or replacement occurs every 6 months at the time the major shutdowns.
  Interlock with FBIS
APTM in the Tuning Dump:
  Limit the power deposition near vacuum seals
  Avoid thermal stress in the beam pipe itself.
  One location is foreseen just upstream of the dump.
  Interlock with FBIS
APTM in the NSW:
  Observe the beam halo, protect the shield wall components such as the pipe itself and the BPM located in the centre of the wall.
  Permit tuning and to accommodate risk of damage from the beam
  Required location for the NSW APTM is just upstream of the NSW centre, but still within the NSW itself.
  Interlock with FBIS

## Slide 9: Concepts for the APTM
J-PARC APTM (courtesy S. Meigo)
Thermocouples (type K)
Metallic Blades (Ni)
J-PARC APTM Prototype (ESS / IMP)
Thermocouples: charge induced temperature:
2x per position: gives temperature gradient
Slow (many pulses) but highly sensitive
Metallic Blades: charge induced current
Several blades per side: finer halo mapping
Fast detection, µs scale, but less sensitive

## Slide 10: Concepts for the APTM
Beam: 400 MeV, 50mA, 50µs, 0.4Hz
Beam: 2 GeV, 50mA, 400ns, single shot
Scan APTM
Intense signal:
Full beam intercepting theo. peak current: IB = 2mA
Temperature rise and decay as expected
Strong 26kHz coupling: can detect rastering

## Slide 11: Concepts for the GRID
Measure beam current density 1D profile in both axis
Measure Horizontal and Vertical position of the beam
Detect 1D current density out of nominal condition
Detect position of the edges of the beam out of nominal values
Detect the beam to be rastered
Interlock with the FBIS
Operating range over peak current range condition (6mA to 62.5mA):
  permit fast tuning
  Permit production and monitor beam during production

## Slide 12: Concepts for the GRID
J-PARC APTM (courtesy S. Meigo)
J-PARC GRID Prototype (ESS / IMP)
Three layers of wires: Hor. Vert. and Diag.
Diag. wires for bias
Fast detection, µs scale, µA sensitivity
Detection of current in each wire
Detection of raster

## Slide 13: Concepts for the GRID
Bias on APTM: -30V
Grid on
APTM 24mm out
Beam: 50mA, 500µs, width = ~2mm
Fast sampling: can detect raster
Retrieval of profile

## Slide 14: BLE and Electronics Projects
Two instruments types with many similarities:
Current reading range
Radiation hard environment
Interlock to FBIS (intra-pulse) and to BIS (pulse period or longer)
Same Electronics will be used: one Electronics development serving both instruments types
Same MC simulations to support BLE design
Same time scale, both to be delivered at various times for integration and to be ready for beam on Dump and on Target
BLE project run in-house
Electronics design by LTH collaborators (Anders J Johansson)

## Slide 15: Project, Design, Interfaces

## Slide 16: Operation and Maintenance
Supporting Beam Tuning and Operation:
Assure signal quality at all time: high signal to noise ratio in current range 6-70mA
Assure reliable detection of errant beam condition: fast signal processing (µs range) to permit FBIS trigger
Permit Fast beam tuning: deliver single shot beam and halo properties
Radiation damage and errant beam condition damage: prepare maintenance plan
  NSW and TD: risk of damage from high power small beam, possible activation of the APTM
  PBIP and PBW: regular change of the plugs, due to radiation damage

## Slide 17: Project planning
PDR
Proposal for next test series at J-PARC
2019
2021
2020
Demonstration Electronic platform with ICS supported FMC
MC simulation and Thermo-mechanical model
First Complete Design
4 APTM and the GRID
Production prototype head #2
Final Mechanical Design APTM
Delivery of PBW
Final Electronic Design
CDR
Prototype tests at J-PARC
BLE for NSW, TD, PBIP
Electronics for
NSW, TD, PBIP
Electronic on MEBT Scrappers
Delivery of PBIP
BLE APTM PBW
Beam on Dump
Beam on Target

## Slide 18: Concluding remarks
Conceptual design advanced in many aspect:
  Functionalities defined
  Interfaces identified and in place
  Performance derived from interfaces and functionalities
  Prototypes designed and tested in real and similar condition
Projects Time Schedule and milestones defined and aligned with other dependant projects
Support from collaborations J-PARC and SNS
Note: Constraints of physical integration in Target Monolith (not mentioned): risk (accepted) of performance reduction

## Slide 19: Review Outline
Electronics for the APTM and GRID
Anders J Johansson
Target Plugs for the Beam on Target instrumentation: PBIP and PBW
Mattias Wilborgsson
Control Systems for the APTM and GRID
Joao Paulo Martins
MC calculation for the APTM and GRID
Dr. Elena Donegani