## Slide 1: Electronics for the APTM and GRID
Anders J Johansson, Lund University

## Slide 2: Main requirements 1: Accuracy and speed. (Position accuracy is set by grid geometry.)
Beam Peak Density Measurement Accuracy
  Peak proton beam density at the Proton Beam Window and at the target surface shall be measured with an accuracy of 20% of the nominal values.
Latency to detect errant beam
  Beam-on-target instrumentation shall detect errant beam conditions and will communicate this to the MPS within one pulse period less the MPS latency.
    I.e. before next pulse, giving the MPS time to inhibit that pulse.
Latency to detect damaging beam pulse
  Beam-on-target instrumentation shall detect damaging beam pulse conditions and will communicate this to the MPS within 10 μs after the start of the damaging condition.

## Slide 3: Main requirements 2: Buffers and saved sampling frequency.
Pulse-by-pulse measurement frequency
  A summary of the measurement over an entire pulse shall be acquired and buffered for every pulse at a rate up to 14 pulses per second.
Intra-pulse measurement period
  Waveforms containing intrapulse measurements shall have a sampling interval of 1 microsecond
Pre-trigger buffer depth
  Data acquisition buffers shall store at least 140 pulses of pre-trigger data.

## Slide 4: Main Requirements 3: Bandwidth
Slow tuning pulse length
  All instrumentation devices shall be capable of measuring 100 microsecond pulses at a 1 Hz repetition rate.
Fast tuning pulse length
  All instrumentation devices shall be capable of measuring 5 microsecond pulses at up to a 14 Hz repetition rate.

## Slide 5: Baseline design concepts
APTM and GRID
Damaging beam pulse condition detected in logic (FPGA).
  (GRID) Decided unilaterally from one by one axis measurements, i.e. horizontal and vertical plane can trigger MPS independently.
Errant beam pulse detected in software (under EPICS).
  If permissible.
  Calculations done in CPU in crate.
Complete beam pulse measurement calculated in software.
  This includes fitting to a beam profile, averaging, windowing etc.
  Calculations done in CPU in crate.

## Slide 6: Baseline design: APTM
Temperature sensing
CPU
(IOCs)
Current sensing
Motion control
MPS
EPICS

## Slide 7: Baseline design: GRID
Current sensing: Vertical
CPU
(IOCs)
Current sensing: Horizontal
Bias generator
MPS
EPICS
MPS

## Slide 8: Baseline design: GRID, alternative solution
Current sensing: Vertical
Bias generator
Bias-T
Bias-T
CPU
(IOCs)
MPS
MPS
Current sensing: Horizontal

## Slide 9: MTCA.4 chassis and layout
Chassis selection according to ESS hardware standards.

## Slide 10
1410
1410
CPU
Timing
Timing
MCH
Fan Tray x 2
PSU x 2
CB control on backplane
Timing triggers
MCH supervision
External I/O
Ethernet on backplane
EPICS
230 V AC
Bias
Schematic view of MTCA.4 chassis.
Number of 1410-boards depend on number of channels, up to a maximum of 10 (80 channels).
Final interface solution to MPS tbd, one solution is illustrated.
MPS

## Slide 11: Current sensing: FMC-Pico-1M4
4 channels
20 bit / channel ADC
1 MHz sampling speed
Bipolar +/- 1uA to +/- 10 mA
10 kHz to 300 kHz BW
Simultaneous sampling
Floating inputs (+/- 300V)

## Slide 12: Fast data processing: IOxOS IFC_1410
FMC carrier
CPU: four core dual thread processor NXP QorIQ T2081 @ 1.8 GHz
2 GB DDR3L-1866 SDRAM
FPGA: Xilinx Kintex UltraScale KU040
Dual FMC slots

## Slide 13: Timing: MRF mTCA-EVR-300U
Timing receiver for the ESS standard timing system.

## Slide 14: Temperature sensing: Beckhoff EtherCAT
Standard solution for ESS.

## Slide 15: Motion control: Beckhoff EtherCAT
Standard solution for ESS.

## Slide 16: MPS interface
TBD according to ESS standards.

## Slide 17: Bias generator
TBD, depends on bias range and polarity.
Will be defined after next round of J-PARC experiments.

## Slide 18: J-PARC experiment 2018
Full MTCA-integrated setup.
Used FMC-Pico-1M4 current samplings FMC cards.
Used properiatory CAEN-ELS carrier board (not 1410).
Fully integrated into EPICS.
No MPS functionality.
External bias-T used for testing.

## Slide 19: J-PARC experiment 2018