nBLM verification plan

|  | Name | Role/Title |
| --- | --- | --- |
| Owner | Clement Derrez | <<Role/ Title>> |
| Reviewer | Irena | <<Role/ Title>> |
| Approver | Irena | <<Role/ Title>> |

| Table of content | Page |
| --- | --- |

1.	Introduction	3

2.	System characteristics	3

2.1.	System purpose [3]	3

2.2.	System overview	4

3.	test setup	9

4.	verification identification	10

5.	Differents validations before and after the MEBT installation	16

6.	Other documentation	17

7.	Glossary	17

8.	references	17

9.	Document Revision history	17

# Introduction

This document describes the verification plan of the neutron Beam Loss Monitor (nBLM) System for the ESS LINAC. In this document, an nBLM testbench and test plan is defined as the complete system needs to be verified before, during and after its installation.

This is the first version for CDR1.2 which will take place on February 12, 2019. This is a preliminary version. As described in document CDR1.2 final [2], each detector will be identified in a database that will include various reception, integration and validation test documents with information on each of the systems that are part of the nBLMs (front end, gas, electronics, complete detector characterization, etc.). These documents as well as the expected values are being finalized and will be updated in the coming months when the various systems are received and during the first assembly / verification test.

# System characteristics

## System purpose [3]

The nBLM system is a new Beam Loss Monitor based on the detection of neutrons using a Micromegas detector. It is complementary to other BLM systems in ESS. It was originally conceived to operate in counting mode, in the region where others BLMs have a small sensibility, i.e. in the low energy region of the accelerator. The expected rate during normal operation discussed during the conceptual design of the system was of few n/s. In addition, in the low energy region, only neutral particles (neutrons and photons) are able to escape from the accelerator. The detector we propose is only sensitive to neutrons and insensitive to gammas, presenting a big advantage compared with icBLM where an x-ray background from the RF could hide or mimic a lost from the accelerator.

Figure 1 : New nBLM detectors design. The “fast” module is seen in the top side of the detector, while the “slow” module is attached below it. As in the previous design both detectors can be used assembly together or separated.

Micromegas detectors are a type of MPGD (MultiPattern Gaseous Detectors) used very extensively in a wide range of particle physics experiments, including neutron detection, for example in the nTOF experiment. As indicated by its name it is a gaseous detector where the charged particles arriving to the chamber will ionize the gas and this ionization is amplified and detected in the Micromegas detector.

The expected signal in a Micromegas detector are of few mV. The requirements from ESS for such a system is to be able to give a response in 5 μs and to be able to monitor losses of 100 mW/m.

The detectors we propose are in fact the combination of two detectors, that we call “Slow” and “Fast”, based on their time response. However, they have also another differences between them, mainly determined by the way the neutrons are converted into charged particles that produce a signal in the chamber. A Micromegas detector, as any ionization detector, detects charged particles, therefore, we need to use a convertor in order to detect the neutrons. In the aforementioned document a comparison between the signal formation, the expected sensitivity and the expected function of each one is discussed. They were conceived to have a complementary functionality. In Table 1 the characteristics of each detector are summarized.

Table 1 : Summary of the characteristics of each detector. Note that “En” is the energy of the neutron detected.

## System overview

System overview…

Table 2 : listed of nBLM detectors along the different accelerators sections [3]

Figure 2 : nBLM detector location along the Linac (blue slow detector, red fast detector) [3]

Figure 3 : Schematic overview of the nBLM system for one detector module[3]

| Components list- system |  |  |
| --- | --- | --- |
| Component ID | Acronym | Description |
| AMC | AMC | 1 per up to 6 detector box, IFC1410 with FPGA and CPU |
| Digitizer | FMC | ioXos, 8 channels ADC ADC311, 1 per AMC |
| Timing receiver | EVR | 1 per MTCA chassis, event receiver for trigger, clock, calibration announcement and beam/machine mode information |
| Chassis | MTCA | 1 per system |
| MicroTCA Carrier Hub | MCH | 1 per MTCA chassis |
| MicroTCA Power Supply | PS | 1 per MTCA chassis, 600/1000 W |
| MicroTCA CPU | CPU | 1 per MTCA chassis, intel i7 |
| HV | HV | 1 per n 24 detectors, CAEN A7030 |
| LV | TBD | 1 per 24 detectors, CAEN LV A2519 |
| Rack patch panel signal | PP | 1 per n systems, top of rack |
| Rack | U | 1 per n ACCTs, rack for electronics |
| Rack patch panel for LV | TBD | One per LV card |
| Beam line patch panel ?? | BLPP | 1 per system, on the faraday cup |
| Gas rack | TBD |  |
| LV cables from rack to tunnel | TBD | 4 cards * 4 cables per card |
| LV cables in tunnel | TBD | 84 (1 per detector) |
| Signal cables | TBD | 84 |
| Gas tubes (from gas bottle storage to rack) | TBD | 2 |
| Gas tubes from rack to tunnel | TBD | 12 |

| Components list per detector |  |  |
| --- | --- | --- |
| Component ID | Acronym | Description |
| Front End Electronics | FEE | 1 per detector box |
| Detector unit | nBLM | 1 per detector box, fast or slow unit |
| Beam line patch panel ??? | BLPP | 1 per system, on the faraday cup |
| HV Cables | TBD | 2 per unit |
| LV cables | TBD | 1 per unit |
| Signal cable | TBD | 1 per unit |
| Gas lines in detector | TBD | 2 per unit (1 goes from one to the next) |
| Shielding boxes | TBD | Integrated in unit (42 slow) |
| Mechanic support | TBD | 2xMEBT + long rail DTLs + 1 per detector in rest |
| HV card | TBD | 1 per unit |
| Convertor | TBD | 1 per unit |
| Cathode fast glued | TBD | 1 per unit |
| Cathode slow | TBD | 1 per unit |
| Gas fast connectors | TBD | 2 per unit |

| MEBT FRUs delivery schedule |  |  |  |
| --- | --- | --- | --- |
| Component | Quantity | Delivery schedule | Remarks |
| Detectors | 4, 2 slow and 2 fast units | April | CEA delivery |
| mTCA crate, populated | 1 | 12 weeks after CDR | ICS delivery, apart from FMC which is CEA delivery |
| HV & LV supply & crate, populated | 1 | 4 months after CDR | CEA delivery |
| Gaz distribution system | 1 | September | CEA delivery |
| FEE | 4 | See Philippe Legou’s report | TBD |
| Mechanics support | 2 | 3 months after approval in CDR | TBD |
| He + 10% CO2 bottle | 1 | TBD | TBD |
| Cables HV / LV | TBD | TBD | TBD |
| Cables signal | TBD | TBD | TBD |
| Long Pipe for gas | TBD | TBD | ESS installation? |
| Short gas tubes + fast connectors | TBD | April | TBD |

# test setup

The nBLM verification setup is composed of the instruments listed below, interfaced using EPICS CA.

Test data is stored in Insight under the corresponding object in HDF5 format.

| Test instruments: list here all Gen/Meas instruments used in the system verification. Uncertainties on each test result are traced back to each instrument using this table |  |  |  |
| --- | --- | --- | --- |
| Instrument ID | Instrument type | Link to Database | Used in: |
|  | DMM |  |  |
| TBD | oscilloscope | TBD | Detector, electronic systems validation |
| TBD | Rad source | TBD | Detector validation |
| TBD | MCA | TBD | Detector validation |
| TBD | Multimeter | TBD |  |
| TBD | High voltage probe | TBD | HV card validation |
| TBD | Pulse generator and card test bench | TBD | FEE validation |
| TBD | Desktop HV power supply | TBD | Detector validation |
| TBD | Gas leak detector | TBD | Detector validation |
| TBD | microTCA | TBD | Detector validation |
| TBD | More instruments | TBD | TBD |

# verification identification

| Gas rack Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: TBD |  |  |  |  |
| Note: |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Tightness verification Firm certification for each rack | TBD |  |  |
| nBLM_2 | Visual inspection. Check visually the mechanical assembly for any apparent damage and/or wrong assembly versus assembly drawings | No apparent damage |  |  |
| nBLM_3 | Connect gas rack to power supply. | TBD |  |  |
| nBLM_4 | Check digital inputs. Switch 0 to 1 and 1 to 0 and check PLC readback signal with Siemens TIA Portal. | TBD |  |  |
| nBLM_5 | Repeat nBLM_2 for each digital input channel. | TBD |  |  |
| nBLM_6 | Check analog inputs. Check PLC readback signal and converted signal with Siemens TIA Portal. | TBD |  |  |
| nBLM_7 | Repeat nBLM_4 for each analog input channel. | TBD |  |  |
| nBLM_8 | Check digital outputs. Control channel with Siemens TIA Portal and check the equipment state. | TBD |  |  |
| nBLM_9 | Repeat nBLM_6 for each digital output channel. | TBD |  |  |
| nBLM_10 | Check analog outputs. Control channel (3 values: 0%, 50%, 100%) with Siemens TIA Portal and check the equipment state. | TBD |  |  |
| nBLM_11 | Repeat nBLM_8 for each analog output channel. | TBD |  |  |

| HV supply Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable: Metrology results, any detailed tests report. |  |  |  |  |
| Note: The tests described below must be performed sequentially upon reception of the system |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Visual inspection. Check visually the mechanical assembly for any apparent damage and/or wrong assembly versus assembly drawings | No apparent damage |  |  |
| nBLM_2 | Install HV unit, connect load (detector or ~ 100pF capacitance, apply test voltage on mesh detector. Readout current | V_mesh = 400 V, _i ≤ 10 nA |  |  |
| nBLM_3 | apply test voltage on drift detector. Readout current | V_drift = 600 V, _i ≤ 10 nA |  |  |
| nBLM_4 | apply test voltage on mesh and drift detector. Readout current | V_mesh = 400 V, _i ≤ 10 nA, V_mesh = 400 V, _i ≤ 10 nA |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

| LV supply Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable: Metrology results, any detailed tests report. |  |  |  |  |
| Note: The tests described below must be performed sequentially upon reception of the system |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Visual inspection. Check visually the mechanical assembly for any apparent damage and/or wrong assembly versus assembly drawings | No apparent damage |  |  |
| nBLM_2 | Current readout with LV on and FEE connected | TBD |  |  |
| nBLM_3 | Voltage readout with LV on and FEE connected | TBD |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

| HV mezzanine card Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: TBD |  |  |  |  |
| Note: |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Visual inspection. Check visually the mechanical assembly for any apparent damage and/or wrong assembly versus assembly drawings | No apparent damage |  |  |
| nBLM_2 | Connect HV mezzanine card to power supply. Apply 2kV at V_in , readout V_out | V_in = V_out |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

| Front End Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable. |  |  |  |  |
| Note: The tests described below must be performed sequentially upon reception of the system See Philippe Legou’s document / presentation |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Noise level with LV on | ± 100 mV  ± 1 mV | Noise RMS |  |
| nBLM_2 | Gain of FAMMAS. Connect a pulse generator with the test bench card | TBD |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

| Digitizer unit Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable. |  |  |  |  |
| Note: These tests are performed separately, upon reception of the AMC. Tests results exchange with ICS is tbd. |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | ADC SNR | ADC SNR >65 dBc |  |  |
| nBLM_2 | ADC SFDR | 65 dBc |  |  |
| nBLM_3 | ADC SINAD | 60 dBc |  |  |
| nBLM_4 | ADC ENOB | 11.5 bits |  |  |
| nBLM_5 | Crosstalk | < 0.1% FS |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

| Detector Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable. |  |  |  |  |
| Note: These tests are performed separately, upon reception of the AMC. Tests results exchange with ICS is tbd. |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | PCB MM verification : company certification | TBD |  |  |
| nBLM_2 | Mesh verification (non damage) | TBD |  |  |
| nBLM_3 | Bulk test in air : apply V_mesh = 900V, V_signal = 0V, current readout | I < 10 nA |  |  |
| nBLM_4 | Put cathode : Visual inspection | TBD |  |  |
| nBLM_5 | Detector assembled : repeat step nBLM_3 | TBD |  |  |
| nBLM_6 | Detector assembled : gas tightness verification with a gas leak detector and the gas distribution (crate n°3) | TBD |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

The next tests require a complete system installation.

Prerequisites:

- Full Control System

- Long haul cables installed and tested

- nBLMs gas system

- Timing

- MPS

| Assembled system: detector box, uTCA electronics Cold check out Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable. |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Gas tightness test with the gas distribution | TBD |  |  |
| nBLM_2 | Connect detector to the power supply, verify all electrical parameters at the OPI | TBD |  |  |
| nBLM_3 | For each detector box, Repeat nBLM_2 step | TBD |  |  |
| nBLM_2 | Analyze signal output, excitation is a neutron source, extract MPV (Rise Time). HV & LV on. Use a long signal cable, same as in tunnel | ~30-50 ns |  |  |
| nBLM_3 | Analyze signal output, excitation is a neutron source, extract MPV (Pulse width). HV & LV on. Use a long signal cable, same as in tunnel | ~100- 150 ns |  |  |
| nBLM_4 | Analyze signal output, excitation is a neutron source, extract MPV/mean (Amplitude for the slow/fast). HV & LV on. Use a long signal cable, same as in tunnel | TBD once nominal voltages are fixed |  |  |
| nBLM_5 | Count Rates for slow and fast placing the source at same position for all detectors | TBD by first calibration. Depends on the source |  |  |
| nBLM_6 | Additional possible tests: Gain stability. Depends on PT environmental variations. Under control conditions | Better than 10% |  |  |
| nBLM_7 | Additional possible tests: Gain stability. Depends on PT environmental variations. No control conditions | Better than 30% |  |  |
| nBLM_8 | Baseline fluctuation (sigma) | <= 1mV |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

| System self-test SW version: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable. |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Apply nominal voltages, and check current | <20 nA (above offset) |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

| Tests with beam Part ID: NA |  |  |  |  |
| --- | --- | --- | --- | --- |
| Detailed tests results document: links on CHESS, if applicable. |  |  |  |  |
| Test ID | Verification procedure | Threshold | Result | Pass / Fail |
| nBLM_1 | Apply nominal voltages, and check current | <20 nA (above offset) |  |  |
| nBLM_2 | Analyze signal output, excitation is a neutron source, extract MPV (Rise Time). HV & LV on. Use a long signal cable, same as in tunnel | ~30-50 ns |  |  |
| nBLM_3 | Analyze signal output, excitation is a neutron source, extract MPV (Pulse width). HV & LV on. Use a long signal cable, same as in tunnel | ~100- 150 ns |  |  |
| nBLM_4 | Analyze signal output, excitation is a neutron source, extract MPV/mean (Amplitude for the slow/fast). HV & LV on. Use a long signal cable, same as in tunnel | TBD once nominal voltages are fixed |  |  |
| nBLM_5 | Tune voltages according to local rates | No sparks at normal beam conditions |  |  |
| nBLM_6 | Dead time | N/A |  |  |
| nBLM_7 | Time response for slow and fast | Fast (<10ns, immediate) Slow ~200 µs |  |  |
| nBLM_8 | Count Rates for slow and fast placing the source at same position for all detectors | TBD by first calibration. Depends on the source |  |  |
| nBLM_9 | Additional possible tests: Gain stability. Depends on PT environmental variations. Under control conditions | Better than 10% |  |  |
| nBLM_10 | Additional possible tests: Gain stability. Depends on PT environmental variations. No control conditions | Better than 30% |  |  |
| nBLM_11 | Baseline fluctuation (sigma) | ≤ 1mV |  |  |
| nBLM_12 | Perform pedestal setting: The pedestal value must be set before all other settings. It is configured for each detector separately in two ways: - By checking the pedestal value monitored on the firmware level. - By histogramming the raw data extracted through the DoD feature. |  |  |  |
| nBLM_13 | Tune Neutron algorithm settings: To increase the statistic of detected neutrons, controlled losses are planned to be used to tune these values. The DoD feature is used to extract event amplitude and TOT together with rise time and charge distributions from either raw data or EventInfo data stream. The distributions serve as a guide to select the algorithm settings. Note that the settings must be tuned for each detector separately. |  |  |  |
| nBLM_14 | - Trigger delays / Monitor accumulated loss (or neutron counts) over the beam pulse: (ie. loss over BEAM_ON period inside the pulse period) - For each trigger delay setting an average value for this loss is computed - Report the trigger setting with highest average value |  |  |  |
| nBLM_15 | - Monte Carlo Simulation verification and equivalent loss scaling factor: - Define a set of controlled losses for simulation verification - Compare simulated and measured results. In case of larger discrepancies find the source and modify simulation model accordingly. - Once the simulation geometry model is verified, use the results to define scaling factors. - For each detector connect number of lost protons to measured current for particular loss scenario. - This gives a factor that can be used to calculate "equivalent lost protons" from the measurement during operation. - Here each group of detectors has a loss scenario assigned to. |  |  |  |
| nBLM_16 | - Protection function commissioning - Identify controlled loss scenarios with different loss time evolution (different time constants) - Produce these controlled losses to tune the protection algorithms |  |  |  |
| nBLM_17 | - MP thresholds - Identify a few likely accidental scenarios that are most damaging - Use Monte Carlo simulations (Geant4) coupled with thermo-mechanical simulations (ANSYS) to understand damage potential. - Produce risk matrix which serves as a bassline to select MP thresholds. |  |  |  |
| Review pass/fail results: possible risks assessment when test is not passed etc. Explain why a particular step is ignored (if applicable) |  |  |  |  |

# Differents validations before and after the MEBT installation

| Test | Before | After (with beam ?) |
| --- | --- | --- |
| Gas tightness | Ok |  |
| HV settings | Ok |  |
| Preamplifier settings | Ok |  |
| FEE settings |  |  |
| Distribution gas settings |  |  |
| Detector pedestal |  |  |
| Detector HV settings |  |  |

# Other documentation

What other documentations are required for the SAR ?

# Glossary

| Term | Definition |
| --- | --- |
| FRU | Field replaceable unit |

# references

- <<Sample reference to CHESS document: ESS Document (ESS-00XXXXX)>>

- L. Segui, NBLM PROJECT CDR1.2 – FINAL, https://indico.esss.lu.se/event/1173/, February 2019

- L. Segui, “Modes Covered by nBLM System,” https://indico.esss.lu.se/event/948/, 2017.

- I. D. Kittelmann, “Requirements and Technical Specifications - ESS nBLM System,” https://jira.esss.lu.se/browse/BIG-1245, October 2018.

# Document Revision history

| Revision | Reason for and description of change | Author | Date |
| --- | --- | --- | --- |
| 1 | First issue | Clement Derrez | 2018-10-02 |