## Slide 1
Cryogenic Supply for the Series Test Facility (STF) for SIS100 Magnet Testing
Claus Schroeder and Holger Kollmus
for the CRY Department at GSI
Lund, 07.-08.10.2019

## Slide 2
Lund, 07.-08.10.2019
Overview
  FAIR Project
  Cryogenic Infrastructure @ GSI
      HeSu
      R3B and FoPi plant
  Cryo Testing
    PTF
    STF
STF
  Mechanical concept
  Status
  Additional tasks
  Modifications
Quadrupole testing
String planning
Outline
Notes: 1

## Slide 3
Facility for Antiproton and Ion Research
UNILAC
SIS
ESR
GSI-Today
Protons to Uran
max 1 GeV/u Uran and 2 Gev/u for Neon
GSI in futue / FAIR
Rare isotopic beams
Anti protons
Higher Beam Current
At higher energies
100 m
SIS 100/300
HESR
Super
FRS
NESR
CR
RESR

## Slide 4
SIS100 with SuperFRS and CBM
SIS18
NuSTAR
SIS100, SuperFRS and CBM are using SV magnets supplied by one central plant
SuperFRS
CBM
CBM
SIS 100

## Slide 5
Cryogenic Supply for FAIR
SIS 100
(SIS 300)
SuperFRS
CBM
R3B
Panda
APPA

## Slide 6
Overview of GSI-campus, with the location of the existing PTF and the
STF-area.
existing PTF with annex,
housing compressor and
transformers
given area for STF in
existing hall
new annex building, housing
compressor, cold box, distribution
box and electrical infrastructure
GSI campus
HeSu (Helium Supply plant)
central liquefier for Helium-Dewar
filling
Overview of FAIR cryogenic Infrastructure
at GSI-campus

## Slide 7
Helium Liquefier (HeSu) Operation since 2015
Liquefaction rate ~ 20+ l/h
3000 l LHe storage, 8000 l delivered so far
Decant Station and campus wide recovery system

## Slide 8
R3B Cryo Plant and FoPi Cryo Plant Refurbished and Migrated to UNICOS
TCF50 including new cabinet
Cabinet
UNICOS + WinCC OA
Supported by CERN and INFN:
M. Pezzetti, Ph. Gayet, R. Pengo and B. Gallese

## Slide 9
Cryogenic infrastructure at PTF:
70.000 h operation of the PTF
in operation at GSI since 2005
2.200 h with 60 g/s
90.500 h total operation of the cryo plant
only 300 m³ Helium inventory
40-80% Helium losses/year
Preparation area,
Cryogenic
infrastructure is
located in the
middle of the
hall
Distribution box (blue)
Feed boxes and
universal cryostat
Cryo plant: | PTF
Cooling power @ 4,5 K | 300 W
Helium circulation pump | 60 g/s- 200 g/s
Helium liquefaction rate | 2 g/s
Max. power consumption | 250 kW
PTF Infrastructure

## Slide 10
The Series Test Facility Operating since 2015
1.5 kW 4 K equivalent, 6 g/s liquefaction, four test benches
SIS 100 dipoles, quadrupoles and SuperFRS Magnets can be tested
Testing of SIS100 and SuperFRS local cryogenics

## Slide 11
Coldbox in the cryo room (1.5 kW@4.5K eq.)
Helium tanks and annex building with 1,1 MW water cooling
„Cold End“ in the STF testing hall
Easy excess to the connection area
STF: Overview and Key Paramenters
Cryogenic infrastructure at STF:
25.000 h operation of the STF
sub-cooling in the Feedboxes
in operation since 2015
1.000 m³ Helium inventory @
warm
up to 3.000 l Helium at cold
Cryo plant: | STF
Cooling power @ 4,5 K | 700 W
Cooling power @ 50 K – 80 K | 2000W
Helium liquefaction rate | 6 g/s
Max. power consumption | 434 kW
Notes: 1

## Slide 12
Mechanical concept I: PTF versus STF
Magnet connection at PTF:
2 meters extra length
lot of flange connections
hard to install
needs strut sections at the bellow
needs additional bus bar
Endbox (Endcap) is:
direct mounted to the magnets
needs a crane
Connection area at PTF
Magnet connection at STF:
magnet penetrates FBx and EBx
easy flange connections
easy to install
no bellows needed
direct connection to the main current leads (MCLs)
Endbox (Endcap) is:
movable on the rail system
no crane needed
Connection at STF
Evolution from prototype- to serial-facility

## Slide 13
Mechanical concept II
From first ideas…
…to final concept
Feedbox connection at STF:
save space
easy excess to the connecting zone
as less as possible crane action
less flanges
no bellows, no additional bus bar

## Slide 14
1305+/- 5
1305+/- 5
Mechanical concept III
Rail system enables dipole and quadrupole installations;
high precision was
requested:
± 0,1 mm horizontal
± 1 mrad
Rail system is prepared to also test SIS100-quadrupols

## Slide 15
HMI for the operation of one Feedbox during magnet test
Serial Testing Status
Sketch of process lines
SIS 100 Dipole:
cooling by forced-two-phase-Helium-flow
at entrance subcooled
108 pieces
aperture: 130x60 mm²
ramp rate. 4T/s
max. field: 1,9 T
cold mass length: 3060 mm
over all length: 3700 mm
Notes: <Foliennummer>

## Slide 16
Serial Testing Status
Scatch of test area
Preparation area; incoming and outgoing tests
View in the testing hall with warm calibration magnet in front
Half of the SIS100 Dipole magnets are successfully testet
100 % of MCLs
50 % of SIS100-Dipols
one bypass-line
one Beam-vacuum-chamber
2 cryo-pump-prototyps

## Slide 17
Additional Testing: SIS100 By-Pass Line (Polish In-Kind)
Successfully tested at GSI STF
Series production in started
Production time until 2022

## Slide 18
Additional Testing II: MCL SIS100, Cryo-Adsroption Pump, Beam Pipe Chamner
Jumper-line during installation on the side of one Feedbox
Electrical switch for 14 kA between 2 benches
The 4. bench operated mainly for the serial testing of main current leads (MCLs)
Cryogenic pump
One key item of SIS100 are the cryo-cooled, bended beam-vacuum-cambers

## Slide 19
SIS100-quadrupol at STF
Straight section modules
types 2.123 and 2.13s
  17 modules
  Stand-alone module types are in between two warm sections
  Cryogenic supply through attached by-pass line
Arc termination modules
types 2.4 and 1.E
  10 modules
  US- and DS-termination of the arc section of SIS100
  Cryogenic supply by attached by-pass line
  Additional support for vacuum forces
Special modules
types 2.4x and 1.Ei
  2 modules
  US- and DS-terminations of sector 5 in SIS100
  Additional Injection- and extraction QP-doublets separately supported
  Cryogenic supply by attached by-pass line
Arc section modules
types 2.5 to 2.9D
  54 modules
  Arc section module type incline of the arc-section
  Cryogenic supply in line of the arc section
4 different configurations

## Slide 20
Beam Position Monitor
Vacuum Chamber
Defocusing QD
Upstream unit VQD (QD + sextupole)
Downstream unit SF2B (F2 + steerer)
Beam direction
Cryocatcher
Focusing quadrupole F2
Steering
Magnet - ST
SIS100-quadrupol at STF
SIS100 Quadrupole-module is a highly complex system of two quadrupole magnets and various additional components

## Slide 21
Mounting scheme at Series Test Facility station
1. Magnet units, vacuum chambers, beam diagnostic and other module components available at Bilfinger Noell for prototype assembly.
2. Integration of QDM prototype advanced, module will be delivered in October 2019 to GSI.
3. Testing campaign of the magnet module is planned and will start in November 2019 at GSI series test facility.
SIS100-quadrupol at STF
Assembly of prototype at Bilfinger-Noell

## Slide 22
String plannings
Side view of string
Top view of string
Next big step in preparation is the installation of a string at STF.
After testing the first quadrupole-module we will immediately start building up a string with one quadrupole-module and two dipoles using also existing ring parts as feed-in-line and feed-in-cryostat.
Missing parts (red) as special end-box and connecting part are in design phase.
After first tests with electrical dipole- and quadrupole circles in series, we will be also able to operate the dipole circles and quadrupole circles independently.
During installation we will test all possible installation processes we will also have later during ring installation.
3-D-model of planned string

## Slide 23
Modifications of the STF
Cryoplant was planned from very beginning to serve in a later stage superconducting cw-linac via the String-End-box.
By installing new distribution-box between Cold-box and test-area we will enable cw-linac-tests from the very beginning.
As a first step the new box and 20m transfer line was installed during shutdown-phase beginning 2019.
In 2. step we will install the residual transferlines and needed cryogenic infrastructure in the cw-linac-testbulding beginning October 2019 by DeMaCo.
This infrastructure will be later used to serve also the real cw-linac without any additional modifications.
Cw-linac superconducting accelerator cell
We used some space between cold-box and first distribution box to cut the fourfold- and singelfold-transferline and installed additional distribution box for cw-linac supply.
New box is already implemented to the HMI of the cryoplant.

## Slide 24
Layout properties
Multigap CH cavities
Cavities with short lengths (<1m) and small transverse dimensions (<0.5 m)
Modular construction with 4 cryomodules
Each containing 3 CH cavities, 1 buncher, 2 solenoids
Ea= 7.1 MV/m enables compact linac design
First step  Demonstrator project
Layout of demonstrator cryomodule
Superconducting
9.3 T solenoids
Support frame
Superconducting
217 MHz CH cavity
2.2 m
CW-Linac at GSI

## Slide 25
Thank you for your interest
Magnet side of jumper-line; all bus bars are shorted