## Slide 1
Status of the Multi-Blade detector for reflectometry
ESTIA and FREIA
Francesco Piscitelli
on behalf of the ESS detector group
IKON17
Lund
2019/09/11

## Slide 2
August 2015
TARGET
ESTIA
FREIA
Multi-Blade
Boron-10-based detector
Reflectometers

## Slide 3
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
cathode strips
anode wires
holding frame
10B4C substrate
32 Wires
A cassette (unit)
A cassette (unit)
10B-detector for reflectometers
A cassette (unit)
front
neutrons
130mm
100mm
A cassette (unit)
back

## Slide 4
June 2013
Multi-Blade MB2
November 2015
MB15
March 2017
September 2018
MB16S
MB18
MB18
MB16T
MB16S
MB15
MB1
MB2
MB0
2012
2013
2015
2017
2018
2019
FINAL DESIGN
MB300

## Slide 5
October 2017
CRISP reflectometer @ ISIS
2015 - 2016
ATHOS @ BNC
November 2019
D50 and D17 @ ILL
2015 - today
STF, Lund
November 2018
AMOR reflectometer @ PSI

## Slide 6
efficiency | 44% @ 2.5 Å 57% @ 4. Å 82% @ 12 Å
spatial resolution | 0.5 x 3.5 mm2
uniformity | ±1.5%
stability | 2% (over days)
counting rate capability (peak rate) | >3.5 kHz/mm2 (lower limit) >60kHz / 30mm2 (lower limit)
gamma-ray sensitivity | < 10-7 (with 100keV threshold)
fast neutron sensitivity | < 10-5(with 100keV threshold)
gas gain | 60
overlap | 50% eff. drop in 0.5mm gap
Matching ESS requirements
x3 better than state-of-the-art
x20 better than state-of-the-art
G. Mauri et al., Results of the tests at AMOR (in preparation).
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS, JINST 13 P05009 (2018).
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers. JINST 13 P03004 (2018).
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 P03013 (2017).
F. Piscitelli et al. Study of a high spatial resolution 10B-based thermal neutron detector for neutron reflectometry: the Multi-Blade prototype, JINST 9 P03007 (2014).
Multi-Blade characterization
x100 better than state-of-the-art
as good as state-of-the-art
~

## Slide 7
Fast neutron sensitivity (1 – 10MeV)
100 keV threshold
Thermal-cold neutron efficiency (2.5Å) ∼ 0.44
Fast neutron sensitivity ∼ 10-5
Gamma sensitivity < 10-7
Multi-Blade measurements
Helium-3 measurements
100 keV threshold
0.65 ~ Thermal-cold neutron efficiency (2.5Å)
∼ 10-3 Fast neutron sensitivity
< 10-7 Gamma sensitivity
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers,
ArXiv:1712.05614, JINST 13 P03004 (2018).
G. Mauri et al., Evidence of fast neutron sensitivity for 3HE detectors and comparison with Boron-10 based neutron detectors. Submitted EPJ.
Thermal n
Fast n
Gammas
Thermal n
Fast n
Gammas
x100

## Slide 8
MB detector
neutrons
October 2017 - CRISP reflectometer @ ISIS
MB detector
MB detector
sample
neutrons
neutrons

## Slide 9
q
q
DIVERGENT
conventional
l - q encoding
q
q
high intensity
Scientific results from CRISP: Iridium and Si
Multi-Blade
Multi-Blade
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
sample
sample
COLLIMATED

## Slide 10
Scan in angle 0.2-0.8 deg
0.5 6
(Y,T)
(X,Y)
Thickness | 550 Å
Roughness | 15 Å
Nb | 7.3⋅10-6 Å-2
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q=0.3o
(Y,T)
(X,Y)
q
q
Multi-Blade
Ir sample

## Slide 11
Conventional
Conventional
0.5 6
0.5 2 6Å
q
Dependence of the q-resolution on detector spatial resolution
Conventional data reduction
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q
q
DIVERGENT
Multi-Blade
Ir sample

## Slide 12
Conventional
Incl. Res.
Incl. Res.
Conventional
0.5 6
0.5 2 6Å
Constant q
q + Dq
Same l
q - Dq
Integration at constant q,
correcting q with the position
on the detector
q
q
Conventional data reduction
q-correction
Including detector
resolution
Dependence of the q-resolution on detector spatial resolution
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q
q
DIVERGENT
Multi-Blade
Ir sample
ESTIA
High Intensity Mode

## Slide 13
0.5mm resolution MB detector
2mm resolution detector
conventional 0D detector
0.2 degrees
0.4 degrees
0.8 degrees
Dependence of the q-resolution on detector spatial resolution
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q
q
DIVERGENT
Multi-Blade
sample
ESTIA
High Intensity Mode

## Slide 14
Fe/Si sample
q
q
Multi-Blade
Parallel component of the scattering vector q.
Investigation of in-plane structures.
Specular
Specular
Off-Specular
Off-Specular
(Y,T)
(X,Y)
Specular
Off-Specular
Scientific results from CRISP: Fe/Si supermirror
Off-Specular
Specular
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).

## Slide 15
Off-specular scattering from Fe/Si neutron supermirror
Yoneda wings
Yoneda wings
Sample Horizon
Supermirror edge
(m=3.8)
Supermirror
edge
(from SuperADAM at ILL)
Scientific results from CRISP: Fe/Si supermirror
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
Si edge
Correlated
roughness
domains
Spin-flip scattering
signal

## Slide 16
August 2015
Something we did not like!

## Slide 17
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
sample
Background
Background
0.5 2 6Å
Multi-Blade
detector
(Log scale)

## Slide 18
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Background
Background
Gating between 6ms and 20ms
0.5 2 6Å
(Log scale)
(Log scale)
sample
Multi-Blade
detector

## Slide 19
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Background
Background
Gating between 6ms and 20ms
0.5 2 6Å
(Log scale)
(Log scale)
sample
Multi-Blade
detector
Projection over Y-axis

## Slide 20
2Å
90%
98%
61%
45%
Nominal 10B4C coating 7um
Actual 10B4C coating 4um
0.5Å
100 neutrons
strips Y
wires X
substrate
10B4C
55 neutrons get scattered @ 0.5Å
sample
Multi-Blade
detector
~4⋅10-3

## Slide 21
August 2015
November 2018 - AMOR reflectometer @ PSI

## Slide 22
neutrons
detector
MB detector
Selene guide
sample
flight tube
November 2018 - AMOR reflectometer @ PSI
WP5 Data Acquisition software chain has been developed during BrightnESS (WP5 - i.e. DMSC/Data) and tested @ AMOR

## Slide 23
neutrons
detector
MB detector
Selene guide
sample
flight tube
November 2018 - AMOR reflectometer @ PSI
flight tube
detector
ESTIA
sample

## Slide 24
flight tube
Multi-Blade
Selene guide
sample
FRAME OVERLAP
MIRROR
SAMPLE
SLIT 4
MID FOCAL POINT
MB DETECTOR
4.9 m
L ~ 3.4 m
SLIT 1
FOCAL POINT
CHOPPER
SELENE 1
SELENE 2

## Slide 25
Beam transmitted
Beam blocked
Cd slit to block the beam far from the detector with shutter open
FRAME OVERLAP
MIRROR
SAMPLE
SLIT 4
MID FOCAL POINT
MB DETECTOR
4.9 m
L ~ 3.4 m
SLIT 1
CHOPPER
SELENE 1
SELENE 2

## Slide 26
Beam transmitted
Beam blocked
Y-Projection

## Slide 27
Beam transmitted
Beam blocked
Y-Projection
1mm Al window
50um Al foil

## Slide 28
Y-Projection
~ 3⋅10-5
~ 10-4
~ 2⋅10-5
1mm window
Foil window
Instrument background

## Slide 29
Y-Projection
CRISP (ISIS) demonstrator
4.4um 10B4C
AMOR (PSI) demonstrator
8um 10B4C
100 neutrons
strips Y
wires X
substrate
10B4C
55 neutrons get scattered @ 0.5Å
~4⋅10-3
~ 3⋅10-5
~ 2⋅10-5
Problem Solved!

## Slide 30
Y-Projection
~ 3⋅10-5
~ 2⋅10-5
Foil window
Instrument background
~ 10-3
He-3 detector
Multi-Blade
~ 3⋅10-2
ESTIA Instrument Proposal
x100 better than He-3

## Slide 31
August 2015

## Slide 32
Technical characterization on CRISP
Gas gain variation across wires
Uniformity plot, projection on wires, summed over strips
PHS per wire (1 cassette)
Li alpha
Common Th (black)
±10%
Staircase Th (blue)
±7%
Channel by channel Th (red)
±1.5%
Mechanical defects
Mechanical defects
Well characterized
efficiency drop
at blades overlap
Uniformity

## Slide 33
August 2015

## Slide 34
Selene guide using the ReFocus principle
sample
detector
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
ESTIA Instrument Proposal

## Slide 35
NiTi sample
q
q
DIVERGENT
Multi-Blade
Selene guide using the ReFocus principle
sample
detector
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
Divergence of the beam : 1.6 degrees
ESTIA Instrument Proposal

## Slide 36
NiTi sample
q
q
DIVERGENT
Multi-Blade
Selene guide using the ReFocus principle
sample
detector
Angle 1: 1.4deg
Angle 2: 2.2 deg
Angle 3: 3.2deg
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
Divergence of the beam : 1.6 degrees
2D of the detector
ToF vs vertical coordinate
(x,y)
(ToF,y)

## Slide 37
NiTi sample
q
q
DIVERGENT
Multi-Blade
Selene guide using the ReFocus principle
sample
detector
Angle 1: 1.4deg
Angle 2: 2.2 deg
Angle 3: 3.2deg
Combination of 3 angles
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
Divergence of the beam : 1.6 degrees
(ToF,y)
(q,l)
(q,l)
(q,l)
(q,l)

## Slide 38
Combination of 3 angles
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
(q,l)
(q,l)
ESTIA Instrument Proposal
1
1
2
3
4
5
5
4
3
2
Measured
Simulated

## Slide 39
August 2015

## Slide 40
350um
8mm
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
3.5kHz / 0.35x4 mm2
60kHz / 4x8 mm2
Multi-Blade

## Slide 41
350um
8mm
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
3.5kHz / 0.35x4 mm2
60kHz / 4x8 mm2
Multi-Blade

## Slide 42
350um
8mm
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
3.5kHz / 0.35x4 mm2
60kHz / 4x8 mm2
Multi-Blade
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
D17 @ ILL - He3-based detector
*Courtesy of Philipp Gutfreund

## Slide 43
350um
8mm
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
D17 @ ILL - He3-based detector
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
0.3kHz / 0.3x8 mm2
3.5kHz / 0.35x4 mm2
*Courtesy of Philipp Gutfreund
3kHz / 3x8 mm2
D17 @ ILL
60kHz / 4x8 mm2
Multi-Blade

## Slide 44
350um
8mm
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
D17 @ ILL - He3-based detector
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
0.3kHz / 0.3x8 mm2
3.5kHz / 0.35x4 mm2
*Courtesy of Philipp Gutfreund
3kHz / 3x8 mm2
D17 @ ILL
60kHz / 4x8 mm2
Multi-Blade
>20x

## Slide 45
MB300 for ESTIA & FREIA

## Slide 46
280mm
140mm
10B4C
Multi-Blade for ESTIA & FREIA: MB300
FINAL DESIGN

## Slide 47
flight tube
detector
ESTIA
detector
flight tube
ESTIA
Multi-Blade 300 for ESTIA

## Slide 48
flight tube
detector
ESTIA
detector
flight tube
ESTIA
upgrade
Multi-Blade 300 for ESTIA

## Slide 49
Multi-Blade
Top view
Side view
262mm
Ar/He 1bar+
Ar/CO2 1bar+
Al foil
~600
~900
Multi-Blade 300 for ESTIA

## Slide 50
Multi-Blade
Top view
70mm
262mm
262mm
upgrade
~600
~900
Side view
Multi-Blade 300 for ESTIA

## Slide 51
flight tube
detector
ESTIA
Multi-Blade 300 for ESTIA
sample
user
1 blade
Curvature towards the sample
50 blades
4m
sample
FINAL DESING

## Slide 52
G. Mauri et al., Results of the tests at AMOR (in preparation).
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS,
ArXiv:1803.09589, JINST 13 P05009 (2018).
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector,
ArXiv:1804.03962, Proc. R. Soc A 474: 20180266 (2018).
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers,
ArXiv:1712.05614, JINST 13 P03004 (2018).
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, ArXiv:1701.07623, JINST 12 P03013 (2017).
F. Piscitelli et al., Neutron reflectometry on highly absorbing films and its application to 10B4C-based neutron detectors,
ArXiv:1510.01085v1, Proc. R. Soc. A 472, Issue 2185 (2016).
F. Piscitelli et al. Study of a high spatial resolution 10B-based thermal neutron detector for neutron reflectometry:
the Multi-Blade prototype, ArXiv1312.2473, JINST 9 P03007 (2014).
F. Piscitelli and P. Van Esch, Analytical modeling of thin film neutron converters and its application to thermal
neutron gas detectors, ArXiv:1302.3153, JINST 8 P04020 (2013).
References

## Slide 53
Thank you.

## Slide 54
August 2015
Backup slides

## Slide 55
Si sample
q
q
DIVERGENT
Si sample
q
q
COLLIMATED
The spatial resolution and the counting rate capability of the detector is needed to measure in divergent mode
DIVERGENT
COLLIMATED
LOWER RATE
HIGHER RATE
Multi-Blade
Multi-Blade
120 min measurement
14 min measurement
Scientific results from CRISP: Silicon
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).

## Slide 56
November 2018
- AMOR reflectometer @ PSI
Counting Rate Capability

## Slide 57
Technical characterization on CRISP
measured
calculated
Efficiency
Measured PHS
across wires
0.14 degrees
7Li
alpha
Strip signals
Wire signals
Cluster Multiplicity
Multiplicity 2
Cassette shadowing
due to overlap
(not a dead area)
Shadows removed
BN mask
Gain compensation
by HV in final detector
Garfield simulation
gas gain

## Slide 58
40mm
60mm
sample
Masks
BN mask
Multi-Blade
detector

## Slide 59
40mm
60mm
sample
BN mask
Multi-Blade
detector
(Log scale)
Masks

## Slide 60
40mm
60mm
sample
BN mask
shadow
Multi-Blade
detector
(Log scale)
Masks
There is a shadow

## Slide 61
strips
wires
40mm
60mm
shadow
sample
BN mask
10B4C
substrate
shadow
Multi-Blade
detector
(Log scale)
Masks
There is a shadow

## Slide 62
strips
wires
40mm
60mm
shadow
sample
Results
BN mask
Not a dead area but the overlap avoids dead zones
10B4C
substrate
shadow
shadow
No shadow
Multi-Blade
detector
(Log scale)
There is a shadow

## Slide 63
40mm
60mm
sample
BN mask
shadow
Multi-Blade
detector
(Log scale)
Masks
There is a shadow
The shadow is changing width

## Slide 64
40mm
60mm
sample
BN mask
shadow
Multi-Blade
detector
(Log scale)
Multi-Blade
sample
neutrons
4m radius
We measured at 2.3m!
Masks
There is a shadow
The shadow is changing width

## Slide 65
Masks
(Log scale)
Shadows removed

## Slide 66
Masks
(Log scale)
Shadows removed
Reduced sensitivity

## Slide 67
strips
wires
shadow
sample
Not a dead area but the overlap avoids dead zones
10B4C
substrate
shadow
No shadow
dead area
No shadow = No overlap
Shadow = Overlap
There is a shadow
Multi-Blade
Technical characterization on CRISP
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer, ArXiv 1803.09589, accepted JINST (2018).

## Slide 68
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
Neutron reflectometry measurements on CRISP
(X,Y,T)
(Y,T)
(X,Y)

## Slide 69
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
detector area
detector ToF
Neutron reflectometry measurements on CRISP
(X,Y,T)
(Y,T)
(X,Y)

## Slide 70
Neutron reflectometry measurements on CRISP
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
detector area
detector ToF
(X,Y,T)
(Y,T)
(X,Y)
(q,l)

## Slide 71
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
detector area
detector ToF
Momentum transfer
(X,Y,T)
(q)
(Y,T)
(X,Y)
(q,l)
Neutron reflectometry measurements on CRISP

## Slide 72
August 2015
D17 @ ILL
sample
q
q
Conventional Mode
detector
sample
High-intensity mode
detector
Selene guide using the ReFocus principle
High detector requirements
1 10 100
1 = state-of-the-art
Instantaneous
Rate Capability
several kHz/mm2
1 2 3 4
Reflectometry
Spatial Resolution
Reflectometry
0.5 mm
He3-based detector
64 multi-tube
sample
detector
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
0.3kHz / 0.3x8 mm2
3kHz / 3x8 mm2
2 x 8 mm2 spatial resolution
*Courtesy of Philipp Gutfreund