## Slide 1: VMM3a news
IKON 17 in Lund
September 9th – 12th
Ramsey Al Jebali (ESS)
Alexander Backis (ESS)
Morten Jagd Christensen (DMSC)
Isabelle John (ESS)
Michael Lupberger (University of Bonn)
Francesco Messi (ESS)
Hans Muller (CERN)
Dorothea Pfeiffer (ESS)
Francesco Piscitelli (ESS)
Lucian Scharenberg (CERN, University of Bonn)
Martin Shetty (DMSC)

## Slide 2: VMM3a ASIC What is an ASIC ?
Wikipedia says…
  An application-specific integrated circuit (ASIC) is an integrated circuit (IC) customized for a particular use
  As feature sizes have shrunk and design tools improved over the years, the maximum complexity (and hence functionality) possible in an ASIC has grown from 5000 logic gates to over 100 million.
  Modern ASICs often include entire microprocessors, memory blocks including ROM, RAM, EEPROM, flash memory and other large building blocks
=

## Slide 3: VMM3a ASIC Developed for the ATLAS new small wheel
For high luminosity upgrade of the LHC at CERN, the ATLAS experiment replaces parts of its muon detectors
VMM3a ASIC developed by Brookhaven National Labs for the Micromegas (MPGD) muon detectors
At ESS, VMM3a will be used as readout frontend for NMX Gd-GEM (MPGD)
For instruments using MultiGrid and Multiblade, the VMM3a is also the first readout choice
Iakovidis, Georgios. "The Micromegas project for the ATLAS upgrade." Journal of Instrumentation 8.12 (2013): C12007.
http://www.atlasexperiment.org
small wheels

## Slide 4: VMM3a ASIC Features
130 nm CMOS technology
64 input channels, each w/ preamplifier, shaper, peak detector, several ADCs
Pos. & neg. polarity sensitive
Digital block w/ neighbouring logic, FIFO, multiplexer
Adjustable gain 0.5 – 16 mV/fC
Adjustable shaping time from 25 ns – 200 ns
Input capacitance from few pF – 1 nF
Maximum time resolution ~ 0.2 ns
Maximum hit rate per channel 2MHz
38 bit per hit

## Slide 5: RD51 VMM3a hybrid Version 4.0 2019
Detector tests with VMM3a so far used the RD51 VMM3a hybrid as front end
Focus of tests is to validate the appropriateness of the VMM3a ASIC (analog matching) for different detector technologies at ESS (GEM, MG, MB)
For these tests, front end read with the RD51 SRS readout
Successful integration of VMM3a into the SRS funded by BrightnESS grant as collaborative effort between ESS and CERN
RD51 VMM3a hybrid can be easily be integrated into the ESS readout with the help of an assister card

## Slide 6: RD51 SRS systems
New mini crate and D-Card have been developed in 2019
One SRS system and three VMM3a hybrids (3x128 = 384 channels) available in STF at Lund University for MG test
VMM3a and SRS has been integrated into DMSC DAQ event formation unit for Gd-GEM and Multigrid
Full readout chain from frontend to DMSC available and working
Mini crate
FEC and DVM card

## Slide 7: RD51 VMM3a hybrid Cooling
56.5 C
VMM3a dissipates about 1 W
Cooling necessary to have sufficient life time
Passive convection cooler keeps chip temperature below 56.5 C
In combination with fan or controlled ambient temperature, chip temperature around 45 C
Grooves for thin copper pipes (water or air cooling) part of cooler design
Estimated life time of ~ 10 years for a chip temperature of 50 C

## Slide 8: RD51 VMM3a hybrid ID and geographical position encoding
Since each VMM3a needs specific settings, it is important to have the correct mapping of the chips to the detector position
Each hybrid has ID chip with unique 128 bit ID
Hirose connector has 6 pins that are not needed for readout, these pins are used to encode one of 64 geographical position by grounding different pins
ID and position read out via I2C

## Slide 9: VMM3a noise measurements (digital data), gain 3mV/fC
VMM 0
channel 18
VMM 14
channel 18
Noise distribution varies between chips, channels, polarity
Positive polarity signal has wider noise distribution
Automatic threshold equalization needed for optimal threshold setting
VMM noise performance understood

## Slide 10: VMM3a Automatic ADC offset and slope correction
Uncorrected testpulse ADC distribution
Corrected testpulse ADC distribution
Uncorrected ADC distribution 55 Fe source
Corrected 55 FE source ADC distribution
1 mV/fC
12 mV/fC
Calibration similar to what Francesco showed you in MultiBlade talk earlier. Fully understood !
Only necessary if charge is needed to determine neutron position (center –of-mass)

## Slide 11: VMM3a Automatic time offset and slope correction
VMM #4
Channel #2
Offset -7.9 ns
Slope 0.974
VMM #5
Channel #2
Offset 0.1 ns
Slope 0.952
Ideally:
Offset should be 0 ns, slope 1
Calibration only necessary for instruments needing << 25 ns time resolution (e.g. NMX for uTPC analysis)

## Slide 12: VMM3a configuration
VMM3a very complex chip with a lot of configuration parameters
During detector R&D phase, VMM3a configuration via graphical slow control tool https://gitlab.cern.ch/mguth/VMM-software-RD51
After determination of best settings, porting to EPICS configuration
Most settings identical for all chips
Exception (settings per chip per channel)
  Threshold (online)
  ADC and time correction (offset/slope) are offline corrections
  Slow control generates JSON files, which are read by event formation unit DMSC DAQ
Best slow control settings for the VMM understood

## Slide 13: VMM3a Performance measurements
PhD student Lucian Scharenberg at CERN
Thesis ongoing about VMM3a performance and applications
VMM3a characterized at CERN with GEM detectors and X-rays (Position resolution, rate capabilities, time resolution)

## Slide 14: RD51 VMM3a hybrid Firmware developments
Team at University of Bonn: 1 post doc (Michael Lupberger), 1 PhD student, 1 Master student
Focus on improving rate capabilities
At the moment, VMM3a data reliably read out with 80 MHz double data rate, resulting in 320 Mbit/s, 800 Mbit/s obtainable
Challenge: Ensure unique non-ambiguous time stamp at very high rates
General aim is to have a performant user friendly system that can be operated by non-experts

## Slide 15: VMM3a Gd-GEM for NMX
NMX:
3 detectors
Gd-Cathode with support
GEM with frame and spacers
Detector with VMM3 ASICS
Detector Demonstrator Zita
80 VMM3a per detector (5120 channels per detector)
3 detector pannels
Drawing by Patrik Thuiner

## Slide 16
Beam focused on lower three rows of Cd mask
Clearly resolved holes with minimum separation (centre to centre) of about 2.0 mm horizontally, 1.6 mm vertically, and 1.3 mm diagonally
VMM3a NMX detector test at BNC July 2018
VMM delivering quality data from demonstrator

## Slide 17
Gamma 63°
Gamma 62°
Gamma 60°
Crystal orientation Omega 55°
VMM3a performance sufficient to
resolve weak reflections
Crystal diffraction sample: Triose phosphate isomerase
w/ 2-phosphoglycolate (2PG) inhibitor
VMM3a NMX detector test at D16 at ILL October 2018

## Slide 18
Injecting charge into wires of MultiBlade Boron-10 detector
Injecting charge into strips of MultiBlade Boron-10 detector
Signal on monitoring output
Successful test of VMM3a hybrid with MB
MB: Charge injection into:
  Wires (negative polarity, AC coupled)
  strips (positive polarity, DC coupled)
Successfully read out with VMM3a via analog monitoring output and digital data in continuous mode
Gain 1 mV/fC , 200 ns shaping time
VMM3a Multiblade test at Utgard December 2018

## Slide 19
VMM3a Multigrid test at STF Summer 2019
Started in May to test VMM3a together with Multigrid
Huge team effort between detector group (Ramsey, Alex, Isabell, Doro) and DMSC (Morten, Martin)
As expected, biggest challenges at the beginning: Mapping and noise
Setup with MG , VMM3a and SRS
VMM3a hybrids with adapter for MG

## Slide 20
VMM3a Multigrid test at STF Summer 2019
Analog matching of VMM to MG succesful
After solving mapping and noise issue, successful reconstruction of source position
Wires and grids read out with VMM3a gain of 1mV/fC
Optimization: Matching of detector gain to VMM3a gain to use full range of 10 bit ADC
VMM works on MultiGrid detector!

## Slide 21: Summary
RD51 VMM3a hybrid has been developed and integrated into the SRS as collaborative project between ESS and CERN, funded by BrightnESS grant
After the end of BrightnESS large interest and ongoing collaboration with several RD51 institutes, especially Bonn University (Michael Lupberger)
VMM3a, originally developed for Micro Pattern Gaseous Detectors (MPGD) has been successfully matched to Boron based gaseous detectors (Multiblade and Multigrid)
Full readout chain from frontend via readout to DAQ/DMSC available for tests in Lund
Integration into ESS/STFC backend via assister card foreseen for beginning of 2020