## Slide 1: Overview of Hazard Analysis, Radiation Safety, Radiation Protection & PSS Workflow

## Slide 2: Hazard Identification (HazId) Checklist
Covers all Hazards both conventional and radiological.
Shall be used for equipment, area, system or tasks.
Identifies hazards but does not estimate the risk associated with them, if necessary a risk assessment shall be done afterwards.
Preferably before a risk assessment is done and could also be used as a check afterwards.
Serves as an aid to decide which regulations must be complied with, and which standards could be applicable. Therefore some questions are very similar (but with small differences regarding the above).
Contains ≈ 100 questions covering hazards in 13 different categories.
All in query form. If the answer to a question is Yes, action must be taken -> Escalate hazard to Instrument Hazard Analysis
At this moment as draft but will be available from CHESS soon (if someone want the draft, please contact Maurice Looft at Quality Division).
CE Marking Session with Maurice tomorrow (Weds 11th) at 15:30
in Sten Broman

## Slide 3: Instrument Hazard Analysis & Assessment
Contains both conventional and radiological hazards
  Separate worksheets in same Excel workbook
  Instrument Hazard Analysis Template : ESS-0100583
  Conventional hazard classification to follow ESS-0379511 - Risk Matrix for Conventional Safety Risks
  Radiological hazard classification to follow “GSO” ESS-0000004 – General Safety Objectives
Instrument Hazard Analysis covers hazards for Operation & Maintenance – installation hazards are dealt with in the installation binder using RAMS process.
Operational modes and operating conditions are defined in the instrument ConOps – the IHA should not be inventing new ways of running the instrument!

## Slide 4: Conventional Hazards
Identify hazard
Classify hazard
Identify mitigation, mitigation owner and if/how mitigation has been implemented
Classify residual hazard

## Slide 5: Radiological Hazards
Radiological hazard classification to follow “GSO” ESS-0000004 – General Safety Objectives as guiding principle
Detailed guidelines to performing this analysis for hazards to the worker is under development but instrument teams should at least NOW be :
  Identifying radiological hazards
  Classifying according to frequency
  Determining dose rate or dose for hazard
  Determining mitigations
Design of shielding is based on H1 & H2 scenarios
Radiological hazards need to be reviewed and developed in cooperation with ES&H team. Teams should develop their hazard list and then organise meeting/workshop to review. DON’T LEAVE THIS UNTIL TG3!
NB : Procedure/guidelines for performing hazard analysis when hazard to the public exists is released and found in ESS-0041755 - ESS Guideline for Radiological Hazard Analysis. Currently it is assumed, but not officially determined, that instruments do not pose a hazard to the public.

## Slide 6: Personnel Safety System (PSS)
Design of PSS requires specific, detailed requirements from the instrument team based on the Instrument Hazard Analysis. Note, radiation area monitoring is the responsibility of the RP group, not PSS.
Example:
From the IHA, combined with the design of the instrument, actual requirements for PSS can then be developed e.g. interlocks, formalised search procedure etc
Example of analysis done for NCL in ESS-1419309
This needs to be done by PSS team in cooperation with instrument team – develop the IHA with good mitigation statements and then have a meeting/workshop with PSS to develop requirements. DON’T LEAVE THIS TO TG3!
NCG01 | rad | 135 | Ionizing Radiation | Electron emission from the DTL cavities creating X-rays and gammas | Non-exposed worker in NCG01 area at the start of operation is exposed to radiation | >1mSv/year | H1 | H | - Electron emission is stopped by metal casing but X-rays may pass through vessel wall. - Personal EPD monitors shall be worn in the PSS1 controlled area. - Means shall be available within PSS1 controlled area in order to manually switch off the RF systems in case of emergency. - A formalised search shall be carried out before operation of the linac. This shall be formally linked to the PSS system. - PSS1 shall prevent beam and RF systems operation when personnel are inside the PSS1 controlled area and the removeable waveguide is not removed. | >1mSv/year | <10-6 | L

## Slide 7: Summary
Instrument teams need to ensure that they are considering safety aspects during the whole design process so …
  Start now!
    Safety isn’t an add-on, it needs to be designed in
  Write it down!
    We can’t know that you have done your due diligence on safety if you don’t!
    Don’t worry about the format too much – the content is the key. We have a template and approach for conventional hazards and one is coming for radiological hazards, but don’t wait for it!
  Discuss and review with ES&H teams – we are here to help!
    Have meetings and workshops during the design process – don’t just wait until formal reviews.