## Slide 1: Heavy Shutter Panel Discussion IKON 17
Gabor Laszlo
NSS Lead Instrument Engineer
www.europeanspallationsource.se

## Slide 2: Recent shutter designs
LoKi
TBL
SKADI

## Slide 3: Parameters
Geometry parameters | LoKi | TBL | SKADI
Envelope (WxHxL) | 872mmx2550mmx673mm | 883mmx1740mmx1510mm | 1000mmx2760mmx2470mm
Size of attenuator (WxHxL) | 250mmx650mmx385mm 300kg | 220mmx180mmx1510mm 350kg | 222mmx230mmx1200mm 350kg

## Slide 4: Requirements/Considerations 1.
Radiation
<3µs in the cave or regularly accessed area
<25µs in potentially, but not regularly accessed area (Chopper pit)
The area around the shutter should be accessible (<25µs ) after 2 days cool-down time, or use other solutions to extract the shutter if necessary
  Cask
  RH
Designed for 'hands on' serviceable within 30 days of beam off (full power) if service interval <15years
(Minimum effect to the neighboring instruments.)
https://confluence.esss.lu.se/display/SPD/Stakeholders+and+Requirements

## Slide 5: Requirements/Considerations 2.
Maintenance
Service interval >5 years (radiation hardness)
PSS interface switch positions should be in a position where easy maintenance can be carried out
Maintainable or hot (>25us after 2days) components shall be Fully RH-compatible in the short sector.
The maintenance shall not change the guide alignment.
Standard motion control components
https://confluence.esss.lu.se/display/SPD/Stakeholders+and+Requirements

## Slide 6: Reqirements/Considerations 3.
Operation and PSS
The repeatability of the open position shall be better than 0.05mm
Closing time after close command should be less than 2 minutes.
The shutter should be able to close automatically in case of loss of power or compressed air
The shutter motion control mechanism requires a hard interlock (eg. actuator power removal) to ensure it may not open without approval from the PSS
The PSS detector for the shutter should not be affected by the environment (other instruments and irradiated components) and the upstream shielding will be interlocked.
https://confluence.esss.lu.se/display/SPD/Stakeholders+and+Requirements

## Slide 7: Recommendations/Standards
PSS Switch
1 mechanical switch+ 1 proximity sensor in closed position
Pneumatic control: Standard solution will be provided by ESS
Siemens
3SE5212-0KB01
3SE5112-0KD02
https://confluence.esss.lu.se/display/PS/Hardware+for+PSS
Siemens
Switching element: 3SE6605-3BA
Switching magnet: 3SE6704-3BA

## Slide 8: Test Beam Line Shutter
If the TBL shutter fulfills your special needs,
ESS can customize and procure it for you.
The approximate price without optics, control and switches is below 40 000 EUR.
You have to:
do neutronics calculation
check if it fits

## Slide 9
Thank you!