## Slide 1
MEBT Beam instrumentation & Control Developments at ESS Bilbao
ESS-Bilbao
Accelerator Division
Idoia Mazkiaran Zelaia
(On behalf of ESS Bilbao team)
Warsaw, October – 2019

## Slide 2
2
Introduction
ESS-Bilbao develops the MEBT as in-kind contribution to ESS
LPS
  G. Harper: gharper@essbilbao.org
PBI, CH, Q, BU Control
  I. Mazkiaran: imazkiaran@essbilbao.org
Interceptive PBI: FC, EMU, WS, SC
  A. R. Páramo: arparamo@essbilbao.org
Non-Interceptive PBI: BPMs, BCMs, Chopper
  S. Varnasseri: svarnasseri@essbilbao.org
The MEBT PBI operates at:
3.63 MeV – 62.5 mA (230 kW)
Nominal Operation: 2.86 ms / 14 Hz
  Commissioning Modes:Fast Tuning: 5 μs – 14 Hz,
  Slow Tuning: 50 μs – 1 Hz.

## Slide 3
3
MEBT General Layout

## Slide 4
4
Control HW & SW
µTCA standard:
MCH: NAT-MCH-PHYS-R1
CPUs: IFC_1410 QorIQ T2081 & Kintex UltraScale AMC by IOxOS.
  AM 900/412-99 by Concurrent Technologies.
Digitizer boards:
    3 IOxOS FMC ADC 3117, 20 analog channels, 5 MSPS, 16 bits (FC (1), EMU (2)).
Event Receiver board: MRF MTCA EVR-300U.
Beckhoff for I/O and motion control.
Motion EPICS SW embedded, based on ecmc module (FC, EMU).
Environments & architecture:
Concurrent CPU Hard disk:
Linux system. The MicroTCA CPUs: local hard drive with CentOS 7.
Plans: Embedded Linux (Yocto Project) on diskless configuration for final operations.
SW: FPGA firmware, Linux modules drivers, integration EPICS drivers and modules.
EPICS environment: E3 ESS defined.

## Slide 5
5
Instruments controller: Mtca HW
µTCA technology:
µTCA crates
IFC1410 CPU
NAT-MCH-PHYS
ADC3117 Digitizer
AM 900/412 Concurrent
MRF MTCA EVR-300U

## Slide 6: Interceptive Instruments
WS
FC
EMU
SC
WS
EMU/Gr
EMU/Sl
SC
FC

## Slide 7
FC Actuator
Category | Parameter | Value
Beam Parameters | Beam Energy | 3.63 MeV
 | Beam Current | 62.5 mA
 | Mode I: Fast Tuning | 5 μs – 14 Hz
 | Mode II: Slow Tuning | 50 μs – 1 Hz
 | Beam Power | 230 kW
 | Avg. Beam Power | 16 W
 | Nominal Beam Size | σx=σy=2.5 mm
 | Nominal Beam Power Flux | 6 kW/mm2
Operational Conditions | Temperature 20-200 ºC | Ambient
 | Vacuum He Leak Test Rate | <1⋅10-9 mbar l/s
 | Refrigeration | 25ºC, 5 bar, 2 l/min (Max. 10 Bar)
 | Air Pressure | 6-8 Bar (Max. 10 Bar)
Instrument Characteristics | Weight | 16.6 kg
 | Flange to Axis Distance | 273 mm
 | Position Adjustment | 273 ± 5 mm
 | Actuator Stroke | 80 mm
 | Motion Velocity | > 1 s (< 80 mm/s)
 | Mounting Flange | CF100
 | Mounting Orientation | 45º
 | External Protection | Protective Cover
 | Automatic Retraction | Spring Retraction
Body Cup Head Characteristics | Cup Head Weight | ~890 g
 | Cup Head Diameter | 70 mm
 | Cup Head Beam Aperture | 48 mm
 | Repeller Voltage | +/-1000 V
 | Collector | Graphite SGL R7550
 | Repeller | Copper
 | Refrigerated Body | Steel
 | Insulators | Macor/Alumina
Main Components | Pneumatic Actuator | DSBC-40-80-PPSA-N3
 | ElectroValve | FESTO - VUVG-L18-M52-MT-G14-U-1K7L
 | Air Flow Control Valves | GRLA-1/4-QS-8-D
 | Position Sensor | Festo - SME-8-K-24-S6
 | External MPS Switches | Crouzet - 83 133 035
Notes: Overall overview of the MEBT

## Slide 8
FC Interfaces
Notes: Overall overview of the MEBT

## Slide 9
Faraday Cup HW: Control schematic

## Slide 10
Faraday Cup HW & SW: µTCA platform
Main control IOC and signal acquisition.
µTCA crate: Schroff 3U, for 6 AMC cards, 4 RTMS, redundant Power Supplies, 1 MCH.
MCH: NAT-MCH-PHYS-R1
CPUs: IOxOS IFC_1410 QorIQ T2081 & Kintex UltraScale AMC. (digitizing)
      Concurrent Technologies AM 900/412-99. (FC application)
  Digitizer board: IOxOS FMC ADC 3117, 20 analog channels, 5 MSPS, 16 bits.
  Event Receiver board: MRF MTCA EVR-300U.
External integrated systems controlled from
the previous crate HW:
  FC insert/retraction: solenoid valve.
  Repeller control.
  Interlocks interface.
1. Main FC IOC mebt_fc module.
2. Synchronization IOC: EVR received back plane acquisition trigger signal.
3. Ethercat IOC
4. Slow Interlocks IOC.
Siemens PLC for cooling and slow interlocks functions.

## Slide 11
11
FC Engineering GUIs
Notes: 11

## Slide 12: FC Injector Test
ESS-Bilbao Injector (45 keV-40 mA)
Electrostatic Field Simulation
- 200 V in the axis for Vrep-1 kV
- SE electrons with E up to ~100 eV*
GPT Simulation
-SE Suppression ~500V
* Drexler et al Physical Review A 53-3, 1996
Irradiation

## Slide 13: Emittance Meter Unit
INTEGRATION
DESIGN
The EMU is composed by two slit and grid subsystems separated ~370 mm for measurement of the phase space in the horizontal and vertical directions.
The slits are made of graphite in order to withstand beam irradiation and have an aperture of 100 μm for minimisation of emittance reconstruction error.
24 tungsten wires of 35 μm separated 500 μm are included in a Grid PCB. Then the Grid PCB is assembled in the actuator.
PRODUCTION

## Slide 14: EMU Actuator
Category | Parameter | Value | 
Beam Parameters | Beam Energy | 3.63 MeV | 
 | Beam Current | 62.5 mA | 
 | Mode I: Fast Tuning | 5 μs – 14 Hz | 
 | Mode II: Slow Tuning | 50 μs – 1 Hz | 
 | Beam Power | 230 kW | 
 | Beam Size | SL1, SL2 | σx 3.2 mm σy 3.8 mm
 |  | GR1 | σx 6.37 mm σy 0.35 mm
 |  | GR2 | σx 0.46 mm σy 0.35 mm
 | Nominal Beam Power Flux (Slits) | 3 kW/mm2 | 
 | Nominal Beam Power Flux (GR2) | 2.4 kW/mm2 | 
 | Max. Wire Temperature | 1500 K | 
 | Estimated Signal Read-Out | 0.1-100 μA | 
 |  | SL | GR
Instrument Characteristics | Weight | 15 kg | 14 kg
 | Flange to Axis Distance | 190 mm | 205 mm
 | Hard Limits Stroke | 175 mm | 120 mm
 | Switches Stroke | 155 mm | 100 mm
 | Mounting Flange | 100x140 mm2 | CF100
 | Actuator Pitch | 2 mm/rev | 
 | Motion Velocity | 0.1-10 mm/s | 
 | Vacuum He Leak Test Rate | <2⋅10-10 mbar l/s | 
 | Mounting Orientation | 0/90º | 
 | External Protection | Protective Cover | 
 | Slit Dimmensions (aperture/thickness) | 100x200 μm | 
 | Slit Plates | Mersen 2220 PT | 
 | Max. Grid Bias Voltage | ± 1000 V | 
 | Grid Wire Type | GOODFELLOW_W005131 | 
 | Wire Diam. | 35 μm | 
Main Components | Motor | NANOTEC ST4118M1804-B - Two Phases - Steps: 200 steps/rev - Current: 1.8 A/phase | 
 | Brake | NANOTEC BWA-0 25-5 - Voltage: 24 V | 
 | Encoder | SIKO MSK1000-11-M-E1-2-I-1-1.00 - Voltage: 5 V - Resolution: 1 μm - Reference Signal: 1 mm | 
 | Magnetic Band | SIKO MB100-1-0.3-10-0.01-St-TM-AM-0 | 
 | Limit Switches | OMRON SS-01GL2-3 | 

## Slide 15: EMU Interfaces

## Slide 16
EMU: Slit & Grid system schematics
16

## Slide 17
EMU HW: µTCA platform
µTCA platform for EMU prepared for E3 EPICS environment.
Main control IOC and signal acquisition.
µTCA crate: Schroff 3U, for 6 AMC cards, 4 RTMS, redundant Power Supplies, 1 MCH.
MCH: NAT-MCH-PHYS-R1
CPUs: IOxOS IFC_1410 QorIQ T2081 & Kintex UltraScale AMC. (digitizing)
      Concurrent Technologies AM 900/412-99. (EMU application)
  Digitizer board: 2 IOxOS FMC ADC 3117, 20 analog channels, 5 MSPS, 16 bits.
  Event Receiver board: MRF MTCA EVR-300U.
External integrated systems controlled from the previous HW crate :
  Motion Control: Bus ethercat, linear motion, Beckhoff modules.
  Power supply control for BIAS voltage supplying.
  Additional control functions: plane change, power supplying checking, wire checking, calibration.
  Analogue front-end interrelation for signal conditioning.
  Interlocks interface.
Slow Interlocks Siemens PLC for cooling and slow interlocks functions.
17

## Slide 18
EMU: Front End / Back End
Front End:
Number of channels 24
Minimum input current 100 (200) nA
Maximum input current 100 (200) µA
Bandwidth ~2MHz
Slow tuning mode 50us at 1 Hz
Fast tuning mode 5 us at 14 Hz
Output: Differential +/- 10 V, High Impedance
ADC Resolution 16 bits
Modular Sub-rack with backplane connections
  + Large modularity.
  + Robustness (no internal cabling)
  - Complex assembly design
Back End:
Auxiliar Power supplies for Front End.
BIAS Power Supply.
Features:
1ST STAGE: Transimpedance
2nd STAGE: Voltage Amplification and single-ended to differential
3rd STAGE: Filter and Transmission line driver
Slot 1: Outputs.
Slot 2: Tests control.
Slot 3: Power.
Slot 4: Inputs: EMU Plane.
Slot 5 .. 8: Signal accommodation channels (4 boards x 6 channels).

## Slide 19
19
EMU Engineering GUIs
Notes: 19

## Slide 20: Scrapers
PRODUCTION
DESIGN
SC are designed to remove beam halo at 3 MEBT locations.
Each scraper blade will remove 0.25 % or 625 W of beam under nominal operation.
To withstand beam irradiation TZM plates over a refrigerated steel support are included.
The TZM plate is electrically insulated with a kapton foil and is connected to a triaxial cable for signal acquisition.

## Slide 21: Scraper Interfaces

## Slide 22: Scraper Actuator
Category | Parameter | Value | 
Beam Parameters | Beam Energy | 3.63 MeV | 
 | Beam Current | 62.5 mA | 
 | Nominal Beam Mode | 2.86 ms – 14 Hz | 
 | Beam Power | 230 kW | 
 | Beam Size | SC1&2 | σx 1.8 mm σy 3.2 mm
 |  | SC3&4 | σx 3.1 mm σy 3.8 mm
 |  | SC5&6 | σx 2.5 mm σy 0.9 mm
 | Max. Beam Collimation | 0.25 % (625 W) | 
 | Max. Beam Irradiation (SC5&SC6) | Pc = 80 MW/m2 | 
 | Estimated Signal Read-Out | 500-1000 μA | 
Instrument Characteristics | Weight | 14 kg | 
 | Vacuum He Leak Test Rate | <2⋅10-10 mbar l/s | 
 | Refrigeration | Water at 25ºC, 5 bar, 2 l/min (Max. 10 Bar) | 
 | Flange to Axis Distance | 190 mm | 
 | Hard Limits Stroke | 80 mm | 
 | Switches Stroke | 60 mm | 
 | Actuator Pitch | 2 mm/rev | 
 | Motion Velocity | 0.1-10 mm/s | 
 | Mounting Flange | 100x140 mm2 Rect. Flange | 
 | Mounting Orientation | Vertical | 
 | External Protection | Protective Cover | 
 | Nominal Bias Voltage | ± 100 V | 
 | Max. Bias Voltage | ± 500 V | 
 | Blade Characteristics | TZM Blade (MEBT-SC-1104 B) -54x52x5 mm plate -15º Chamfer | 
 | Cooling Block | Steel | 
 | Insulators | Kapton MT300 | 
Main Components | Motor | NANOTEC ST4118M1804-B - Two Phases - Steps: 200 steps/rev - Current: 1.8 A/phase | 
 | Brake | NANOTEC BWA-0 25-5 - Voltage: 24 V | 
 | Encoder | SIKO MSK1000-11-M-E1-2-I-1-1.00 - Voltage: 5 V - Resolution: 1 μm - Reference Signal: 1 mm | 
 | Magnetic Band | SIKO MB100-1-0.3-10-0.01-St-TM-AM-0 | 
 | Limit Switches | OMRON SS-01GL2-3 | 

## Slide 23
23
Scraper Control Schematic

## Slide 24
Scraper Motion Engineering GUIs

## Slide 25: Wire Scanner
DESIGN
PRODUCTION
The wire scanners are designed for measurement of beam profile at 3 MEBT locations in commissioning modes.
Equipped with 34 μm carbon wire and an alignment toolkit has been manufactured for wire placing.
Each carbon wire is connected to triaxial cabling with biased voltages of ± 100 V for minimisation of signal noises.

## Slide 26: Wire Scanner Actuator
Category | Parameter | Value | 
Beam Parameters | Beam Energy | 3.63 MeV | 
 | Beam Current | 62.5 mA | 
 | Mode I: Fast Tuning | 5 μs – 14 Hz | 
 | Mode II: Slow Tuning | 50 μs – 1 Hz | 
 | Beam Power | 230 kW | 
 | Beam Size | WS1 | σx 2.4 mm σy 1.3 mm
 |  | WS2 | σx 3.9 mm σy 1.4 mm
 |  | WS3 | σx 3.6 mm σy 3.8 mm
 | Max. Wire Temperature | 1500 K | 
 | Nominal Beam Irradiation (for WS1) | Ic = 0.3 A/cm2 Ec = 58 J/cm2 | 
 | Estimated Signal Read-Out | 100-500 μA | 
Instrument Characteristics | Weight | 16 kg | 
 | Vacuum He Leak Test Rate | <2⋅10-10 mbar l/s | 
 | Flange to Axis Distance | 305 mm | 
 | Hard Limits Stroke | 260 mm | 
 | Switches Stroke | 240 mm | 
 | Actuator Pitch | 2 mm/rev | 
 | Motion Velocity | 0.1-10 mm/s | 
 | Mounting Flange | 100x140 mm2 Rect. Flange | 
 | Mounting Orientation | 45º | 
 | External Protection | Protective Cover | 
 | Nominal Bias Voltage | ± 100 V | 
 | Max. Bias Voltage | ± 500 V | 
 | Wire Type | Specialty Carbon Monofilament | 
 | Wire Diam. | 34.5 μm | 
 | Wire Preload | 30 g | 
Main Components | Motor | NANOTEC ST4118M1804-B - Two Phases - Steps: 200 steps/rev - Current: 1.8 A/phase | 
 | Brake | NANOTEC BWA-0 25-5 - Voltage: 24 V | 
 | Encoder | SIKO MSK1000-11-M-E1-2-I-1-1.00 - Voltage: 5 V - Resolution: 1 μm - Reference Signal: 1 mm | 
 | Magnetic Band | SIKO MB100-1-0.3-10-0.01-St-TM-AM-0 | 
 | Limit Switches | OMRON SS-01GL2-3 | 

## Slide 27: BCM in Beam Dump Vessel

## Slide 28
BCM in Combo Box

## Slide 29
Beam Position Monitors (BPM)

## Slide 30
BPM UHV tests

## Slide 31
Return loss of the ports |  |  |  | 
 | S11 (704 MHz) dB | SWR (704 MHz) | BW-3dB (MHz) | Leakage current
Strip 1 | -23,5 | 1,14 | 3037 | 0
Strip 2 | -24,6 | 1,13 | 2829 | 0
Strip 3 | -23,6 | 1,14 | 2965 | 0
Strip 4 | -23,7 | 1,14 | 2981 | 0
 | C-S 3-4 (dB) | C-S 4-2 (dB) | C-S 1-4 (dB)
Coupling | -46,9 | -57,2 | -49,6
BPM RF Parameters

## Slide 32
BPM Frequency response
With the same RF input signal, BPM 4 striplines have the signal peak at the same level.
Rigid wire input: RF signal (704 MHz), variation from 0 to -60 dB
Single strips output shows a linear range of around 50 dB.

## Slide 33
33
Fast Chopper
Chopper control of the pulsers (FID GmbH) proprietary protocol.

## Slide 34
34
Buncher Tuner Mobile
3 Mobile tuners for each Buncher, used at cavity tunning.
Movement EPICS development based on ecmc module.
Actuator: Linear shift mechanism: LSM64-21395-001 (UHV).
Stroke:75 mm.
Pitch: 2,54 mm.
Torque to move 245N load about 0.75Nm.
Stepper motor McLennan 23HT18C230.
12.5:1 spur gearbox.
Renishaw LM10IC010EA10A00 magnetic Encoder: 10µm resolution, index signal 2 mm.
Bekchoff modules for motion control.

## Slide 35
35
Magnets: Steerers Power Supplies
CAENels Bipolar PSs:
22 FAST-PS-2020-400
Current and Voltage-Controlled Bipolar PS.
TCP/IP.
Proprietary protocol encapsulated over TCP/IP frames.
Tests:
Master: CAENels PS.
Slave: EPICS module: bipsquad_MEBT.
AsynDriver, Stream Device.

## Slide 36
36
Magnets LPS and Unipolar PSs
Elytt Unipolar PSs:
Modbus TCP/IP.
Modbus addresses map provided.
Device under manufacturing

## Slide 37
37
Thanks for your attention.
Dziękuję bardzo!
¡Muchas gracias!
Eskerrik asko!