## Slide 1
Multi-Blade detector for reflectometry
ESTIA and FREIA
Francesco Piscitelli
on behalf of the ESS detector group
STAP meeting
Lund
2019/09/24

## Slide 2
August 2015
TARGET
ESTIA
FREIA
Multi-Blade
Boron-10-based detector
Reflectometers

## Slide 3
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
cathode strips
anode wires
holding frame
10B4C substrate
32 Wires
A cassette (unit)
A cassette (unit)
10B-detector for reflectometers
A cassette (unit)
front
neutrons
130mm
100mm
A cassette (unit)
back

## Slide 4
Multi-Blade
sample
neutrons
wires X
substrate
neutrons
θ
strips Y
10B4C
MWPC
q = 5 degrees
A cassette (unit)
10B-detector for reflectometers
A cassette (unit)
front
A cassette (unit)
back
strips Y

## Slide 5
neutrons
neutrons
130mm
100mm
32 Strips
(back)
9 units
32 wires
(front)
vessel

## Slide 6
130mm
100mm
The cassettes (units)
are placed horizontally
Resolution 0.5mm
Resolution 3mm
n

## Slide 7
MB300 for ESTIA & FREIA

## Slide 8
280mm
140mm
10B4C
Multi-Blade for ESTIA & FREIA: MB300
FINAL DESING

## Slide 9
flight tube
detector
ESTIA
detector
flight tube
ESTIA
Upgrade:
Detector 2
Detector 1

## Slide 10
70mm
262mm
262mm
upgrade
~600
~900
Side view
Multi-Blade
Top view
Ar/He 1bar+
Ar/CO2 1bar+
Al foil

## Slide 11
flight tube
detector
ESTIA
Multi-Blade for ESTIA
sample
user
1 blade
Curvature towards the sample
50 blades
4m
sample
FINAL DESING

## Slide 12
For a 30 blades detector: FREIA
 | N blades | N wires / blade | N strips/blade | Blade height (mm) | N w ch | N s ch | N tot ch
full | 30 | 32 | 64 | 294 | 960 | 1920 | 2880
reduced | 25 | 32 | 54 | 248 | 800 | 1350 | 2150
detector
cassette
32 wires
64 strips
xN

## Slide 13
General Features

## Slide 14
efficiency | 44% @ 2.5 Å 57% @ 4. Å 82% @ 12 Å
spatial resolution | 0.5 x 3.5 mm2
uniformity | ±1.5%
stability | 2% (over days)
counting rate capability (peak rate) | >3.5 kHz/mm2 (lower limit) >60kHz / 30mm2 (lower limit)
gamma-ray sensitivity | < 10-7 (with 100keV threshold)
fast neutron sensitivity | < 10-5(with 100keV threshold)
gas gain | 60
overlap | 50% eff. drop in 0.5mm gap
Matching ESS requirements
x3 better than state-of-the-art
x20 better than state-of-the-art
G. Mauri et al., Results of the tests at AMOR (in preparation).
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS, JINST 13 P05009 (2018).
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers. JINST 13 P03004 (2018).
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 P03013 (2017).
F. Piscitelli et al. Study of a high spatial resolution 10B-based thermal neutron detector for neutron reflectometry: the Multi-Blade prototype, JINST 9 P03007 (2014).
Multi-Blade characterization
x100 better than state-of-the-art
as good as state-of-the-art
~

## Slide 15
October 2017
CRISP reflectometer @ ISIS
2015 - 2016
ATHOS @ BNC
November 2019
D50 and D17 @ ILL
2015 - today
STF, Lund
November 2018
AMOR reflectometer @ PSI

## Slide 16
August 2015
https://arxiv.org/abs/1905.12311
Giacomo Mauri on board from March 2017 till Sept 2019, did his PhD thesis with us
Now postdoc at ISIS in the detector group

## Slide 17
Gamma-ray sensitivity
100keV threshold
100keV threshold
<10-7 sensitivity
PHS
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 18
Fast Neutron Sensitivity
100keV threshold
PHS
<10-7 gamma-ray sensitivity
~10-5 fast-n sensitivity
Thermal neutrons
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers. arXiv:1712.05614 (submitted to JINST) (2018)

## Slide 19
Fast neutron sensitivity (1 – 10MeV)
100 keV threshold
Thermal-cold neutron efficiency (2.5Å) ∼ 0.44
Fast neutron sensitivity ∼ 10-5
Gamma sensitivity < 10-7
Multi-Blade measurements
Helium-3 measurements
100 keV threshold
0.65 ~ Thermal-cold neutron efficiency (2.5Å)
∼ 10-3 Fast neutron sensitivity
< 10-7 Gamma sensitivity
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers. JINST 13 P03004 (2018).
G. Mauri et al., Evidence of fast neutron sensitivity for 3HE detectors and comparison with Boron-10 based neutron detectors. Submitted EPJ.
Thermal n
Fast n
Gammas
Thermal n
Fast n
Gammas
x100

## Slide 20
Efficiency
64% @ 5.1Å
measured
calculated
58% @ 4.2Å
45% @ 2.5Å
25Hz
50Hz
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 (3) P03013 (2017).

## Slide 21
Stability

## Slide 22
Masks
1 mm pitch holes

## Slide 23
Self-background

## Slide 24
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
sample
Background
Background
0.5 2 6Å
Multi-Blade
detector
(Log scale)
October 2017 - CRISP reflectometer @ ISIS

## Slide 25
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Background
Background
Gating between 6ms and 20ms
0.5 2 6Å
(Log scale)
(Log scale)
sample
Multi-Blade
detector

## Slide 26
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Background
Background
Gating between 6ms and 20ms
0.5 2 6Å
(Log scale)
(Log scale)
sample
Multi-Blade
detector
Projection over Y-axis

## Slide 27
2Å
90%
98%
61%
45%
Nominal 10B4C coating 7um
Actual 10B4C coating 4um
0.5Å
100 neutrons
strips Y
wires X
substrate
10B4C
55 neutrons get scattered @ 0.5Å
sample
Multi-Blade
detector
~4⋅10-3

## Slide 28
neutrons
detector
MB detector
Selene guide
sample
flight tube
November 2018 - AMOR reflectometer @ PSI
WP5 Data Acquisition software chain has been developed during BrightnESS (WP5 - i.e. DMSC/Data) and tested @ AMOR

## Slide 29
neutrons
detector
MB detector
Selene guide
sample
flight tube
November 2018 - AMOR reflectometer @ PSI
flight tube
detector
ESTIA
sample

## Slide 30
flight tube
Multi-Blade
Selene guide
sample
FRAME OVERLAP
MIRROR
SAMPLE
SLIT 4
MID FOCAL POINT
MB DETECTOR
4.9 m
L ~ 3.4 m
SLIT 1
FOCAL POINT
CHOPPER
SELENE 1
SELENE 2

## Slide 31
Beam transmitted
Beam blocked
Cd slit to block the beam far from the detector with shutter open
FRAME OVERLAP
MIRROR
SAMPLE
SLIT 4
MID FOCAL POINT
MB DETECTOR
4.9 m
L ~ 3.4 m
SLIT 1
CHOPPER
SELENE 1
SELENE 2

## Slide 32
Beam transmitted
Beam blocked
Y-Projection

## Slide 33
Beam transmitted
Beam blocked
Y-Projection
1mm Al window
50um Al foil

## Slide 34
Y-Projection
~ 3⋅10-5
~ 10-4
~ 2⋅10-5
1mm window
Foil window
Instrument background

## Slide 35
Y-Projection
CRISP (ISIS) demonstrator
4.4um 10B4C
AMOR (PSI) demonstrator
8um 10B4C
100 neutrons
strips Y
wires X
substrate
10B4C
55 neutrons get scattered @ 0.5Å
~4⋅10-3
~ 3⋅10-5
~ 2⋅10-5
Problem Solved!

## Slide 36
Y-Projection
~ 3⋅10-5
~ 2⋅10-5
Foil window
Instrument background
~ 10-3
He-3 detector
Multi-Blade
~ 3⋅10-2
ESTIA Instrument Proposal
x100 better than He-3

## Slide 37
Uniformity & Overlap

## Slide 38
Edge
gain not corrected
Garfield Simulation
Measured
Edge
wires
Uniformity and overlap
gain corrected

## Slide 39
Uniformity
Wires projection uniformity ~10%
2D uniformity

## Slide 40
Technical characterization on CRISP
Gas gain variation across wires
Uniformity plot, projection on wires, summed over strips
PHS per wire (1 cassette)
Li alpha
Common Th (black)
±10%
Staircase Th (blue)
±7%
Channel by channel Th (red)
±1.5%
Mechanical defects
Mechanical defects
Well characterized
efficiency drop
at blades overlap
Uniformity

## Slide 41
Technical characterization on CRISP
Gas gain variation across wires
Uniformity plot, projection on wires, summed over strips
PHS per wire (1 cassette)
Li alpha
Common Th (black)
±10%
Staircase Th (blue)
±7%
Channel by channel Th (red)
±1.5%
Mechanical defects
Mechanical defects
Well characterized
efficiency drop
at blades overlap
Uniformity

## Slide 43
40mm
60mm
sample
Masks
BN mask
Multi-Blade
detector

## Slide 44
40mm
60mm
sample
BN mask
Multi-Blade
detector
(Log scale)
Masks

## Slide 45
40mm
60mm
sample
BN mask
shadow
Multi-Blade
detector
(Log scale)
Masks
There is a shadow

## Slide 46
strips
wires
40mm
60mm
shadow
sample
BN mask
10B4C
substrate
shadow
Multi-Blade
detector
(Log scale)
Masks
There is a shadow

## Slide 47
strips
wires
40mm
60mm
shadow
sample
Results
BN mask
Not a dead area but the overlap avoids dead zones
10B4C
substrate
shadow
shadow
No shadow
Multi-Blade
detector
(Log scale)
There is a shadow

## Slide 48
40mm
60mm
sample
BN mask
shadow
Multi-Blade
detector
(Log scale)
Masks
There is a shadow
The shadow is changing width

## Slide 49
40mm
60mm
sample
BN mask
shadow
Multi-Blade
detector
(Log scale)
Multi-Blade
sample
neutrons
4m radius
We measured at 2.3m!
Masks
There is a shadow
The shadow is changing width

## Slide 50
Masks
(Log scale)
Shadows removed

## Slide 51
Masks
(Log scale)
Shadows removed
Reduced sensitivity

## Slide 52
Masks
100x100mm2 BN tiles
0.5 mm pitch / 0.5 mm slits
1 mm pitch / 0.5 mm slits
0.5 mm pitch / 0.5 mm slits
3.5mm
0.5mm pitch
1mm pitch

## Slide 53
Rate

## Slide 54
350um
8mm
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
3.5kHz / 0.35x4 mm2
60kHz / 4x8 mm2
Multi-Blade

## Slide 55
350um
8mm
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
3.5kHz / 0.35x4 mm2
60kHz / 4x8 mm2
Multi-Blade

## Slide 56
350um
8mm
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
3.5kHz / 0.35x4 mm2
60kHz / 4x8 mm2
Multi-Blade
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
D17 @ ILL - He3-based detector
*Courtesy of Philipp Gutfreund

## Slide 57
350um
8mm
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
D17 @ ILL - He3-based detector
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
0.3kHz / 0.3x8 mm2
3.5kHz / 0.35x4 mm2
*Courtesy of Philipp Gutfreund
3kHz / 3x8 mm2
D17 @ ILL
60kHz / 4x8 mm2
Multi-Blade

## Slide 58
350um
8mm
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
D17 @ ILL - He3-based detector
November 2018 - AMOR reflectometer @ PSI
Counting Rate Capability
Selene guide
detector
sample
position
Peak Rates
0.3kHz / 0.3x8 mm2
3.5kHz / 0.35x4 mm2
*Courtesy of Philipp Gutfreund
3kHz / 3x8 mm2
D17 @ ILL
60kHz / 4x8 mm2
Multi-Blade
>20x

## Slide 59
November 2018
- AMOR reflectometer @ PSI
Counting Rate Capability

## Slide 60
Rate
For the Local rate the area is ~40 mm2
(10w X 3s) (10 ⋅ 0.35 mm X 3 ⋅ 4 mm).
For the Global rate the area is ~8000 mm2
(180w X 32s) (180 ⋅ 0.35 mm X 32 ⋅ 4 mm).
The difference between the expected rate (lines) and the measured one (dots) is due both to the clustering and the data acquisition.
(the clustering gives a factor 4-5 less with respect to the raw data this is the first drop we see, after that is JADAQ)
p.s. the green line is dashed just to see also the red line below it.

## Slide 61
Science

## Slide 62
MB detector
neutrons
October 2017 - CRISP reflectometer @ ISIS
MB detector
MB detector
sample
neutrons
neutrons

## Slide 63
q
q
DIVERGENT
conventional
l - q encoding
q
q
high intensity
Scientific results from CRISP: Iridium and Si
Multi-Blade
Multi-Blade
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
sample
sample
COLLIMATED

## Slide 64
Si sample
q
q
DIVERGENT
Si sample
q
q
COLLIMATED
The spatial resolution and the counting rate capability of the detector is needed to measure in divergent mode
DIVERGENT
COLLIMATED
LOWER RATE
HIGHER RATE
Multi-Blade
Multi-Blade
120 min measurement
14 min measurement
Scientific results from CRISP: Silicon
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).

## Slide 65
Scan in angle 0.2-0.8 deg
0.5 6
(Y,T)
(X,Y)
Thickness | 550 Å
Roughness | 15 Å
Nb | 7.3⋅10-6 Å-2
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q=0.3o
(Y,T)
(X,Y)
q
q
Multi-Blade
Ir sample

## Slide 66
Conventional
Conventional
0.5 6
0.5 2 6Å
q
Dependence of the q-resolution on detector spatial resolution
Conventional data reduction
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q
q
DIVERGENT
Multi-Blade
Ir sample

## Slide 67
Conventional
Incl. Res.
Incl. Res.
Conventional
0.5 6
0.5 2 6Å
Constant q
q + Dq
Same l
q - Dq
Integration at constant q,
correcting q with the position
on the detector
q
q
Conventional data reduction
q-correction
Including detector
resolution
Dependence of the q-resolution on detector spatial resolution
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q
q
DIVERGENT
Multi-Blade
Ir sample
ESTIA
High Intensity Mode

## Slide 68
0.5mm resolution MB detector
2mm resolution detector
conventional 0D detector
0.2 degrees
0.4 degrees
0.8 degrees
Dependence of the q-resolution on detector spatial resolution
Scientific results from CRISP: Iridium
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
q
q
DIVERGENT
Multi-Blade
sample
ESTIA
High Intensity Mode

## Slide 69
Fe/Si sample
q
q
Multi-Blade
Parallel component of the scattering vector q.
Investigation of in-plane structures.
Specular
Specular
Off-Specular
Off-Specular
(Y,T)
(X,Y)
Specular
Off-Specular
Scientific results from CRISP: Fe/Si supermirror
Off-Specular
Specular
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).

## Slide 70
Off-specular scattering from Fe/Si neutron supermirror
Yoneda wings
Yoneda wings
Sample Horizon
Supermirror edge
(m=3.8)
Supermirror
edge
(from SuperADAM at ILL)
Scientific results from CRISP: Fe/Si supermirror
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
Si edge
Correlated
roughness
domains
Spin-flip scattering
signal

## Slide 71
August 2015
November 2018 - AMOR reflectometer @ PSI

## Slide 72
Selene guide using the ReFocus principle
sample
detector
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
ESTIA Instrument Proposal

## Slide 73
NiTi sample
q
q
DIVERGENT
Multi-Blade
Selene guide using the ReFocus principle
sample
detector
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
Divergence of the beam : 1.6 degrees
ESTIA Instrument Proposal

## Slide 74
NiTi sample
q
q
DIVERGENT
Multi-Blade
Selene guide using the ReFocus principle
sample
detector
Angle 1: 1.4deg
Angle 2: 2.2 deg
Angle 3: 3.2deg
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
Divergence of the beam : 1.6 degrees
2D of the detector
ToF vs vertical coordinate
(x,y)
(ToF,y)

## Slide 75
NiTi sample
q
q
DIVERGENT
Multi-Blade
Selene guide using the ReFocus principle
sample
detector
Angle 1: 1.4deg
Angle 2: 2.2 deg
Angle 3: 3.2deg
Combination of 3 angles
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
Divergence of the beam : 1.6 degrees
(ToF,y)
(q,l)
(q,l)
(q,l)
(q,l)

## Slide 76
Combination of 3 angles
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
(q,l)
(q,l)
ESTIA Instrument Proposal
1
1
2
3
4
5
5
4
3
2
Measured
Simulated

## Slide 77
August 2015

## Slide 78
August 2015
1D
2D
0.55mm x 3.5mm resolution
640 + 1280 ch = ~2k ch
0.3mm res. / probably 500ch
200mm
100mm
Full 2D detector is
1k + 2k ch = ~3k ch.
2D
300mm
0.55mm x 3.5mm resolution
0.3mm res.
probably 1500ch
probably 2500ch
1D

## Slide 79
August 2015

## Slide 80
August 2015
 | 1D | 2D
Background rejection | More sensitive to background | Coincidence, better background rejection
Electronics | ~1k ch (30x30cm2) with same number of el. ch. 3x detector area (90x30cm2) | 1k + 2k = ~3k ch (~30x30cm2)
Resolution | 0.55mm 0.3mm (devel.) (loses compatibility with ESTIA spare parts) | 0.55mm x 3mm 0.3mm (devel. >1.5y) (loses compatibility with ESTIA spare parts)
Gaps (Unif) | 50% loss in 1 resolution No loss (devel.) | 50% loss in 1 resolution No loss (devel.)
1D and 2D parts (cassettes) can be mixed in the same detector
Developments were targeted to requirements set by STAP and the instrument in 2014
No further time/resources exist for further development now: they need to be found externally
Interchangeability with ESTIA is very attractive. Affects cost. Maintainability more complicated.

## Slide 81
August 2015

## Slide 82
efficiency | 44% @ 2.5 Å 57% @ 4. Å 82% @ 12 Å
spatial resolution | 0.5 x 3.5 mm2
uniformity | ±1.5%
stability | 2% (over days)
counting rate capability (peak rate) | >3.5 kHz/mm2 (lower limit) >60kHz / 30mm2 (lower limit)
gamma-ray sensitivity | < 10-7 (with 100keV threshold)
fast neutron sensitivity | < 10-5(with 100keV threshold)
gas gain | 60
overlap | 50% eff. drop in 0.5mm gap
Matching ESS requirements
x3 better than state-of-the-art
x20 better than state-of-the-art
G. Mauri et al., Results of the tests at AMOR (in preparation).
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS, JINST 13 P05009 (2018).
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector, Proc. R. Soc. A 474: 20180266 (2018).
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers. JINST 13 P03004 (2018).
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 P03013 (2017).
F. Piscitelli et al. Study of a high spatial resolution 10B-based thermal neutron detector for neutron reflectometry: the Multi-Blade prototype, JINST 9 P03007 (2014).
Multi-Blade characterization
x100 better than state-of-the-art
as good as state-of-the-art
~

## Slide 83
G. Mauri et al., Results of the tests at AMOR (in preparation).
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS,
ArXiv:1803.09589, JINST 13 P05009 (2018).
G. Mauri et al., Neutron reflectometry with the Multi-Blade 10B-based detector,
ArXiv:1804.03962, Proc. R. Soc A 474: 20180266 (2018).
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on boron-10 converter layers,
ArXiv:1712.05614, JINST 13 P03004 (2018).
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, ArXiv:1701.07623, JINST 12 P03013 (2017).
F. Piscitelli et al., Neutron reflectometry on highly absorbing films and its application to 10B4C-based neutron detectors,
ArXiv:1510.01085v1, Proc. R. Soc. A 472, Issue 2185 (2016).
F. Piscitelli et al. Study of a high spatial resolution 10B-based thermal neutron detector for neutron reflectometry:
the Multi-Blade prototype, ArXiv1312.2473, JINST 9 P03007 (2014).
F. Piscitelli and P. Van Esch, Analytical modeling of thin film neutron converters and its application to thermal
neutron gas detectors, ArXiv:1302.3153, JINST 8 P04020 (2013).
References

## Slide 84
Thank you.

## Slide 85
August 2015
Backup slides

## Slide 86
Uniformity
Cassette n-1
Cassette n
Garfield Simulation
wires

## Slide 87
Technical characterization on CRISP
measured
calculated
Efficiency
Measured PHS
across wires
0.14 degrees
7Li
alpha
Strip signals
Wire signals
Cluster Multiplicity
Multiplicity 2
Cassette shadowing
due to overlap
(not a dead area)
Shadows removed
BN mask
Gain compensation
by HV in final detector
Garfield simulation
gas gain

## Slide 88
Neutron reflectometry measurements, scientific measurements (CRISP @ ISIS)
detector
sample
l1
l2
l3
q
q
Ir sample
Fe/Si sample
Si sample
0.5mm resolution MB detector
2mm resolution detector
conventional 0D detector
Ir sample: Dependance of the q-resolution on detector spatial resolution
Si sample: Collimated vs divergent mode
DIVERGENT
COLLIMATED
COLLIMATED: 120 min measurement
DIVERGENT: 14 min measurement
Fe/Si sample: Off-specular
(from SuperADAM at ILL)
Sample Horizon
B. Correlated roughness
domains from the layers
A. Spin-flip scattering
signal from the layers
G. Mauri et al., “Neutron reflectometry with the multi-blade 10b-based detector,” PRSA: Mathematical, Physical and Engineering Sciences, vol. 474, no. 2216, 2018.

## Slide 89
strips
wires
shadow
sample
Not a dead area but the overlap avoids dead zones
10B4C
substrate
shadow
No shadow
dead area
No shadow = No overlap
Shadow = Overlap
There is a shadow
Multi-Blade
Technical characterization on CRISP
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer, ArXiv 1803.09589, accepted JINST (2018).

## Slide 90
August 2015

## Slide 91
Masks
1 mm pitch holes

## Slide 92
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
Neutron reflectometry measurements on CRISP
(X,Y,T)
(Y,T)
(X,Y)

## Slide 93
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
detector area
detector ToF
Neutron reflectometry measurements on CRISP
(X,Y,T)
(Y,T)
(X,Y)

## Slide 94
Neutron reflectometry measurements on CRISP
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
detector area
detector ToF
(X,Y,T)
(Y,T)
(X,Y)
(q,l)

## Slide 95
detector
sample
l1
l2
l3
q
q
ToF (Time-Of-Flight) reflectometer
(Y,T)
(X,Y)
detector area
detector ToF
Momentum transfer
(X,Y,T)
(q)
(Y,T)
(X,Y)
(q,l)
Neutron reflectometry measurements on CRISP

## Slide 96
August 2015
D17 @ ILL
sample
q
q
Conventional Mode
detector
sample
High-intensity mode
detector
Selene guide using the ReFocus principle
High detector requirements
1 10 100
1 = state-of-the-art
Instantaneous
Rate Capability
several kHz/mm2
1 2 3 4
Reflectometry
Spatial Resolution
Reflectometry
0.5 mm
He3-based detector
64 multi-tube
sample
detector
STATE-OF-THE-ART
10kHz/tube
(250x8 mm2)
…
0.3kHz / 0.3x8 mm2
3kHz / 3x8 mm2
2 x 8 mm2 spatial resolution
*Courtesy of Philipp Gutfreund

## Slide 97
August 2015

## Slide 99
Backup slides

## Slide 100
Backup slides

## Slide 102
Neutron beam
DAQ VME crate with CAEN digit
On CRISP the MB will be rotated
by 90deg, blades horizontally placed
Backup slides

## Slide 103
Backup slides

## Slide 104
MB on AMOR

## Slide 105
flight tube
detector
ESTIA
detector
ESTIA
sample
sample
detector
detector
flight tube
flight tube
sample
Selene guide