## Slide 1: ODIN: Optical and Diffraction Imaging with Neutrons at the ESS Technical and design update
PSI: M. Morgano, M. Strobl
TUM: E. Calzada, A. Tartaglione, M. Schulz
Notes: Phase 2, May 31st. 2017

## Slide 2: Outline
ODIN Technical Design update in Bunker
  WFMC
  T0 Chopper
  Heavy Shutter
ODIN Technical Design update outside the Bunker
  Cave – MCNP Results - Technical design
  Control Hutch
  Neutron guide Shielding interface

## Slide 3: Instrument design in Bunker
Bunker Wall
Monolith
WFMC
FOC1 – BP1
FOC3
FOC2 – BP2
FOC4
FOC5
HS
T0
In Bunker
Outside theBunker
11m Zone

## Slide 4
Remote Handling Unit 1
WFMC1
WFMC2

## Slide 12
Holders for Laser tracker reflectors
Captive Bolt
Captive Bolt
Kinematic Mounting
Kinematic Mounting
Redundant limit switches

## Slide 13
Rack and Pinion
Displacement measuring system
B4C
Drive Unit

## Slide 14
FOC2 – BP2

## Slide 15
FOC2 – BP2
T0 Chopper
FOC3
Heavy Shutter

## Slide 16
Leveling Wedges
T0 Chopper

## Slide 17
Simulated Attenuation Materials:
10B4C: 3mm
Tungsten: 100mm
Copper : 30mm
Nickel: 30mm
Tungsten: 100mm
Total thickness: 263 mm
Frequency: 28Hz
Attenuation<1E-4 for 1keV - 1MeV Neutrons

## Slide 18
Neutron Guide
Shutter Block
Heavy Shutter requirements:
The basic requirement is to have 3µSv/h inside the cave and 25µSv/h in the Chopper Pit for the FOC5
Design life >15yrs
Service interval >5 years
Reliability >98% (measured over 100 days of beam distribution)
Construction materials shall be selected to achieve activation to levels compatible with access requirements (in the bunker), and instrument maintenance strategies. (i.e. levels comparable or lower than the activation of the current bunker rear wall).
Material selection and construction compatible with facility requirements for decommissioning and disposal of activated equipment (i.e. level of activation, size of components compatible with disposal casks).
Designed for 'hands on' serviceable within 30 days of beam off (full power) if service interval <15years.
Serviceability in-situ strongly preferred.
Installed mass <17t.
Closing time after close command should be less than 2 minutes.
Heavy Shutter

## Slide 19
Pneumatic Actuator
Shock Absorbers
Non-lubricated plain bearings
Neutron Guide
Shutter Block
Heavy Shutter

## Slide 20: Instrument design outside the Bunker
Cave:
Radiation Requirements:
Walls 3µSv/h. Roof 25 µSv/h
Safety factor by ESS: 2
Requires surface detectors of all walls and roof
Operational Scenarios: Steel sample and water sample in experimental cave. Gd slice (2mm) (grating interferometer)
Worst case scenarios: Cd (2mm) in optical or experimental cave

## Slide 21
Cave: MCNP Results: Version “L”
18m
2.4m
13.6m
7m
8.2m
3m

## Slide 22
Cave: MCNP Results
2.137m
N- Beam

## Slide 23
Cave: MCNP Results

## Slide 24
Cave: MCNP Results

## Slide 25
Cave: MCNP Results

## Slide 26
Cave: MCNP Results

## Slide 27
Cave: MCNP Results

## Slide 28
Cave: MCNP Results

## Slide 29
Cave: MCNP Results

## Slide 30
Cave: MCNP Results

## Slide 31
Cave: MCNP Results

## Slide 32
Cave: MCNP Results

## Slide 33
Cave: MCNP Results

## Slide 34
Cave: MCNP Results

## Slide 35
Cave: MCNP Results

## Slide 36
Cave: MCNP Results

## Slide 37
Cave - MCNP Conclusion:
With version “L” of the ODIN cave the desired dose rate levels are not exceeded.

## Slide 38: Instrument design outside the Bunker
Cave: Technical Design

## Slide 39
Cave: Technical Design
30cm
Design of a possible option on the bases to transmit the load from the walls and roof to the floor keeping the load on the floor within the required limits
1.5m
63.7cm
30cm
Floor load <14t/m²

## Slide 40
Cave: Technical Design
18m
2.4m
13.6m
7m
3m
8.2m
Optical Cave
Experimental Cave
Internal Door
Sliding Door

## Slide 41
Cave: Technical Design

## Slide 42
Cave: Technical Design
Beam Stopper
False Floor
45cm Concrete
17cm Steel
63cm Concrete
1.5m

## Slide 43
Cave: Technical Design
Interface Guide Shielding – Cave
Horizontal Cut
Interface Guide Shielding – Cave
Vertical Cut

## Slide 44
Cave: Technical Design

## Slide 45
Cave: Technical Design
Sliding Door:
Round track system
Internal Door:
Air cushion modules

## Slide 46
Cave: Technical Design
Removable plates

## Slide 47
Cave: Technical Design
External Chicanes

## Slide 48
Cave: Technical Design
Removable plates
Concrete block along the beam
Outputs from the supply channel
below the false floor
Feed through
Self-supporting
removable roof Slabs

## Slide 49
Cave: Technical Design
Neutron guide Shielding
Cave

## Slide 50
Cave: Technical Design

## Slide 51
Cave: Technical Design

## Slide 52
Cave: Technical Design

## Slide 53
Cave: Technical Design

## Slide 54
Cave: Technical Design

## Slide 55
Cave: Technical Design

## Slide 56
Cave: Technical Design
Control Hutch
Platform
Lift
Sample preparation Area
Ramp
Racks:
Choppers and Vacuum
Racks:
Motion Control

## Slide 57
Cave: Technical Design
Racks:
Motion Control
Racks:
Data Management
PSS
Racks:
Choppers and Vacuum

## Slide 58: Thank You!