## Slide 1: Updates on the beam monitor common project
Kalliopi Kanaki
Steven Alcock, Ioannis Apostolidis, Vendula Maulerová
IKON18, 25 February 2020

## Slide 2: Overview of beam monitor session
Status of common project (Kalliopi Kanaki, ESS)
Status of V-monitor (Vendula Maulerova, ESS)
Mechanical interfaces and integration (Ioannis Apostolidis, ESS)
DMSC interfaces (Jonas Nilsson, ESS)
Beam monitors at ILL (Fabien Lafont, ILL/ESS)
NitroGEM beam monitor updates (Davide Raspino, ISIS)

## Slide 3: Purpose of the project
Ensure provision of monitors to enable science case of instruments and in particular early science success
Allow efficient commissioning and long term facility monitoring
Design, procurement, installation and cold commissioning of ~55 beam monitors for the ESS instruments from a few kW to 2 MW
People involved:
  Steven Alcock – electronics engineer (15%-20%)
  Ioannis Apostolidis – mechanical engineer (50%)
  Kalliopi Kanaki – neutron scientist (50%)
  Vendula Maulerová – workshop assistant (50%)
Related documents:
  https://confluence.esss.lu.se/display/DG/Beam+monitors+common+project
  Beam Monitors Common Project description (ESS-0502620)
  Neutron Beam Monitor Use Cases for commissioning and operation of neutron instruments and processing of neutron data (ESS-0419542)
Anders left in Dec 2019.

## Slide 4: Participation status
# | Instrument / BM partner | Spatial envelope status | Status of McStas/Vitess model | Indicative participation status
1 | ESTIA / PSI | Reservations in CAD | Up-to-date | Interested in utilities and services (V-monitor)
2 | ODIN / PSI | Reservations in CAD | Up-to-date | YES
3 | DREAM / FZJ | Reservations in CAD | Up-to-date | NO (CDT monitors)
4 | BEER / HZG | Reservations in CAD | Up-to-date | YES (in-bunker BMs), MWPC by HZG (pre-sample)
5 | CSPEC / TUM | Reservations in CAD | Up-to-date | Under consideration
6 | BIFROST / LLB | Reservations in CAD | Up-to-date | YES
7 | MAGIC / FZJ | Reservations in CAD | Update on the way | Undecided at the moment
8 | LoKI / ISIS | Reservations in CAD | Up-to-date | NO (NitroGEM monitors from ISIS)
9 | SKADI / JCNS | Reservations in CAD | Up-to-date | YES
10 | VESPA / CNR, ISIS | Reservations in CAD | Up-to-date | NO (B-GEM monitors from University of Milano-Bicocca)
11 | NMX / ESS | Reservations in CAD | Up-to-date | YES
12 | MIRACLES / ESS Bilbao | Reservations in CAD | Up-to-date | YES
13 | T-REX / CNR | Contact initiated | 2018-09-17 | Discussion ongoing
14 | HEIMDAL / AU | Reservations in CAD | Up-to-date | Undecided at the moment
15 | FREIA / ISIS | Reservations in CAD | Up-to-date | YES?

## Slide 5: Beam monitors per location and function
Location | BM | Desired functionality
Bunker | Fission chamber (flux ✕ eff ≤ 1012 n/cm2/s) | •Flux •Timing •Discrimination of thermal n vs. fast n vs. γ
 | Ionisation chamber (flux ✕ eff ≤ 1012 n/cm2/s) | 
 | GEM (flux ✕ eff ≤ 108 n/cm2/s) | 
Guides/choppers | Quasi-parasitic V-monitor (109 n/cm2/s) | •Flux •Timing (to be investigated further for V-monitor)
 | GEM (flux ✕ eff ≤ 108 n/cm2/s) | 
Pre- and post-sample | MWPC (flux ✕ eff ≤ 105 n/cm2/s) GEM (flux ✕ eff ≤ 108 n/cm2/s) | •Flux •Timing

## Slide 6: Meeting the flux challenge
Optimisation parameters for standard BMs
  Choice of converter (10B4C, 235UO2, N2, 3He)
  Solid converter thickness
  Enrichment
  Gaseous converter partial pressure
  Active gas depth
Optimisation parameters for V-monitor
  Foil size
  Foil thickness
  Active detector area
  Detector distance from foil
Green: neutrons
Gray: BM windows
Yellow: gas
Blue: collector volume for analysis

## Slide 7: B4C conversion efficiency
10B4C (99% enriched)
10B4C (99% enriched)
natB4C
natB4C
Efficiency is primarily defined by:
Enrichment
Coating thickness

## Slide 8: UO2 conversion efficiency
Solid line: 235UO2 (93% enriched)
Dotted line: ± 30% effective thickness variation

## Slide 9: UO2 conversion efficiency
Solid line: 235UO2 (93% enriched)
Dotted line: ± 30% effective thickness variation
Dashed line: natUO2

## Slide 10: 3He conversion efficiency
BM (gas) thickness = 1.3 cm

## Slide 11: 3He conversion efficiency
thickness = 1.3 cm
BM (gas) thickness = 1.3 cm
Dotted line: 1.0 cm < thickness < 1.6 cm

## Slide 12: N2 conversion efficiency
BM (gas) thickness = 1.3 cm

## Slide 13: N2 conversion efficiency
BM (gas) thickness = 1.3 cm
Dotted line: 1.0 cm < thickness < 1.6 cm

## Slide 14: BM efficiency calibration
Draft from 2011!
Honeycomb (hexacomb) detector
Allows absolute flux determination of ~ 1%!
Used for MultiGrid and MultiBlade efficiency characterisation
30 cm

## Slide 15: Impact of multiple BM windows
BM placed at 10 cm distance
from each other
NO converters or gas
Only windows and/or GEM foil
Monochromatic beam
neutrons

## Slide 16: Impact of multiple BM windows
δθ > 0.1⚬
Need a definition for scattering that serves instrument purposes.
transmission
scattering

## Slide 17: Transmission of single BM monitor
2% loss at 4 Å per 2 mm of Al
Transmission > 99% at 10 Å for thin foil BM windows (100 μm) in air placement (1 atm inner pressure)

## Slide 18: Transmission studies with FREIA model
Transmission when all windows – guide (1 mm) & BM (1 mm) – are enabled.
Windows OFF
Windows ON
PRELIMINARY!!!

## Slide 19: Transmission studies with FREIA model
Transmission when all windows – guide (1 mm) & BM (100 μm) – are enabled.
Windows OFF
Windows ON
PRELIMINARY!!!

## Slide 20: Setup with official ESS DAQ readout
BM
Preamp
Shaping amplifier
Back-end electronics
ADC+FPGA
Timing from ICS
EFU
Event Formation Unit
Kafka
aggregator
FileWriter
(hdf5)
Event-mode data
Timestamping
Pulse height spectra
Grafana live visualisation
Possibility for flux and time resolution analysis
V. Maulerová et al., EPL 128 (2020) 52001

## Slide 21: Setup with ESS DAQ readout
BM
Preamp
Shaping amplifier
Back-end electronics
ADC+FPGA
Timing from ICS
EFU
Event Formation Unit
Kafka
aggregator
FileWriter
(hdf5)
V20 setup
Event-mode data
Timestamping
Pulse height spectra
Grafana live visualisation
Possibility for flux and time resolution analysis
V. Maulerová et al., EPL 128 (2020) 52001

## Slide 22: BM data acquisition in Utgård (DG, ICS, ECDC)
8 channels in 2 ADC boxes
EPICS control for HV power supply and back-end readout
Event Formation Unit (EFU)/Kafka
Grafana for EFU/BM monitoring
File writer
Software ADC scope for pulse monitoring
Tests with cosmics on-going

## Slide 23: Next steps
Close out phase 1 (after IKON18)
Write Critical Design Report (March-April)
  Distribute monitor technologies in each reserved location
  Document summary of activities so far
Complete detailed engineering design
  Initial focus on: ODIN, BIFROST, NMX, ESTIA
Write offer documents for common project participants
Write publications
Experimental tests planned (Dubna, J-PARC)

## Slide 24: Check-list for the non-participants of common project
Post TG5 the BMs are delivered to ESS for operation and maintenance.
Use cases satisfaction
Documentation of BM requirements and specs (peak and average flux, timing where relevant, stability, time resolution, attenuation, scattering etc.)
Describe method for absolute flux derivation
Describe BM calibration methods before installation and regularly during operations (efficiency and TOF)
Maintain up-to-date McStas model wrt to engineering model
Shielding and insulation (radiation, electrical, mechanical)
Mechanical & electrical integration (grounding, cable routes, in-vacuum specifics, remote handling modules, etc.)
Clear definition of interfaces between instrument and DG/ICS/DMSC
Front-end electronics
Back-end electronics readout, ICS and DMSC integration (time-stamping and Mantid integration)
EPICS controls, adoption of ESS operation modes, logging parameters
Rack scheme, server characteristics
Software installation
Cost estimate (hardware, labour)
Schedule match
Spares strategy and 10-year maintenance plan

## Slide 25: Backup slides

## Slide 26: TOF and λ calibration
4.03 Å
2.8 Å
2.3 Å
Steel cross sections and scattering angles (α-Fe)

## Slide 27: TOF and λ calibration