## Slide 1
DETECTOR BACKGROUNDS
Fast Neutron and Gamma-ray Sensitivity
of Helium-3 and Boron-10 Detectors
and Effect of Window and Cosmic Neutrons
Francesco Piscitelli
on behalf of the ESS Detector Group
IKON18
Lund
2020/02/25

## Slide 2
He-3 vs B-10
DETECTOR BACKGROUND
Cosmic neutrons
Gamma-rays
Fast Neutrons
Scattered Neutrons

## Slide 3
He-3 vs B-10
Detector x-section
Detector window
Detector
Detector sensitive cells or voxels
Cosmic neutrons
Gamma-rays
Fast Neutrons
Scattered Neutrons
DETECTOR BACKGROUND

## Slide 4
He-3 vs B-10
Detector x-section
Detector window
Detector sensitive cells or voxels
Good events
Unwanted events
LEGEND
Cosmic neutrons
Gamma-rays
Fast Neutrons
Scattered Neutrons
DETECTOR BACKGROUND
Detector

## Slide 5
August 2015
COSMIC NEUTRON
BACKGROUND

## Slide 6
BACKGROUND: cosmic neutrons
Neutrons are created by cosmic ray spallation in the high atmosphere.
Energies from 10-9 to 103 MeV

## Slide 7
BACKGROUND: cosmic neutrons
Neutrons are created by cosmic ray spallation in the high atmosphere.
Energies from 10-9 to 103 MeV
Physikalisch-Technische Bundesanstalt (PTB) - Measurements of Neutron Spectra Induced by Cosmic Radiation at Altitudes of 85m, 1195m and 2650m (2010)
85m
1195m
2650m
Neutron flux
Measured with Bonner spheres in the whole eneregy spectrum
2650m

## Slide 8
BACKGROUND: cosmic neutrons
Neutrons are created by cosmic ray spallation in the high atmosphere.
Energies from 10-9 to 103 MeV
Physikalisch-Technische Bundesanstalt (PTB) - Measurements of Neutron Spectra Induced by Cosmic Radiation at Altitudes of 85m, 1195m and 2650m (2010)
85m
1195m
2650m
Measured with Bonner spheres in the whole eneregy spectrum
2650m
Neutron flux
FAST
25meV
1-100MeV
elevation
THERMAL

## Slide 9
BACKGROUND: cosmic neutrons
Neutrons are created by cosmic ray spallation in the high atmosphere.
Energies from 10-9 to 103 MeV
Physikalisch-Technische Bundesanstalt (PTB) - Measurements of Neutron Spectra Induced by Cosmic Radiation at Altitudes of 85m, 1195m and 2650m (2010)
85m
1195m
2650m
Measured with Bonner spheres in the whole eneregy spectrum
2650m
Neutron flux
FAST
elevation 85m ~ ESS
135 Hz/m2 (10-9 to 103 MeV) by PTB
25meV
1-100MeV
elevation
THERMAL

## Slide 10
BACKGROUND: cosmic neutrons
Neutrons are created by cosmic ray spallation in the high atmosphere.
Energies from 10-9 to 103 MeV
Physikalisch-Technische Bundesanstalt (PTB) - Measurements of Neutron Spectra Induced by Cosmic Radiation at Altitudes of 85m, 1195m and 2650m (2010)
85m
1195m
2650m
Measured with Bonner spheres in the whole eneregy spectrum
2650m
Neutron flux
THERMAL
FAST
100% efficiency
10-3 to 10-5 efficiency
25meV
1-100MeV
elevation 85m ~ ESS
135 Hz/m2 (10-9 to 103 MeV) by PTB

## Slide 11
BACKGROUND: cosmic neutrons
cosmics
Detector window
gas
Solids + Gases
Measurements done in Utgård
He-3 and He-4 detectors
He-4 10bar
He-3 10bar
He-4 5bar
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
*Thanks to Toshiba/Canon Electron Tubes & Devices Co. LTD for the He-4 tubes.
Dimensions:
2.54mm diam. x 250mm length

## Slide 12
BACKGROUND: cosmic neutrons
Bare tubes
He-3 10bar
He-4 10bar
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
He-4 5 bar
He-3 10 bar
He-4 10 bar
Pulse-Height-Spectrum

## Slide 13
BACKGROUND: cosmic neutrons
Covered with 2mm Mirrobor
Bare tubes
He-4 5 bar
He-3 10 bar
He-4 10 bar
He-3 10bar
He-3 10bar
He-4 10bar
He-4 10bar
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
Pulse-Height-Spectrum
Pulse-Height-Spectrum

## Slide 14
BACKGROUND: cosmic neutrons
Covered with 2mm Mirrobor
Bare tubes
Thermal n decrease
He-3 10bar
He-3 10bar
He-4 10bar
He-4 10bar
Fast n
unaltered
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
He-4 5 bar
He-3 10 bar
He-4 10 bar
Pulse-Height-Spectrum
Pulse-Height-Spectrum

## Slide 15
BACKGROUND: cosmic neutrons
Covered with 2mm Mirrobor
Bare tubes
Rate Hz | rate per area Hz/m2 | rate per volume Hz/m3 | detector | Rate Hz | rate per area Hz/m2 | rate per volume Hz/m3
0.148 0.0043 | 23 0.7 | 292 8.5 | He-3 10bar He-4 10bar | 0.007 0.005 | 1 0.8 | 14 10
0.029 Hz / (bar ⋅ litre) of He3
0.0014 Hz / (bar ⋅ litre) of He3
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
He-4 5 bar
He-3 10 bar
He-4 10 bar

## Slide 16
BACKGROUND: cosmic neutrons
Covered with 2mm Mirrobor
Bare tubes
Rate Hz | rate per area Hz/m2 | rate per volume Hz/m3 | detector | Rate Hz | rate per area Hz/m2 | rate per volume Hz/m3
0.148 0.0043 | 23 0.7 | 292 8.5 | He-3 10bar He-4 10bar | 0.007 0.005 | 1 0.8 | 14 10
0.029 Hz / (bar ⋅ litre) of He3
0.0014 Hz / (bar ⋅ litre) of He3
elevation 85m ~ ESS
135 Hz/m2 (10-9 to 103 MeV) by PTB
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
He-4 5 bar
He-3 10 bar
He-4 10 bar
Agreement with rates at FRMII
K. Zeitelhack – private communication
Agreement with rates measured with MG at Utgård (See next talk from Alex)

## Slide 17
August 2015
GAMMA-RAY
BACKGROUND

## Slide 18
BACKGROUND: gamma-rays
Detector window
gas
Solids + Gases

## Slide 19
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
e-
gamma-ray
e-
Photoelectric / Compton
Photoelectric interaction most probable at low energy,
Compton scattering is the majority of the background
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
photoelectric
Compton
pair
N
50keV

## Slide 20
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
gas
solid
electrons
e-
gamma-ray
e-
Photoelectric interaction most probable at low energy,
Compton scattering is the majority of the background
photoelectric
Compton
pair
Photoelectric / Compton
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
50keV

## Slide 21
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
gas
solid
electrons
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
e-
e-
Gas vs Solid
gammas in the walls can only result in a signal if the electron reaches the gas
photons interacting with the gas are guaranteed to generate a signal
gas
gas
solid

## Slide 22
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
gas
solid
electrons
Gas dominates
240keV
gamma
150 keV for
1 atm of ArCO2
and Al walls
is 50/50
30keV
gamma
Solid dominates
(Simulations)
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
g–ray energy
Gas vs Solid
30keV
150keV
240keV

## Slide 23
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
gas
solid
electrons
Gas dominates
240keV
gamma
150 keV for
1 atm of ArCO2
and Al walls
is 50/50
30keV
gamma
Solid dominates
(Simulations)
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
g–ray energy
Gas vs Solid
30keV
150keV
240keV
Energy threshold generally 100-200keV
Solid dominates

## Slide 24
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
Higher pressure = higher energy deposition
(Simulations)
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)

## Slide 25
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
Higher pressure = higher energy deposition
(Simulations)
Higher gamma energy = higher energy deposition
… BUT deposited energies
rarely exceed a few tens
of keV regardless of
the primary energy
8MeV incoming gamma
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
60keV incoming gamma

## Slide 26
BACKGROUND: gamma-rays
gamma-rays
Detector window
gas
Solids + Gases
electrons
Measurements
Multi-Grid B-10
(spectrometers)
Multi-Blade B-10
(reflectometers)
He-3/He-4 detectors
Solid dominates no matter if He-3 or B-10

## Slide 27
August 2015
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
Multi-Grid B-10
He-3
gammas <400keV
gammas 1.1 - 1.3MeV
Solid dominates no matter if He-3 or B-10

## Slide 28
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency
Pulse Height Spectrum - PHS
Sensitivity / efficiency

## Slide 29
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency

## Slide 30
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency

## Slide 31
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency

## Slide 32
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency

## Slide 33
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency

## Slide 34
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency

## Slide 35
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 P03013 (2017) (arxiv: 1701.07623)

## Slide 36
August 2015
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 P03013 (2017) (arxiv: 1701.07623)
Measured with Multi-Blade B-10
PHS
Sensitivity / Efficiency
Cold neutrons
gammas

## Slide 37
August 2015
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 P03013 (2017) (arxiv: 1701.07623)
Sensitivity / Efficiency
Cold neutrons
gammas
50 – 200keV
~3 orders of magnitude
No matter if He3 or B-10
(Solid Dominates)
10-6
10-9
Gamma-ray Sensitivity

## Slide 38
August 2015
FAST NEUTRON
BACKGROUND

## Slide 39
BACKGROUND: fast neutrons
Detector window
gas
Solids + Gases

## Slide 40
fast neutrons
Detector window
gas
N
Solids + Gases
gas
solid
Recoil
BACKGROUND: fast neutrons
Absorption
In either
solid or gas

## Slide 41
fast neutrons
Detector window
gas
N
Solids + Gases
gas
solid
Recoil
BACKGROUND: fast neutrons
Absorption
In either
solid or gas
g
n
n’
X’
X
n
Y’
X
n’
n
X
n’
Recoil
(elastic)
(n,p)
(n,alpha)
(n,gamma)
etc.
Absorption
Inelastic
(elastic)
(n,gamma)
gamma
Recoil nucleus
Emitted particle

## Slide 42
fast neutrons
Detector window
gas
N
Solids + Gases
gas
solid
Recoil
BACKGROUND: fast neutrons
Absorption
In either
solid or gas
g
n
n’
X’
X
n
Y’
X
n’
n
X
n’
Recoil
(elastic)
(n,p)
(n,alpha)
(n,gamma)
etc.
Absorption
Inelastic
(elastic)
(n,gamma)
gamma
Sensitivity <10-6
Recoil nucleus
Emitted particle

## Slide 43
fast neutrons
Detector window
gas
N
Solids + Gases
gas
solid
Recoil
BACKGROUND: fast neutrons
Absorption
In either
solid or gas
X
X
particle in the walls can only result in a signal if the particle reaches the gas
particles interacting with the gas are guaranteed to generate a signal
gas
gas
solid
Solid vs Gas

## Slide 44
Detector window
gas
N
Solids + Gases
gas
solid
Recoil
BACKGROUND: fast neutrons
Absorption
In either
solid or gas
Probability to generate an event
Gas dominates
Gas = x100 Solid
True for both He-3 and B-10
Solid vs Gas
Gas
Solids
Probability
fast neutrons

## Slide 45
August 2015
Gas is main component, then He-3 differs from B-10 …
BACKGROUND: fast neutrons
Detector window
gas
B-10
Window+(conversion+detection)
Window+conversion+detection
He-3
He-3 detector
N
N
gas
Recoil on He-3
Recoil on Argon
Detector window
Fast N
Fast N
B-10 detector
Gas dominates

## Slide 46
August 2015
BACKGROUND: fast neutrons
Detector window
gas
B-10
Window+(conversion+detection)
Window+conversion+detection
He-3
He-3 detector
N
N
gas
Recoil on He-3
Recoil on Argon
gas
He-4
gas
Detector window
N
Fast N
Fast N
Fast N
N
Window+detection
Window+detection
B-10 detector
Exchange He-3 with He-4
Remove B-10
Gas is main component, then He-3 differs from B-10 …
Gas dominates

## Slide 47
August 2015
BACKGROUND: fast neutrons
Detector window
gas
B-10
Window+(conversion+detection)
Window+conversion+detection
He-3
He-3 detector
N
N
gas
Recoil on He-3
Recoil on Argon
gas
He-4
gas
Detector window
N
Fast N
Fast N
Fast N
N
Window+detection
Window+detection
B-10 detector
Exchange He-3 with He-4
Remove B-10
Gas is main component, then He-3 differs from B-10 …
Gas dominates

## Slide 48
August 2015
BACKGROUND: fast neutrons in He-3 detectors
Detector window
gas
Window+(conversion+detection)
He-3
He-3 detector
N
Recoil on He-3
gas
He-4
Fast N
Fast N
N
Window+detection
He-4
10bar
He-3
10bar
He-4
5bar
Exchange He-3 with He-4
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
*Thanks to Toshiba/Canon Electron Tubes & Devices Co. LTD for the He-4 tubes.
Dimensions:
2.54mm diam. x 250mm length

## Slide 49
August 2015
BACKGROUND: fast neutrons in He-3 detectors
Exchanging He-3 with He-4 is a good approximation to evaluate the sensitivity of He-3 to fast n
n
n’
X’
X
n
Y’
X
n’
Recoil
Absorption
He-3
n
n’
X’
X
Recoil
He-4
(n,p)
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)

## Slide 50
August 2015
BACKGROUND: fast neutrons in He-3 detectors
Exchanging He-3 with He-4 is a good approximation to evaluate the sensitivity of He-3 to fast n
n
n’
X’
X
n
Y’
X
n’
Recoil
Absorption
He-3
n
n’
X’
X
Recoil
He-4
(n,p)
1
Similar total cross-section:
s(He-3) ~ s(He-4)
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)

## Slide 51
August 2015
BACKGROUND: fast neutrons in He-3 detectors
Exchanging He-3 with He-4 is a good approximation to evaluate the sensitivity of He-3 to fast n
n
n’
X’
X
n
Y’
X
n’
Recoil
Absorption
He-3
n
n’
X’
X
Recoil
He-4
Almost identical energy deposition ion gas:
He-3++ in He-3 ~ He-4++ in He-4
(n,p)
1
2
Similar total cross-section:
s(He-3) ~ s(He-4)
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)

## Slide 52
August 2015
BACKGROUND: fast neutrons in He-3 detectors
Exchanging He-3 with He-4 is a good approximation to evaluate the sensitivity of He-3 to fast n
n
n’
X’
X
n
Y’
X
n’
Recoil
Absorption
He-3
n
n’
X’
X
Recoil
He-4
Almost identical energy deposition ion gas:
He-3++ in He-3 ~ He-4++ in He-4
(n,p)
1
2
3
Similar total cross-section:
s(He-3) ~ s(He-4)
Different PHS
but
~ same number of events !
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)

## Slide 53
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)

## Slide 54
August 2015
Fast-N
gammas
Fast-N
gammas
He-3 (10bar)
He-4 (10bar)
Sensitivity / Efficiency
He-4 (5bar)
He-4 (10bar)
He-3 (10bar)
He-4 (5bar)
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
Pulse Height Spectrum - PHS

## Slide 55
August 2015
He-3 (10bar)
He-4 (10bar)
10-3
Sensitivity / Efficiency
He-4 (5bar)
200keV
Fast-N
gammas
He-4 (10bar)
He-3 (10bar)
He-4 (5bar)
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
Difference between He3 or B-10
(Gas Dominates)
Fast Neutron Sensitivity
He-3

## Slide 56
August 2015
He-3 (10bar)
He-4 (10bar)
200keV
Sensitivity / Efficiency
He-4 (5bar)
100keV
Fast-N
gammas
He-4 (10bar)
He-3 (10bar)
He-4 (5bar)
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
10-3
Difference between He3 or B-10
(Gas Dominates)
Fast Neutron Sensitivity
He-3

## Slide 57
August 2015
He-3 (10bar)
He-4 (10bar)
200keV
Sensitivity / Efficiency
He-4 (5bar)
100keV
Fast-N
gammas
He-4 (10bar)
He-3 (10bar)
He-4 (5bar)
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
10-3
Difference between He3 or B-10
(Gas Dominates)
Fast Neutron Sensitivity
He-3
NOTE:
high pressure He-3 detectors,
little gain in thermal n efficiency with drastic increase of fast n sensitivity!
S/B matters!

## Slide 58
August 2015
BACKGROUND: fast neutrons
Detector window
gas
B-10
Window+(conversion+detection)
Window+conversion+detection
He-3
He-3 detector
N
N
gas
Recoil on He-3
Recoil on Argon
gas
He-4
gas
Detector window
N
Fast N
Fast N
Fast N
N
Window+detection
Window+detection
B-10 detector
Exchange He-3 with He-4
Remove B-10
Gas is main component, then He-3 differs from B-10 …
Gas dominates

## Slide 59
August 2015
BACKGROUND: fast neutrons in B-10 detectors
B-10
Window+conversion+detection
N
gas
Recoil on Argon
gas
Detector window
N
Fast N
Window+detection
B-10 detector
Remove B-10

## Slide 60
August 2015
BACKGROUND: fast neutrons in B-10 detectors
B-10
Window+conversion+detection
N
gas
Recoil on Argon
gas
Detector window
N
Fast N
Window+detection
B-10 detector
Remove B-10
Compared Titanium, Aluminum, Copper, Kapton
Gas dominates
Gas = x100 Solid
Experimentally (+simulations) proved -> no difference!
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on Boron-10 converter layers, JINST 13 P03004 (2018) (arxiv: 1712.05614)

## Slide 61
August 2015
Pulse Height Spectrum - PHS
Pulse Height
counts
Threshold
Sensitivity / efficiency
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on Boron-10 converter layers, JINST 13 P03004 (2018) (arxiv: 1712.05614)

## Slide 62
August 2015
Sensitivity / Efficiency
BACKGROUND: fast neutrons in B-10 detectors
Fast-N
gammas
Thermal-N
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on Boron-10 converter layers, JINST 13 P03004 (2018) (arxiv: 1712.05614)

## Slide 63
August 2015
Sensitivity / Efficiency
BACKGROUND: fast neutrons in B-10 detectors
Fast-N
gammas
Thermal-N
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on Boron-10 converter layers, JINST 13 P03004 (2018) (arxiv: 1712.05614)
200keV
100keV
10-5
Difference between He3 or B-10
(Gas Dominates)
Fast Neutron Sensitivity
B-10

## Slide 64
August 2015
BACKGROUND
DUE TO
SCATTERED NEUTRONS

## Slide 65
August 2015
Detector window
gas
B-10
Detector window
He-3
Detector cells or voxels
He-3 detector B-10 detector
N
N
gas
BACKGROUND: scattering
N
N
N
N
Scattering at the detector window: B10 has generally thinner windows due to lower pressure
Scattering at other parts of the detector

## Slide 66
Multi-Blade
sample
neutrons
MWPC
32 Wires
A cassette (unit)
10B-detector for reflectometers:
Estia and FREIA
neutrons
130mm
100mm
BACKGROUND: scattering

## Slide 67
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
strips Y
wires X
substrate
10B4C
Detector window
N

## Slide 68
Multi-Blade
detector
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
strips Y
wires X
substrate
10B4C
Detector window
N
N

## Slide 69
strips Y
wires X
substrate
10B4C
Multi-Blade
detector
Detector window
N
Image of the detector
Beam transmitted
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
N

## Slide 70
strips Y
wires X
substrate
10B4C
Multi-Blade
detector
Detector window
N
Image of the detector
1D projection
Beam transmitted
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
N

## Slide 71
strips Y
wires X
substrate
10B4C
Multi-Blade
detector
Detector window
N
Image of the detector
1D projection
Multi-Blade
detector
Beam blocked to evaluate instrument bkg
Beam transmitted
Beam blocked
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
N
N

## Slide 72
strips Y
wires X
substrate
10B4C
Multi-Blade
detector
Detector window
N
Image of the detector
1D projection
DETECTOR WINDOW
50um Aluminium foil
Beam transmitted
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
N
1mm Aluminium

## Slide 73
He-3 – AMOR detector
Multi-Blade – B-10
10-4
10-2
100
10-6
Instrument
background
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
* Thanks to A. Glavic and J. Stahn

## Slide 74
He-3 – AMOR detector
Multi-Blade – B-10
10-4
10-2
100
10-6
Instrument
background
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
(normalized to same area)
Figure-of-Merit ∝ Signal-to-Background
* Thanks to A. Glavic and J. Stahn

## Slide 75
He-3 – AMOR detector
Multi-Blade – B-10
10-4
10-2
100
10-6
(normalized to same area)
Instrument
background
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, (2020) (arxiv: 2001:02965). Accepted for publ. in JINST.
Detector | FoM (2.5-15Å)
He-3 (a few mm Al window) | 4000
Multi-Blade B10 – Al 1mm window | 73000
Multi-Blade B10 – Al foil window | 211000
x20
x50
Figure-of-Merit ∝ Signal-to-Background
* Thanks to A. Glavic and J. Stahn

## Slide 76
August 2015
Conclusions

## Slide 77
August 2015
Conclusions
@ESS 23 Hz/m2 cosmic thermal neutrons and 1 Hz/m2 with shielding
1

## Slide 78
August 2015
Conclusions
@ESS 23 Hz/m2 cosmic thermal neutrons and 1 Hz/m2 with shielding
Aluminium foil as a detector window reduces x50 the background generated by scattered neutrons
1
2

## Slide 79
August 2015
Boron-10 |  | Helium-3
0.5 – 0.8 | Thermal N Efficiency | 0.6 - 1
10-5 | Fast neutron sensitivity (gas dominates) | 10-3
10-6 - 10-9 | Gamma-ray sensitivity (solid dominates) | 10-6 - 10-9
Highly affected by a small variation of the threshold
Conclusions
@ESS 23 Hz/m2 cosmic thermal neutrons and 1 Hz/m2 with shielding
Aluminium foil as a detector window reduces x50 the background generated by scattered neutrons
1
3
2
(*results are general and applicable for any facility)

## Slide 80
Gamma-ray sensitivity
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)
F. Piscitelli et al., The Multi-Blade Boron-10-based Neutron Detector for high intensity Neutron Reflectometry at ESS, JINST 12 P03013 (2017) (arxiv: 1701.07623)
References
Fast neutron sensitivity
F. Piscitelli et al., Verification of He-3 proportional counters fast neutron sensitivity through a comparison with He-4 detectors, sub. to EPJ Plus (2020) (arxiv: 2002.08153)
G. Mauri et al., Fast neutron sensitivity for 3He detectors and comparison with Boron-10 based neutron detectors, EPJ TI 6, no. 1, p. 3, (2019) (arxiv: 1902:09870)
G. Mauri et al., Fast neutron sensitivity of neutron detectors based on Boron-10 converter layers, JINST 13 P03004 (2018) (arxiv: 1712.05614)
Scattered Neutron Background
G. Mauri et al., The Multi-Blade Boron-10-based neutron detector performance using a focusing reflectometer, Accepted for publ. in JINST (2020) (arxiv: 2001:02965).
F. Piscitelli et al., Characterization of the Multi-Blade 10B-based detector at the CRISP reflectometer at ISIS, JINST 13 P05009 (2018) (arxiv: 1803.09589)
G. Galgoczi et al., Investigation of neutron scattering in the Multi-Blade detector with GEANT4 simulations, JINST 13 P12031 (2018) (arxiv: 1810:06241)
E. Rossi, Master. thesis, Characterisation of the Spatial Resolution and the Gamma-ray Discrimination of Helium-3 Proportional Counters (2015) (arxiv: 1702:06501)
J. Birch et al., Investigation of background in large-area neutron detectors due to alpha emission from impurities in aluminium, JINST 10, P10019 (2015) (arxiv: 1507:00607)
J. Birch et al., In-beam test of the Boron-10 Multi-Grid neutron detector at the IN6 Time-of-Flight spectrometer at the ILL, J. Phys. Conf. Ser. 528:1, 012040 (2014)

## Slide 81
August 2015
BACKUP SLIDES

## Slide 82
August 2015
Detector window
gas
B10
Detector window
Window+(conversion+detection)
Window+conversion+detection
He3
Detector cells or voxels
He3 detector B10 detector
N
N
gas
Usually thicker for He3

## Slide 83
August 2015
BACKGROUND
DUE TO
SCATTERED NEUTRONS

## Slide 84
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
sample
Background
Background
0.5 2 6Å
Multi-Blade
detector
(Log scale)
October 2017 - CRISP reflectometer @ ISIS

## Slide 85
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Background
Background
Gating between 6ms and 20ms
0.5 2 6Å
(Log scale)
(Log scale)
sample
Multi-Blade
detector

## Slide 86
detector area
detector ToF
reflected intensity
ToF of the
reflected intensity
ToF (or l)
6ms (2Å)
20ms (6Å)
Background
Background
Gating between 6ms and 20ms
0.5 2 6Å
(Log scale)
(Log scale)
sample
Multi-Blade
detector
Projection over Y-axis

## Slide 87
2Å
90%
98%
61%
45%
Nominal 10B4C coating 7um
Actual 10B4C coating 4um
0.5Å
100 neutrons
strips Y
wires X
substrate
10B4C
55 neutrons get scattered @ 0.5Å
sample
Multi-Blade
detector
~4⋅10-3

## Slide 88
neutrons
detector
MB detector
Selene guide
sample
flight tube
November 2018 - AMOR reflectometer @ PSI
WP5 Data Acquisition software chain has been developed during BrightnESS (WP5 - i.e. DMSC/Data) and tested @ AMOR

## Slide 89
neutrons
detector
MB detector
Selene guide
sample
flight tube
November 2018 - AMOR reflectometer @ PSI
flight tube
detector
ESTIA
sample

## Slide 90
flight tube
Multi-Blade
Selene guide
sample
FRAME OVERLAP
MIRROR
SAMPLE
SLIT 4
MID FOCAL POINT
MB DETECTOR
4.9 m
L ~ 3.4 m
SLIT 1
FOCAL POINT
CHOPPER
SELENE 1
SELENE 2

## Slide 91
Beam transmitted
Beam blocked
Cd slit to block the beam far from the detector with shutter open
FRAME OVERLAP
MIRROR
SAMPLE
SLIT 4
MID FOCAL POINT
MB DETECTOR
4.9 m
L ~ 3.4 m
SLIT 1
CHOPPER
SELENE 1
SELENE 2

## Slide 92
Beam transmitted
Beam blocked
Y-Projection

## Slide 93
Beam transmitted
Beam blocked
Y-Projection
1mm Al window
50um Al foil

## Slide 94
Y-Projection
~ 3⋅10-5
~ 10-4
~ 2⋅10-5
1mm window
Foil window
Instrument background

## Slide 95
Y-Projection
CRISP (ISIS) demonstrator
4.4um 10B4C
AMOR (PSI) demonstrator
8um 10B4C
100 neutrons
strips Y
wires X
substrate
10B4C
55 neutrons get scattered @ 0.5Å
~4⋅10-3
~ 3⋅10-5
~ 2⋅10-5
Problem Solved!

## Slide 96
He3 – AMOR detector
Multi-Blade - B10
10-4
10-2
100
10-6
(normlized to same area)
Foil ~ x3 ~ 1mm window ~ x20 ~ He3
Foil ~ x60 ~ He3
Intrument
background

## Slide 97
August 2015

## Slide 98
August 2015
FAST NEUTRON
BACKGROUND

## Slide 99
August 2015
THEORY
Cross-sections
We can consider only recoil and (n,p)
Similar probability:
to happen recoil in He4
as recoil+(n,p) in He3!

## Slide 100
August 2015
THEORY
Macro corss-sect
And Probability
for 2.54cm
(tube diam.)
Same number density for He3 and He4 -> n = 2.43e19 1/cm3 @ 1 bar
Mass density: He4 0.00016 g/cm3 and He3 0.00012 g/cm3 @ 1bar

## Slide 101
August 2015

## Slide 102
Set-up and Methodology. Background Measurements @ CRISP reflectometer
2. Indirect Calculation
Spectrum sketch scenarios:
O: chopper in phase with the proton pulse. Beam passes through the chopper
C: chopper not in phase with the proton pulse. Neutron beam hits the chopper
TS2: 1 every 5 pulse is sent to Target Station 2
25Hz

## Slide 103
2. Indirect Calculation
Spectrum region of interest:
(T): 𝚽 total flux integrated in the full spectrum t = 0-200 ms
(O): 𝚽tn thermal flux integrated in t = 120-160 ms
(C): 𝚽p background flux integrated in t = 100-105 ms
Measurements & Calculation. Background Measurements @ CRISP reflectometer

## Slide 104
2. Indirect Calculation
Spectrum region of interest:
(T): 𝚽 total flux integrated in the full spectrum t = 0-200 ms
(O): 𝚽tn thermal flux integrated in t = 120-160 ms
(C): 𝚽p background flux integrated in t = 100-105 ms
Measurements & Calculation. Background Measurements @ CRISP reflectometer

## Slide 105
2. Indirect Calculation
Spectrum region of interest:
(T): 𝚽 total flux integrated in the full spectrum t = 0-200 ms
(O): 𝚽tn thermal flux integrated in t = 120-160 ms
(C): 𝚽p background flux integrated in t = 100-105 ms
Measurements & Calculation. Background Measurements @ CRISP reflectometer

## Slide 106
2. Indirect Calculation
Indirect calculation
Calculated fast neutron flux
Measurements & Calculation. Background Measurements @ CRISP reflectometer

## Slide 107
Measurements & Calculation. Background Measurements @ CRISP reflectometer
2. Indirect Calculation
(T)
(O)
(C)
 | 𝚽 | 𝚽tn | 𝚽p | 𝚽fn
MB | 6.6∙104 ± 250 | 1.56∙104 ± 120 | 870 ± 30 | 900 ± 140
3He | 3.95∙105 ± 600 | 2.45∙104 ± 150 | 7.38∙104 ± 300 | 7.4∙104 ± 200

## Slide 108
August 2015
GAMMA-RAY
BACKGROUND

## Slide 109
August 2015
A. Khaplanov et al., Investigation of gamma-ray sensitivity of neutron detectors based on thin converter films, JINST 8 P10025 (2013) (arxiv: 1306:6247)

## Slide 110
August 2015

## Slide 111
August 2015

## Slide 112
August 2015