Discussion baseline for potential rig designs of Rig2 and Rig3 for the ESS.

- All rigs have the option to rotate a sample inside under load.

- Please update/complete the document and save as “RigDesigns-Rig+12_ESS_v(n+1)”

Contents

Table with Specs	2

Psylotech	4

Deben stages	4

FRM2 rig(s)	5

Robin rig(s)	6

RAMS	7

# Table with Specs

| Rig | Sample rotation | Axial Force compression | Axial Force max | Max Torsion | Stroke | height | Weight | Additional options | Distance between bars: | Costs | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ESS Rig1 (exists at NPI) No sample rotation | No |  | 100kN (Nominal load: 60kN) | None |  | 2300mm (Can be shortened as well) | 600kg | Vacuum chamber Cooling system is in place | 600mm |  | Integration with ESS SW and HW ongoing |
| ESS Rig2 (conceptual) | Yes |  | 50kN | 100Nm |  |  | <200kG | Furnace (gas flow enough, i.e. no vacuum))? Enough width between columns Adjustable height |  |  |  |
| ESS Rig3 (conceptual) | Yes ? |  |  |  |  |  |  | Low cycle Fatigue? |  |  |  |
| Psylotech xTS | Yes |  | 1.6kN - 45 kN | 350Nm | 50-100mm | 40cm -110 cm 600mm | 6 – 135kg | temperature chamber available |  | xTS 45kN Base System* - $195K 300Nm Torsion Add-On - $50K xTS Temp Chamber(-100C to 300C) - $65K CT Scanner Integration - $18K On-Site Installation - $4900 |  |
| Deben CT20K | yes | -20kN | 20kN | 100Nm |  |  | 120kg | optional furnace 1000C |  |  |  |
| FRM2 rig (for diffraction and not imaging so far) | Not full 360deg? Need to check. | 50KN(?) | 50kN | 100Nm (max angle 49deg) | 75mm | ? | ? | Mirror furnace up to 1000C. Euler cradle |  |  |  |
| Robins rig (UTK) | Yes: freely | -24kN | Range of 5 load cells: 900N, 9kN, 11kN/11Nm, 22kN/111Nm, 44kN | 11Nm (110Nm) | 80mm | 70cm | 55kg | Flexible design (height, width can be adjusted). Alignment software developed (LabView) based on ASTM standard. |  | 6500 Euro parts 12000 Euro Controllers (NI) 8000 Euro for set of load cells (no development/machining included) |  |
| RAMS | Yes | 2-kN | 2kN | None |  | 73cm | 73kg |  |  |  |  |

# Psylotech

https://www.psylotech.com/loadframe/

# Deben stages

Commercial systems that are optimized for uXCT and synchrotron. The CT20K looks intertsing.

https://deben.co.uk/tensile-testing/open-frame-torsion-tension-and-compression-stage/

# FRM2 rig(s)

Hoelzel, M., Gan, W. M., Hofmann, M., Randau, C., Seidl, G., Jüttner, P., & Schmahl, W. W. (2013). Rotatable multifunctional load frames for neutron diffractometers at FRM II—design, specifications and applications. Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, 711, 101-105.

Developed at FRM2 and in use at Stress-Spec and SPODI. HZB (E3) bought uniaxial version.

# Robin rig(s)

https://trace.tennessee.edu/utk_graddiss/3375/

CAD available in SolidEdge

# RAMS

Shade, P. A., Blank, B., Schuren, J. C., Turner, T. J., Kenesei, P., Goetze, K., ... & Lienert, U. (2015). A rotational and axial motion system load frame insert for in situ high energy x-ray studies. Review of Scientific Instruments, 86(9), 093902.

Designed for small samples with 1mm x 1mm cross section. Interesting alignment features that could be copied.

Focus on alignment fixtures: