## Slide 1
Status of the ESS Cryomodule for Elliptical cavities
M-ECCTD
Unité mixte de recherche CNRS-IN2P3
Université Paris-Sud 11
91406 Orsay cedex
Tél. : +33 1 69 15 73 40
Fax : +33 1 69 15 64 70
http://ipnweb.in2p3.fr
Gilles OLIVIER
for the CEA/DSM/IRFU and CNRS/IN2P3/IPNO teams

## Slide 2: Cryomodule for elliptical Cavities
SUMMARY
Introduction
Medium and high beta cryomodules standardization
General design
Heat balance
Cryogenic circuits
Tooling and assembly process
Assembly and test facilities

## Slide 3: Introduction
IPN ORSAY in charge of the cryostat,
design and fabrication
Vacuum vessel
Space frame – cavity supports
Thermal screen
Superinsulation
Internal cryogenic pipes
Instrumentation
CEA SACLAY in charge of of the “cavity package”, the cryomodule assembly and RF power tests
Cavities + helium tank
Power coupler design, manufacturing, RF tests
Piezo tuner
Magnetic shield
Tooling: field flatness, cavity preparation, cryomodule assembling, …
Cryomodule assembly
Tests stand for cryogenic and RF power tests
Tests of the ECCTD cryomodule
Development of a technology demonstrator
with medium beta elliptical cavities (M-ECCTD)
Work done within the FR-SW agreement - collaboration IPNO and IRFU
The design of the cryomodule is compatible with both medium and high beta cavities

## Slide 4: High and medium β cryomodules standardization
High beta (0,86):
5 cells
Length 1316,91mm
Medium beta (0,67):
6 cells
Length 1259,40mm
57,5 mm

## Slide 5: High and medium β standardization
Identical sub-assemblies:
Vacuum vessel
Spaceframe
Thermal shield
Cryogenics circuits
Position of the nozzle
Ø418
20
188
900
120,3 (HB)
115,3 (MB)
180
908
Blocks for half ring fasteners
Position of the nozzle
1316,91 (HB) / 1259,4 (MB)
Main identical dimensions
Main differences
Same mechanical interfaces in the tunnel
250
794
Specific components:
Inter cavities bellows
Cold/warm transitions
1500 between couplers

## Slide 6: OVERALL DIMENSIONS
506,3
1500
6329,6
6587,6
1500
1500
450
4300 (between supports)

## Slide 7: OVERALL DIMENSIONS
2745
Ø324
1500
341
604.5
293
Tunnel
1245
685

## Slide 8: General design

## Slide 9: overview
Beam

## Slide 10: overview
HIGH BETA CRYOMODULE
Vacuum vessel support (x2)
Supporting jacks
Exchanger
Low pressure line
Spaceframe
Coupler
Door knob
Bias system

## Slide 11: Heat balance

## Slide 12: Cryogenic diagram
Nuno ELIAS 2014-03

## Slide 13: Helium Pressure relief devices
Working pressure = nominal pressure
WP
PSET Set Pressure
PS
Maximum allowable overpressure
PS + 10%
Bursting disk
Safety devices
PSET
POPENING = + 5 % PSET
(overpressure when opening)
PMAX APERTURE = + 10 % POPENING
PCLOSING = + 5 % PSET
PMIN CLOSING = - 10 % PCLOSING
(hysteresis before closing)
PSET + 50mbar
PSET - 50mbar
Relief valve
Bursting area
Opening area
Cavity mechanical behaviour limit
European rules compliance
ESS Helium factory, circuits
and operating modes
WP upper limit
Max. allowable press.
+10%

## Slide 14: Pressure vessels European rules
Vessels
Pipes
PS<0,5bar
The equipment is not on the scope of the 97/23/CE directive
Article 3-3
The equipment must be designed and manufactured according to workmanlike way
Category I
The manufacturing must be more documented, especially with internal production control
THE GOAL Article 3-3

## Slide 15: Helium low pressure circuit
2 bursting disk at each tip
+ upstream safety relief valve (with he guard)
36°
Heat exchanger
The circuit is designed to reduce as low as possible the overpressure in case of beam vacuum failure by using a continuous DN100 diameter for the diphasic pipe, large curvatures and 2 DN100 bursting disks at each extremity.
A 36° angle is set up for the tank nozzle in order to allow the insertion of the cavity string and the cooling circuit inside the spaceframe
To valve box
Worst scenario: beam vacuum failure
38KW/m2 applied on the cavity wall
Accidental overpressure: 550mbar reduced
to 230mbar after rupture of the disks

## Slide 16: Assembly procedure
Acknowledgments:
Amaury MARTIN (engineer student thesis)
Nicolas BAZIN (CEA)
ASSEMBLY PROCEDURE AND TOOLING FOR THE M-ECCTD
(Medium beta prototype)

## Slide 17: COUPLER Assembly
Choice of the vertical position below the cavity for the coupler during the assembling operation

## Slide 18: Cavity string assembly
Clean room ISO 5

## Slide 19: Cavity string assembly
Installation of the central trolley/rail

## Slide 20: Cavity string assembly
Installation of the extremity trolleys

## Slide 21: Cavity string Assembly
Installation of the extremity trolleys

## Slide 22: Cavity string assembly
Installation of support posts

## Slide 23: Cavity string assembly
Assembly, positioning and pre-alignment of the cavities

## Slide 24: Cavity string assembly
Assembly of the inter-cavity bellows

## Slide 25: Cavity string assembly
Assembly of the cold to warm transitions

## Slide 26: Exit of Clean room

## Slide 27: Exit of Clean room

## Slide 28: Exit of Clean room
Removing of the extremity trolley

## Slide 29: Exit of Clean room
2 m

## Slide 30: Exit of Clean room
2 m

## Slide 31: Exit of Clean room

## Slide 32: Transfer frame
4 parallel rails
Length about 8m
Gaps: 500 and 900mm

## Slide 33: Assembly outside the clean room
Alignment and fixing of the trolley

## Slide 34: Assembly outside the clean room
Transfer of the cavity string

## Slide 35: Assembly outside the clean room
End of transfer

## Slide 36: Assembly outside the clean room
Assembly of:
Cold Tuning System
Insulation
Magnetic shield
Cryogenic lines
(not shown)

## Slide 37: Assembly outside the clean room
Assembly of the spaceframe supporting system

## Slide 38: Assembly outside the clean room
Assembly of the spaceframe supporting system

## Slide 39: Assembly outside the clean room
Assembly of the spaceframe and the thermal screen (not shown)

## Slide 40: Assembly outside the clean room
Thermal shield not shown
Moving of the spaceframe and the thermal shield around the cavity string
support of the cavity string and the thermal shield inside the spaceframe by means of the supporting system (TA6V rods)

## Slide 41: Supporting system
Rods
Magnetic
shield
Half rings linked to the tank (under the magnetic shield)
Helium tank
Cross rods fixed
on 2 titanium half rings fastened (on their middle only) to the helium tank on one side
On the spaceframe on the other side
3000N pre-stress applied on the rods
Maximum force 8500N per rod after cooling down

## Slide 42: Thermalization of the supporting rods Hanging of the Thermal shield
Elastic and deformable boxes
Thermal shield
Supporting rods
The boxes are blocked only at the fixed point of the thermal shield
Locking jaws and thermalization
FOLDED BOXES ENSURE THE AXIAL MOVING AND THE THERMALIZATION OF THE THERMAL SHIELD.
THEY PERMIT ONLY LOW DISPLACEMENTS ON OTHER DIRECTIONS
3 prototypes have been made (thickness 0,5 , 1 et 1,5mm)

## Slide 43: cryostating
Insertion of the whole assembly inside the vacuum vessel
2 wheels are fixed to the spaceframe at each extremity.
The guiding is ensured by two rails welded to the vacuum vessel

## Slide 44: cryostating

## Slide 45: cryostating

## Slide 46: cryostating
End of cryostating.
Finishing operations: final positioning, valves and jumper connections, closing of the cryomodule…

## Slide 47: Alignment of the cavity string
The same fiducials are available before and after cryostating
Reflector brackets
(coupler flange)
Reflector brackets
(cavity flange)
Reflector brackets
(vessel alignment)
Alignment made by laser tracker
Alignment of the cryomodule in the tunnel (source ESS)
1,5" Corner cube reflector
(source ESS)
Alignment of the cavity string within 1,5mm of the beam axis
Additional fiducials will be set up inside the ECCTD cryomodule to check the alignment after cooling down.

## Slide 48: Test facilities at CEA saclay
Development and installation:
Cryogenic lines (L>50m) + Dewar + valves box + jumper line
C/C system
RF wave guide line at 704MHz
Modification of the klystron modulator to increase the pulse length up to ESS requirements (2ms => 3ms)

## Slide 49
The new clean room at Saclay
La salle blanche ISO7-ISO5 a été construite pour réaliser les assemblages de cavités supraconductrices dans les conditions de faible empoussièrement.
La salle blanche ISO5 accueillera les montages. Elle est équipée de filtres (FFU - Fan Filter Unit).
Le HPR (rinçage Haute Pression) permet de rincer l’intérieur des cavités. La cavité est fixée sur le plateau monte/baisse. La canne est fixe.
Intérieur de l’ISO5
Local HPR
1
2
4
5
1
4
Intérieur de l’ISO7
Dans la salle blanche ISO7, l’extérieur des cavités est nettoyé (SAS humide) et les cavités sont préparées.
2
3
Le local d’eau ultra-pure contient
le système de filtration par résine.
3
1 | Stéphane BERRY | 4 | Alain PEROLAT
2 | Alain CATALDI | 5 | Richard LE GUENNIC
3 | Fabien EOZENOU | Mardi 13 Mai 2014 | 
Local d’eau Ultra-Pure
La centrale de traitement d’air permet de faire circuler l’air dans la salle blanche à XXXM3/h
5
Centrale de traitement de l’air

## Slide 50: Thank you for your attention