## Slide 1: Science Directorate Instrument Technologies division Neutron Chopper Group Update IKON6
Iain Sutton
Neutron Chopper Group Leader
www.europeanspallationsource.se
February 26, 2014
Notes: Good morning
My name is Iain Sutton
I am currently the Group leader for Neutron choppers
And In the following 10mn I will attempt to present the Status and strategy of NG Tech work package

## Slide 2: The story so far….

## Slide 4: Spallation source instruments
PHOTO ISIS Instrument
- overlay schematic components
CHOPPER 3
CHOPPER 1
CHOPPER 2
NEUTRON OPTICS
- GUIDE
NEUTRON OPTICS
- POLARIZER
NEUTRON OPTICS
- COLLIMATION
DETECTOR
-ARRAY
DETECTOR
-BEAM MONITOR
DETECTOR
-BEAM MONITOR
FLIGHT TANK
BEAM LINE SHIELDING
SAMPLE
ENVIRONMENT
DETECTOR
-BEAM MONITOR
NEUTRON OPTICS
- IN MONOLITH GUIDE
BEAM SHUTTER
-LIGHT
BEAM SHUTTER
-HEAVY
Cut away of LET on TS2
Courtesy of ISIS
Notes: ----- Meeting Notes (2/26/14 12:36) -----
The basic anatomy of a neutron spectrometer at a spalation source
Suffice to say that choppers are one of the principal components of most of them and that is unlikely to change at ESS

## Slide 5: Bunker – Preliminary view
Notes: What we may me doing a little differently that at other facilities is our approach to shielding
A highly intergrated approach between beamlines which currently supposes a common bunker housing the instruments fornt end
Including around 2 thirds of all the facilities choppers

## Slide 6: Chopper cavities in Common shielding concept – side view
Bulk shield
- upper –
Heavy concrete
Bulk shield
Mid section
Bulk shield
- lower -
Bulk shield
- Access plug –
Steel (?)
Neutrons
R6m
R10m
Notes: Bulk shielding – heavy concrete – grey
Removable sections for acces in green. (in steel)
Chopper and other equipment will be housed in cavities
Current thinking is to design and install as modules to aid removal
Removable sections less than 10.

## Slide 7: Beam line separation < 10 deg Complicated chopper installation
Notes: For this reason the installation of choppers within the shielding has come in for a lot of thought.
Particularly with respect to a decision on beamline separation
with an angular separations in multiples of 5 degrees seem probable.
Short and medium length instruments will likely be separated by 10 or 15 degrees
Long Instruments situated in Hall 3 would be separated by as little as 5
As one might imagine, close to the monolith this limits the space available for the all important 1st choppers and compromises will have to be made

## Slide 8: Intergrated Systems
NCG / E & E Group
NCG / E& E / Integrated Control Systems
NCG / E&E / DMSC
Neutron Chopper System
N0&S Group
N0&S Group
N0&S Group
Notes: Choppers are not only tighly integrated with shielding
But also within the other systems composing and supporting an instrument
Much of this equipment is the responsibility of other groups within the company an the issues have to be worked out together …

## Slide 9: Chopper Interfaces & Stake holders
DMSC
Servo Control
Electronics & Parameters
Control box
Software
(device driver)
Motion control
Software
Control Network
HW / SW
Electronic
Hardware &
Middleware
Instrument
control Software
ICS
E &E Group
Control
Box
Chopper
Hi level Control algorithms
Communication
protocols
Data acquisition
Software
Chopper Integrated Slow motion
Control
Facility
Instrument specific components
Instrument specific installations
Instrument # 22
Instrument specific control parameters
Instrument #2
Instrument availability
Machine Protection Systems
Personnel Safety
Systems
Timing
signals
Instrument #1
Health &
Safety
Safe working
practices
Containment
Contamination
Detectors
Optics &
Shielding
Guide-Chopper Interface
Alignment
Neutron optics
Vacuum systems
Crane access & capacity
Chopper
Control
HW/SW
Chopper
Support
Systems
Building Services
Cooling /
heat recovery
Electrical Power & UPS
Ventilation
Integrated
Beamline
shielding
Sample
Environment
Operating schedule
Target
Monolith / instrument interface
Chopper
Mechanic
systems
Time stamping
Data Veto
Neutron Chopper Systems
Notes: In fact if you just try to identify the primary stakeholders and proximate issues concerning choppers you come up with something like this
I think you will agree it’s quiet a bundle
Establishing interfaces and tying up the loose end is currently proving quiet a challenge
But it’s a vital building block to enable the design and operation of instruments
Lets look at just one of these interfaces …

## Slide 10: Interfaces: Chopper-Optics-Shielding
Key Objectives
  Highest neutronic performance for instruments
  Low back ground @ Detector
  Highest instruments availability
  Lowest lifetime costs
Input to instrument construction projects
  Interface definitions
    Chopper-Optics
  Guidelines for design
    Installation
    Access / Serviceability
    Integration of shielding
  Policies
    Alignment /Vacuum /Access /Activation
  Equipment standards for procurement
Common housing strategies
Vertical split
Guide removed with Choppers
Horizontal split
Guide remains in lower housing
Notes: Very obvious and real interface Choppers / Guides
With the increased complexity of ESS choppers systems the numer of these interfaces is increased
Intergration
with guides
with shielding
Ensuring serviceability within shielding…
The list goes on…
The result of this effort will be condensed into a number of documents – many in the form of guidelines for instruments
Its important to the design of each instrument
But equally to the operation of the instrument as a whole
We do …. so you do not have to.

## Slide 11: A preliminary review of chopper systems included in current proposals
Proposal round 2013
Notes: Turning aside from the current preparations for construction I would like to share with you a brief look at some of chopper systems might look like

## Slide 12
BEAM PORT
# 1-2 3-4
BW SELECTION
FUNCTION
NOTES
NMX
GUIDE
2PA-02 (FTF)
14Hz, D700
L 20 50 156
2PA-02 (FTF)
BW SELECTION
14Hz, D700
GUIDE
GUIDE
Status: P1-P.Design
Chopper system schematic
EQUIP ACCESS ZONE
CLASS : BLACK
EQUIP ACCESS ZONE
CLASS : GREEN
Sector: W
Beam-port: W12
Notes: Wide range of designs and system complexities
Ranging from No choppers to more than 20
These are simple ‘stick representations’ and I do not propose to analyse them at length
But flick though a number to give you an over view
Source on the left
Sample on the right

## Slide 13
2PA-01-M-350 (CR)
BEAM PORT
GUIDE
GUIDE
# 1-2 3-4 5 6-7 8-9
PULSE SHAPING
FUNCTION
NOTES
CS/ VOR
Phase 0: Proposal
Schematic
PA-02 (-M-14)
GUIDE
MONOCHROMATOR
PHASE(D) / CO
RRM
BAND WIDTH
2PA-01-M-350 (CR)
2PA-01-M-!$) (FTF)
2PA-01-M-400(CR)
COMMON SHIELDING
SHUTTER
SHUTTER
POSSIBLE SHUTTER POSITION
POSSIBLE SHUTTER POSITION
L 8 16-16.05
Sector: Beam port: NA

## Slide 14
BEAM PORT
7-14Hz, D700
# 1-2 3 -4 5 6-7 8 9-10 11
WAVELENGTH BAND DEFINING
FUNCTION
NOTES
RF/ FREIA v2.0
GUIDE
FRAME OVERLAP
L 6.5/6.6 7.1/7.5 8,5 10/10.1 11.1 15 15,6 25
2PA-02 (FTF)
PA-02 (DC)
2PA-02 (FTF)
PA-02 (DC)
GUIDE
GUIDE
GUIDE
WAVELENGTH BAND DEFINING
WAVELENGTH BAND DEFINING
FRAME OVERLAP
FRAME OVERLAP
PULSE SHAPING
7-14Hz, D1300
7-14Hz, D1300
56Hz, D1300
56Hz, D1300
42Hz, D1300
28Hz, D1500
EQUIPT ACCESS CLASS : ORANGE
EQUIPT ACCESS CLASS : GREEN
GUIDE
2PA-02 (UO)
2PA-02 (FTF)
PA-02 (DC)
GUIDE
GUIDE
Status: P0-Proposal
Schematic
Sector: E/N Beam-port: ?

## Slide 15
BEAM PORT
PA-02 (DC)
GUIDE
GUIDE
280Hz, D700
# 1 2 3 4 5 6 7
PULSE SHAPING (CT1)
FUNCTION
NOTES
PD/ HEIMDAL
Status: P0-Proposal
Schematic
2PA-02 (FTF)
GUIDE
FRAME OVERLAP (CT3)
L 6,5 7 9 13 78 78,6 S = 169m
PULSE SELECTION (CT2)
PA-02 (DC)
14Hz, D1360
14Hz, D700
PA-02 (DC)
BAND DEFINITION (CC1)
14Hz, D1280
FRAME OVERLAP (CC3)
14Hz, D1280
PA-02 (DC)
FRAME OVERLAP (CC3)
14Hz, D1280
EQUIPT ACCESS CLASS : BLACK
PA-02 (DC)
EQUIPT ACCESS CLASS : TBD
Notes: The typical number of choppers on a concept is 7.5
This figure is 3 time higher than the average of instruments at reactor sources and twice that of short pulse spallation sources
The is nothing surprising in this, it is a result of techniques required to get the most out of a long pulse

## Slide 16
BEAM PORT
# 1-2 3-4 5 6-7 8-9 10-11 12-13 14-15 16-17 18-19
PPS
FUNCTION
NOTES
IM-ODIN
28Hz, D700
L 6,5 7 9,25 10 11 12 16 20 30 40 60
2PA-02 (FTF)
BAND WIDTH
GUIDE
GUIDE
PA-10 (PPSc)
2PA-02 (FTF)
2PA-02 (FTF)
2PA-02 (UO)
2PA-02 (UO)
PULSE SHAPING
PULSE SHAPING
56Hz, D700
BAND WIDTH
56Hz, D700
3,5-14Hz, D700
3,5-14Hz, D700
2PA-02 (FTF)
2PA-02 (FTF)
2PA-02 (FTF)
SUB FRAME OVERLAP
SUB FRAME OVERLAP
SUB FRAME OVERLAP
SUB FRAME OVERLAP
14-42Hz, D700-2000
14-42Hz, D700-2000
2PA-02 (FTF)
SUB FRAME OVERLAP
14-42Hz, D700-2000
14-42Hz, D700-2000
14-42Hz, D700-2000
GUIDE
GUIDE
GUIDE
GUIDE
GUIDE
GUIDE
Status: P1-P.Design
Sector: S
Beam-port: S12
Schematic
EQUIP ACCESS ZONE
CLASS : BLACK
EQUIP ACCESS ZONE
CLASS : GREEN

## Slide 17
A Summary of Choppers for 22 Instruments, based on 3+16 instrument concepts
22 Instruments = 150 – 160 axis in 6 ‘platforms’
LOW
SPEED
(50-70)
PA-02 (-M-14)
(15-35)
INTERMEDIATE
SPEED
PA-02 (-M-14)
HIGH
SPEED
(27)
2PA-02 (UO)
LARGE
ROTOR
(30)
2PA-02 (FTF)
(3 -5)
FAN
PPSc
(4-8)
PA-10 (PPSc)
Notes: Average choppers 7.5
Which 22 inst 150 -160 axis
By configuration, performance and requirements can be sub divided into 6 broad families
Small rotored disc chopper
Low / intermediate / high speed
Large rotored choppers

## Slide 18: Platform 1 Disc chopper – Small rotor – Low speed
DC-SR
(50-70)
Key requirements
Rotation speed: Low ( 7 - 96 Hz)
Openings: Large
Attenuation at short wavelengths
High reliability in radiation environment
Low lifetime cost
Principal characteristics
Type: PA-1-H- (horizontal axis disc)
Rotors : Simple + Robust
Diameter: 600-800mm
Enabling Technologies .
Rotors materials: Metallic
Bearings: Magnetic

## Slide 19: Platform 2 Disc chopper – Small rotor – Intermediate speed
DC-SR
(15-35)
Key requirements
Rotation speed: (96-192 Hz)
Openings: (Multiple) Small or Large
High reliability in radiation environment
Principal characteristics
Type: PA-1-H- (horizontal axis disc chopper)
Rotors: Simple or Optimized
Diameter: 600-800mm
Enabling Technologies .
Rotor material: CFRP / Alu
Bearings: Magnetic

## Slide 20: Platform 3: Disc chopper – Small rotor – High speed
Key requirements
Rotation speed: 192- 400+ Hz
Openings: (Multiple) Small
Minimal guide interruption
High reliability in radiation environment
Principal characteristics
Type : PA-1-H- (horizontal axis disc)
Rotors design: Optimised
Diameter : 600-700mm
Enabling Technologies .
Rotor material: CFRP / Ti / MMC
Bearings: Magnetic
DC-SR-HP
(25-35)
2.PA-1-H-M

## Slide 21: Platform 4: Disc chopper – Large rotor
DC-LR
(25-35)
Key requirements
Rotation speed: (7 - 56 Hz)
Openings: Multiple V.Large, Asymmetric
High closing speed
High reliability in radiation environment
Principal characteristics
Type: PA-1-H- (horizontal axis disc chopper)
Rotors: Optimized
Diameter: 1200 - 2000mm
Enabling Technologies .
Rotor material: CFRP / Alu
Bearings: Magnetic or Contact

## Slide 22: Platform 5: Disc chopper – Fan
DC-SR-FAN
(3-5)
Key requirements
Rotation speed: (7 - 56 Hz)
Openings: Variable width, Asymmetric
Principal characteristics
Type: XPA-1-H- (horizontal axis disc chopper)
Rotors: multiple concentric ‘Fan’
Diameter: 500 - 800mm
Enabling Technologies .
Rotor material: Aluminium
Bearings: Contact

## Slide 23: Platform 6: Prompt pulse suppression chopper PPSc
PPSc
(4-8)
PA-10 (PPSc)
Key requirements
Rotation speed: (7 - 56 Hz)
Closure: 3ms, Symmetric or Asymmetric
Attenuation : 90% @ Prompt pulse energies
Extreme radiation resistance
Principal characteristics
Type: PA-10-H- (horizontal axis chopper)
Rotors: 300 – 400 thick , single or double
Diameter: 500 - 600mm
Enabling Technologies .
Rotor material: Nickel alloy / Tungsten
Bearings: Magnetic

## Slide 24: approach
Objectives
Maximize instrument performance by considering the needs of the instrument suite and that of the facility globally and not case by case.
Increase the effectiveness of collaborations by matching partner competence and capacity to equipment needs.
Reduce development costs by avoiding duplication of effort and sharing equipment across several instruments
Ensure high serviceability to achieve highest standard of support to instruments in operation.
Notes:   Identify the common requirements of Science & the facility Serviceability, maintainability, instrument availability, and even cost effectiveness…
  Increase serviceability and reduce costs through parts commonality.
  Explore the potential for solutions which can meet ‘sub sets’ of requirements
  Explore the potential for providing solutions within a family of choppers.
  Match build partners technical & supply potential to family groups.

## Slide 25: Procurement - Some options : for ‘standard’ systems
SYSTEMS INTERGRATORS
TECHNOLOGY PROVIDERS
SYSTEMS
Notes: Despite the fact that ESS unique many of the systems may be constructed for existing industry standard or ‘near standard’ components
Typical disc choppers / fermi
Commercial suppliers

## Slide 26: Procurement options - some ideas for ‘long pulse specific’ systems ?
SYSTEMS INTERGRATORS
Notes: ESS unique requirements
Unique Long pulse will require some equally unique solutions
Special PPSc choppers adapted to the long pulse and higher energies
Completely unique large rotor choppers
Diameters 2 or more meters
Critical to employing the WFM and (RRM ?) techniques an harnessing the potential of the long pulse
No european commercial suppliers for complete systems
Current spallation sources design own
Suppliers of critical subsystems are known and we propose to work with them in much the same manner

## Slide 27
Development & Build partners the discussion is open …
LOW
SPEED
INTERMEDIATE
SPEED
HIGH
SPEED
LARGE
ROTOR
FAN
PPSc
POSSIBLE BUILD PARTNERS ?
BUILD PARTNER ‘F’
BUILD PARTNER ‘C+G’
BUILD PARTNER ‘B’
BUILD PARTNERS ‘A+H’
BUILD PARTNER ‘E’
BUILD PARTNER ‘D’

## Slide 28: Chopper system breakdown
Neutron Chopper System
Mechanical systems
Control systems
Support systems
Pit systems
Chopper 1-2 Low speed
Chopper 3-4 Low speed
Rack & Cabling
Control Hardware
Software
Vacuum
Power supply
Cooling
Monitoring
Pit 1 module
Pit 2 module
Extraction tooling
Service pack
Tooling
Critical Spares
Chopper 5 PPsc
Chopper 6–7 Hi-Speed
Procurement value ~70 % of total
Procurement value ~10%
Predominantly generic systems
Procurement value ~20%
Predominantly Instrument specific
Notes: The facilities goals of performance, cost and quality are attained by the retaining of high level control over design and procurement parameters and the selection of equipment on and ‘Instrument suite wide basis’.
Establish Standards, Guidelines
Manage development & procurement through technical groups via the delegation of instrument work units.
An optimal match to build partners and a maximal use of in-kind potential.
Facility does not benefit from the ‘Fruit cake’ model of building instruments.
  Commonality is lost
  Cost advantages of scale are lost
  Duplication of effort

## Slide 29
 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025 | 2026 | 2027 | 2028
Operations Schedule |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#5 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#6 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#7 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#9 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#10 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#11 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#12 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#13 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#14 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#15 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
#16 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Initial Design | Final Design | Procurement & Fabrication | Construction | Installation | Cold Commissioning | Hot Commissioning
1st Neutrons
Instrument projects overview

## Slide 30: Chopper & Instruments
Initial Design | Final Design | Procurement & Fabrication | Construction | Installation | Cold Commissioning | Hot Commissioning
1st Neutrons
On site Installation
Delivery, Integration & Test
Procurement
Development
Commissioning Interventions
Target Operations Schedule |  |  |  |  |  | 
#1
#2
#3
T-60m T-48m T-36m T-24m T-12m T0

## Slide 31: Science directorate / Neutron technologies division - Interim mechanical workshops -
Skiffervagen 26, 224 78 ,Lund
Space
Building ESS’s capacity to innovate….
CURRENT STATUS
Tools
Equipment

## Slide 32: Neutron technologies workshops
Chopper test cell
Real work is commencing!
Basic machine shop
Prototype assembly area
Optics group test station
Electronic assembly area

## Slide 33: THE END
Thank you for you attention

## Slide 34: A chopper extraction concept

## Slide 35: The End