## Slide 1: 10B-based MultiGrid detectors as an alternative to 3He: Achievements and construction of a large area demonstrator
Mathieu FERRATON
ILL, Neutron Detector Group
ESS Detector Group seminar

## Slide 2: Outline
Basic principles and design of MultiGrid detectors
Achievements with previous prototypes
The IN5 demonstrator

## Slide 3: 3He crisis and need for alternatives
3He gas detector
  demonstrated performance
  relatively simple use
3He is a byproduct of nuclear industry and nuclear weapons
  Product of tritium beta-decay (T1/2 ~ 12 years)
  available quantity is limited
Around 2008, USA government began to keep their 3He stock for Homeland Security Department
-> 3He shortage for scientific applications
-> cost of 3He increased by a factor ~40 (from ~100€/L to ~4000€/L)
-> a new IN5 detector (the largest at ILL) would cost about 10M€ just for the gas itself
ESS and other neutron scattering science facilities need hundreds (thousands) of square meters active area detectors
-> 3He is not an option anymore for the largest detectors
-> alternative technologies are needed shortly!!
We have a solution!!!!

## Slide 4: Basic principles of MultiGrid detectors Gas neutron detectors based on solid 10B converters
Solid 10B4C layers are used as neutron converters, coupled with a gas amplification system
Unlike 3He gas detectors, the neutron converter is not the amplification medium
~40 µm of 10B4C is required to convert 90% of thermal neutrons
Ranges of fragments between 1.7 and 3.3 µm
neutron
4He
7Li
Amplification gas
10B4C layer
substrate
4He
7Li
α (4He)
7Li
Fragments need to reach the gas to produce a signal
  thinner layers (~1µm) increase the probability for the fragments to reach the gas
  At least 30 layers (1 µm thick) to reach a good efficiency

## Slide 5: Basic principles of MultiGrid detectors The MultiGrid concept
aluminum samples (blades) are coated on both sides with 10B-enriched B4C layers (~1 µm thick)
15 blades (or more) are mounted on a frame, orthogonally to the incident neutron direction -> grid
1 grid is made of ~60 voxels (2 x 2 x 1 cm3)
-> active area = size of one blade (~10x2 cm²)
neutrons
Grids are stacked in column to increase the sensitive area (and kept insulated from each other)
Superimposed unit cells from different grids form rectangular tubes, mounted with anode wires (20 µm diameter)
-> proportional counters (gas amplification in 1 bar Ar/CO2)
Stack of grids
Rectangular tubes (mounted with 20 µm diameter wires)
Position sensitive detector:
-> coincidences between wire and grid signals
(position resolution = size of a voxel)

## Slide 6: 10B MultiGrid prototypes Previous versions of MultiGrid prototypes
12-grids prototype (2010):
  8x24 cm² active area
  60 anode wires
  optimized for the 2.5 Å neutrons from the CT1 test beam-line at ILL:
  -> 30 layers, 1 µm of 10B4C each
96-grids prototype (2011)
  Stack of 96 grids
  8x200 cm² active area
Several prototypes were tested to show that performances are compatible with 3He standards
  Neutron efficiency @2.5 Å well predicted by MC calculation:
    Calculation: 53.5%
    Measurement: 53.1 %
Neutron efficiency measurement @2.5 Å
Proved the feasibility of larger detectors with the MultiGrid technique
-> first step toward up-scaling the Multi-Grid concept

## Slide 7: 10B MultiGrid prototypes Previous versions of MultiGrid prototypes
The IN6 prototype (2012-2013):
Designed to be mounted on the IN6 time-of-flight spectrometer at ILL (replacing part of the regular 3He tubes)
96 frames arranged as 6 stacks of 16 frames (sensitive area 32 x 48 cm²)
60 anode wires per stack (total 360 wires)
Gas (1 bar Ar/CO2) flushed in the vessel to avoid contamination due to outgasing (detector is not
456 channels to read out the detector
-> a strategy had to be found to reduce this large number of channels
neutrons

## Slide 8: 10B MultiGrid prototypes Toward the reduction of channel number
Reduction of channel number :
  30 wires connected together via resistive networks
  -> charge division allows determination of the targeted wire
    12 resistive networks × 2 ends = 24 readout channels
  3 frames connected together horizontally
    96 / 3 = 32 readout channels
-> 56 readout channels to determine position and depth (reduction by a factor 8)
R
2R
R
R
R
anode wires
Charge Preamplifier
R
R
Charge division card

## Slide 9: The IN6 measurement campaign The IN6 time-of-flight spectrometer
Measurement
  3 weeks measurement with users’ samples
  24 h dedicated beam time
Side by side comparison of performances
Timing information from the IN6 chopper -> much more observables
3He tubes (3 bar, diameter 3.2 cm, length 30 cm)
ToF spectrometer in the range of thermal/cold neutrons (incident wavelengths : 4.1, 4.6, 5.1, 5.6 Å)
Detection system : 337 3He tubes (sensitive area 4 m²)

## Slide 10: The IN6 measurement campaign Efficiency comparison with IN6 tubes
  -> optimized design of the MultiGrid (square tubes, minimized dead space) lead to an efficiency acceptable from the point of view of users
3He tubes have higher intrinsic efficiency
10B Multi-Grid have less dead spaces thanks to optimized geometry
YGarnet @4.1 Å
Si @4.6 Å
rate(10B) / rate(3He) = 1.08
Bragg peaks in the detector -> well defined wavelength
rate(10B) / rate(3He) = 0.97

## Slide 11: The IN6 measurement campaign Gamma sensitivity
Time spectrum NaI
Time spectrum multigrid
3He (gas converter)
10B (solid converter)
gammas and fission fragments well separated
gamma and fission fragment spectra widely overlapping
Multi-Grid detector:
  -> a similar time structure is observed for large values of the high voltage (above 900 V)
  -> no sign of this structure down from 900V
NaI scintillator:
  -> Strong gamma background from IN6 (see spectrum)
  -> well defined time structure
Gamma efficiency measurement
(662 keV 137Cs source)
Gamma efficiency below 10-7 at normal operating voltage was measured (compatible with 3He standards)
α
7Li

## Slide 12: The IN6 measurement campaign Time resolution
Measured ToF depends on the interaction depth in the detector
ToF value of interest is at the entrance window of the detector
  -> ToF value has to be corrected for depth
Wavelength (Å) | FWHM (µs) |  | (10B -3He) / 3He
 | 3He | 10B | 
4,1 | 45,3 | 50,7 | 12,0%
4,6 | 45,9 | 53,0 | 15,5%
5,1 | 57,3 | 65,2 | 13,7%
10B@4.6 Å
δdepth = 1cm → δToF = 8.6 µs (@4.6 Å)
Elastic neutrons

## Slide 13: The IN6 measurement campaign Scattering in aluminum
4.6 Å
z = 6
4.6 Å
z = 12
  24 h dedicated beam time
  Bragg peaks in the detector, well defined wavelength
  2 wavelengths and 2 samples (YGarnet@4.1 Å, Si@4.6 Å)
  Bragg scattering angles :
    4.1 Å : θ = 122°
    4.6 Å : θ = 160°
θ
Cd shielding
  -> Shielding the back of the detector will suppress the bump we observed in time-of-flight spectra

## Slide 14: The IN6 measurement campaign Background
4.4 Hz flat background (30 times higher than IN6’s) was observed (no time structure)
  -> independent of the IN6 instrument
Energy structure of the background (Pulse Height Spectrum measured on wires) :
  significant rate up to more than 3 MeV (excluding neutrons and gammas)
  background energy spectrum overlaps fission fragments spectrum (no threshold discrimination possible)
ToF Spectrum (ZnO@5.1 Å)
What is the source of this background?
PHS measured on wires

## Slide 15: 10B MultiGrid detectors Background
Aluminum is known to be contaminated by alpha emitters (238U, 232Th and daughters)
  Emission rate at the surface: 0.05 to 0.25 cm-2.h-1
  our detector is made of 360 tubes, 192 cm² each
  -> expected total rate between 0.96 Hz and 4.8 Hz (whole detector)
  -> the measured 4.4 Hz background rate is coherent with these values
Two solutions have been investigated with the 12-grid prototype:
  - Ni layer (thickness > 23 µm) deposited at the Al surface
  - ultra-pure Al (U,Th concentration < 1 ppb)
Background PHS on wires
(no B4C coating)
alpha energy
(from MC simulation)
  -> Background reduced by a factor of 20
10 mm

## Slide 16: Toward large area MultiGrid detectors
Good performances of MultiGrid detectors are now well established
Our capacity to up-scale the technology still needs to be demonstrated
-> a large area detector is now under fabrication

## Slide 17: The IN5 demonstrator
IN5 chosen as a benchmark for the construction of a large area demonstrator (largest detector @ILL)
-> same dimensions as an IN5 module (1/12 of the instrument detector) : 3 x 0.8 m²
Challenges :
  Up-scaling means much more grids, blades, B4C coating, readout channels,… (an order of magnitude larger than the largest prototypes ever built)
    Mass production of the parts (grids, blades, inner mechanics,…)
    Optimized design to facilitate mounting
  Operation inside a vacuum chamber: outer pressure variation from vacuum to atmospheric pressure should not affect the detector
    strong pressure vessel required to stand these mechanical constraints

## Slide 18: The IN5 demonstrator
Pressure vessel as large as an IN5 module
8 columns, 128 grids each (1024 grids)with new design (thinner walls)
  less scattering/absorption
34 B4C layers in depth
  More efficiency
544 anode wires, 3m long (~ 1.6 km of 20 µm diameter wire)
1500 readout channels

## Slide 19: Resolution of IN6’s problems Scattering
Backscattering
  2 possible solutions
    Cadmium: blades after (of instead of) the last blade
    Gadolinium: epoxy/Gd2O3 painting
    Cd is heavy, expensive, and toxic -> Gd painting was chosen
    Test measurement at LLB (Saclay)
    -> efficiency of Gd painting as a neutron shielding
    -> determination of the optimum amount of Gd2O3
Scattering in the walls of the vessel (IN6 was shielded with Cd foils)
  large surface -> Gd2O3 painting is more convenient
  Wall painting -> 0.5 mm thick layer successfully tested in a 3m long tube (not in the IN5 vessel…)
    insertion could be forbidden by this additional thickness
  Side of the multi-part grid will be painted directly
Background:
  Combination of Ni-plating and pure Al
    Ni-plating on one part frames (Al is far to expensive)
    Blades: ½ Ni plated, ½ made of pure Al (cost effective solution?)

## Slide 20: The IN5 demonstrator Pressure vessel
Requirements:
  Small dead spaces (strong constraints on mechanical tolerances)
  Thin entrance window for transparency
  Strength against pressure constraints
  -> the last 2 pull in opposite directions!!
Design:
  4 mm thick entrance window
  4 mm thick walls between columns to help standing the pressure constraints
  25 mm thick backplate to hold the grids
  Electron beam welding -> minimum deformation of the parts due to heat
  total weight of the vessel: 350 kg
-> long awaited but now we have it!!

## Slide 21: The IN5 demonstrator Production of the blades @Linköping University
Blades:
  0.5 mm thick aluminum blades
    1/2 from purified Al alloy
    1/2 from standard Al plated with 25 µm Ni
B4C layer coated in the “Samson” PVD reactor at Linköping University :
  3 different coating thicknesses (0.75 µm, 1 µm and 1.5 µm)
  8 week fulltime operation of the Samson deposition machine @Linkoping University
People working hard in Linköping at changing/ packing samples, sandblasting,…
  More than 17000 blades have to be produced

## Slide 22: The IN5 demonstrator Production of the grids @ESS
1024 frames to machine/assemble and mount with blades
Frames:
  512 frames machined from standard aluminum alloy (spark erosion)
  512 frames assembled from several standard parts (good mechanical properties
  Chemical Ni-plating (25 µm Ni layer)
  Gd2O3/epoxy painting to prevent from scattering
    back of the frames
    Side of the multipart frames
Mounting of the grids
  Blades mounted in Lund
  128 grids delivered so far @ILL, 128 more next week

## Slide 23: The IN5 demonstrator Readout
More than 1500 channels for the whole detector, 64 for an IN5 3He module (factor of 25)
Same principles as for the IN6 prototype are applied
  Wires: Charge division readout, 34 channel connected together
  -> only 32 channel (instead of 544)
  Grids: connected 4 by 4 horizontally
  -> “only” 256 channels (instead of 1024)
grids have to be connected across the inner walls of the vessel -> mechanism mounted on backbones pushing some mobile PCBs through slits
32 x 4 grids

## Slide 24: The IN5 demonstrator Mounting procedure of the grid columns
Grids
  128 grids/column
  gap between two grids: 0.5 mm
  -> Spacers between grids to keep them insulated

## Slide 25: The IN5 demonstrator Wire mounting
3 m long, 20 µm diameter W/Th, 80 g tension
544 wires will be inserted in the tubes (1.6 km)
We will probably break some (we bought 5 km)

## Slide 26: The IN5 demonstrator Insertion of a column
First insertion test in the IN5 vessel

## Slide 27: The IN5 demonstrator Mechanical precision: vessel
Mechanical precision is a key parameter
  proper mounting
  good detection performances (gain homogeneity)
Mechanical precision of the vessel guarantee the possibility to insert the column:
    Flatness of the back plate
    Dimension of the tube (only 1.25 mm space between grids an wall)
-> deformation was measured along the columns direction maximum deformation 0.8 mm (over 3m long!!)
Tubes globally respect the specified tolerances (+/- 0.5 mm) except for the edge ones
-> Some simple tools will help guiding the insertion
Chart: mapping of the surface
| 295 | 480 | 663 | 847 | 1037 | 1218 | 1400 | 1588 | 1770 | 1958 | 2140 | 2323 | 2505 | 2690
100 | 0 | 0.18 | 0.4 | 0.5 | 0.55 | 0.645 | 0.655 | 0.645 | 0.585 | 0.51 | 0.44 | 0.31 | 0.17 | -0.005
200 | 0 | 0.17 | 0.42 | 0.485 | 0.54 | 0.685 | 0.72 | 0.715 | 0.64 | 0.52 | 0.485 | 0.385 | 0.19 | 0.075
300 | 0 | 0.19 | 0.455 | 0.63 | 0.63 | 0.77 | 0.77 | 0.7 | 0.66 | 0.625 | 0.55 | 0.38 | 0.22 | 0.05
400 | 0 | 0.245 | 0.56 | 0.68 | 0.71 | 0.85 | 0.85 | 0.855 | 0.83 | 0.825 | 0.725 | 0.575 | 0.385 | 0.235
500 | 0 | 0.255 | 0.555 | 0.68 | 0.66 | 0.765 | 0.785 | 0.765 | 0.725 | 0.72 | 0.59 | 0.41 | 0.18 | 0.07
600 | 0 | 0.255 | 0.535 | 0.625 | 0.595 | 0.72 | 0.715 | 0.71 | 0.67 | 0.575 | 0.48 | 0.33 | 0.12 | -0.065
700 | 0 | 0.16 | 0.37 | 0.47 | 0.515 | 0.6 | 0.58 | 0.58 | 0.56 | 0.425 | 0.32 | 0.16 | -0.05 | -0.225
2.5 mm
1.25 mm

## Slide 28: The IN5 demonstrator Mechanical precision: backbones
Some backbones show bending beyond tolerances (>1mm)
5 mm
x
y
Two cases
  Bending in the y direction-> not a problem (thinner -> more flexible). We can unbend them at ILL.
  Bending in the x direction -> no proper insertion possible
  -> at the moment, we only have 2 “good” backbones
  -> discussion ongoing with the supplier to solve this problem
Expected shape
Actual bended shape

## Slide 29: The IN5 demonstrator The flaking problem
Flaking of the Ni-plating on blades
  Probably due to heating during deposition process (it was never observed on frames)
  Flaking triggered when the blade is inserted in the frame (friction)?
  Possible consequences:
    Efficiency reduced
    Short-circuits between grids
    Discharges in the detector
  Pure Al blades do not flake!!!
  -> they might be used for most of the detector (the coating team is working on it)

## Slide 30: The IN5 demonstrator First tests
Single-column vessel
Mounting of the vessel last Friday (only a few wires)
Gas flushed from Monday
Detector successfully biased up to 1400V (standard operating voltage is only 900-1000V)
-> First neutrons detected 2 day ago!!!!!!!

## Slide 31: The IN5 demonstrator Test campaign at ILL
An IN5 measurement appears to be impossible:
  weight
    Our detector will weigh almost 900 kg
    The mounting rail of IN5 is designed for the 350 kg of its 3He modules
  Gas flow system
    IN5 is not equipped with flanges
  Strength of the vessel under vacuum constraints still to be demonstrated
CT2 beamline (2.5A only)
  3 different thicknesses -> further check reliability of calculations
Other measurements
  Background with the 2 uncoated columns (no need of a neutron shielding)
  Gamma sensitivity
  homogeneity
When we are done, we need to test the vessel strength

## Slide 32: The IN5 demonstrator And more tests
The IN6 measurement highlighted the interest of a pulsed monochromatic beam for characterization
  Scattering
  Time resolution
We need to find such a beam line in a NSS facility

## Slide 33: The IN5 demonstrator Where are we now?
Production phase still going on
  Coating is almost done @Linköping University
  Production of frames and mounting of blades will last for another few months @ESS
Mounting of the detector going on at ILL
  The bending problem need solution soon
We still have a few problems to solve
Our detector can detect neutrons
Many tests will be carried out

## Slide 34: Conclusion
Construction of this large area demonstrator is the final phase of a huge research and engineering program that lasted for more than 5 years
Many and more people have been involved at all stages of its development
  Theoretical studies
  Mechanical design
  Development of the coating process
  Prototyping
  Test campaigns
This last phase requires a strong collaboration between the 3 labs involved, and huge efforts from everybody
  Coating is almost done @Linköping University
  Production of frames and mounting of blades will last for another few months @ESS
  Mounting is going in at ILL as soon as grids are available
The success of this project will show the Neutron Scattering Science community that we have a reliable and performing alternative to 3He, and that we can build it in a large scale
We will make it work!!

## Slide 35: Thank you for attention!!!