## Slide 1: 10B-based detector with stacked macrostructured cathodes and resistive wire readout
Irina Stefanescu, Richard-Hall Wilton, Carina Höglund
ESS AB, Lund, Sweden
Ilario Defendi, Machiel Zee, Karl Zeitelhack
MLZ (FRM2), Garching, Germany
Jens Birch, Lars Hultman
IFM, Linköping University, Sweden
www.europeanspallationsource.se
03/06/14

## Slide 2: The working framework
Main objective: development of alternative technologies to replace the 3He tubes in large-area detectors for neutron scattering applications.
German In-Kind Contribution to the ESS Upgrade Phase (WP 2.3)
FP7 – NMI3-II (Task 21.2)
International Collaboration for Neutron Detector Development
(ICND).
icnd.org

## Slide 3: The detector concept
Stack of MWPCs with Boron-lined “macrostructured” cathodes
3D regular pattern consisting of grooves that can be created in the substrate material by milling, extrusion, forming, etc.
Sense wire planes
mm
α
7Li(α)
α(7Li)
Al
10B
n
εabsorption  d
n
α
εabsorption  d/sin(α)
d‘
Chart: Angle (deg)
| 3.0 | 5.0 | 10.0 | 15.0 | 20.0 | 22.5 | 25.0 | 30.0 | 40.0 | 50.0 | 60.0 | 70.0 | 80.0
Inclined, 1 um, E>120 | 70.622 | 62.206 | 46.738 | 36.763 | 30.397 | 27.77 | 25.736 | 22.587 | 18.137 | 15.802 | 14.021 | 13.123 | 12.721
ε～12%
ε～28%
λ=4.7 Å
x
α=45°
Concept exploits the increase of detection efficiency with angle
22.5°

## Slide 4: Model simulations
GEANT4 simulations
Optimal geometry of the groove to maximize the detection efficiency of the coated plate.
 = 4.7 Å
I. Stefanescu et al., Nuclear Instruments and Methods in Physics Research A727 (2013) 109.
tside = ttop * sin(α/2)
tside = ttop * (x/h)
(uniform coating, ideal)
(coating non-uniform, realistic)
Effects of the conformality of the coating on the detection efficiency.
h
α

## Slide 5: Model simulations
GARFIELD simulations
Electric field inside a MWPC with macrostructured cathodes.
Drift lines and times for electrons.
I. Stefanescu et al., Nuclear Instruments and Methods in Physics Research A727 (2013) 109.
Drift time ≈700 ns (400 ns for flat cathodes)
A-C distance = 6.5 mm.

## Slide 6: Proof-of-concept measurements
=4.7 Å
TREFF@FRM2
I. Stefanescu et al., Nuclear Instruments and Methods in Physics Research A727 (2013) 109.
Test detector, 10 cm x 10 cm active area.
First proof-of-concept measurements with a small size test detector and grooved cathodes made by milling and coated by magnetron sputtering at the Univ. of Linköping, Sweden.

## Slide 7: Efficiency of stacked MWPCs
I. Stefanescu et al., JINST 8, P12003 (2013).
 = 4.7 Å
Flat
Macrostructured
A realistic detector must incorporate several Boron layers in order to be competitive with the He-3 tube.
Detector stack used to test the efficiency of up to 5 MWPCs (10 Boron layers).
The larger efficiency obtained with the stack with macrostructured cathodes could be beneficial especially for instruments using thermal neutrons.

## Slide 8
n
10B4C-coated plates for the demonstrator (Univ. Linköping, Sweden).
Cathode panel, 40 cm x 40 cm (8 plates).
5 cm x 40 cm macrostructured plate fabricated by extrusion (MIFA, Holland).
40 cm x 40 cm active area
Incorporates a stack with 3 layers of macrostructured planes (2 MWPCs) each coated with 1.4 μm 10B4C by magnetron sputtering.
Operated in a continuous flow of Ar/CO2 gas.
Meant to answer questions concerning upscaling and procurement of the various components.
The 10B-demonstrator
Design by Ilario Defendi, FRM2

## Slide 9: The 10B-demonstrator
Frame with resistive wires
Stycast 1266
Conductive glue
Stablohm 875 (resistive) wire Ø=17 μm
wire pitch = 5 mm
anode - cathode distance: 7 mm
72 wires/plane
each wire 40 cm long
resistance wire: ~3 kΩ
Winding tool designed and built in house.
Gluing of anode wires
Stack with 3 MWPCs with macrostructure cathodes and
resistive wires

## Slide 10: Resistive wire readout concept
Concept*:
group the resistive wires in pairs by taking every other wire to create 2 independent, interleaved chains (4 preamp channels, Y-position).
connect all wires through a resistive chain (2 preamp channels, X-position).
Front-end boards, multichannel CSP and shapers delivered by Mesytec.
Aimed position resolution: 5 mm x 5 mm.
A (Y-pos)
B (Y-pos)
Global Group
(X-pos)
B
B
*Thanks to R. Schneider, mesytec.com
A
A
5 mm

## Slide 11: Performance tests with a 252Cf source: charge released in detector
Gas: 1 atm Ar-CO2 80/20
12 anode wires per counter grouped for readout
Readout with Mesytec CSPs (MPR-1), Silena amplifiers, and commercial MCA.
n + 10B  7Li (0.84MeV) + a (1.47MeV) + g (0.48MeV) (93%)
 7Li (1.02MeV) + a (1.78MeV) ( 7%)
Charge released by 1.47 MeV alpha
HV=1680 V

## Slide 12: Performance tests with collimated beam: position resolution
n beam
Photograph of the 10B-demonstrator during the in beam tests at TREFF@FRM2 (λ=4.7 Å).
…
…
W1a
W2a
W2b
40 cm
40 cm
4.7 Å collimated beam delivered by the TREFF instrument @FRM2. Detector placed on a XY moving table.
Group of 4 consecutive wires selected, 2 independent chains created by connecting each second wire (lchain = 80 cm).
Mesytec CSPs + CAMDA multichannel analyzer used to collect and analyze the signals at the ends of each chain.
Ypos = W1(2)a/(W1(2)a+W1(2)b)
W1(2)a, W1(2)b: signals readout at the end of each chain

## Slide 13: Position resolution: preliminary tests with the 10B demonstrator
Beam spot: 2 mm x 2 mm.
Charge distributed to maximum 2 wires, which belong to different chains.
Measured FHWM ~ 4 mm along the wire.
W1a
W2a
W1b
W2b
W1a
W2a
W1b
W2b
Beam ⊥ wire W1b
Beam between wires W2a and W1b

## Slide 14: Coming soon….
Next week Mesytec will deliver the MWPC9 custom-designed frontend boards  6 output signal channels: Y-pos. group 0, Y-pos group 1, X-pos.
16-channel preamplifier (MPR-16) and shaper delivered earlier.
Final set-up ready to be assembled at FRM2.
Performance tests with the 252Cf source planned for July at FRM2.
In-beam tests at the TREFF instrument at FRM2 foreseen for the upcoming reactor cycle (starting in Fall ’14).
Mesytec front-end boards
Mesytec MPR-16 and shaper

## Slide 15: Conclusions and outlook
We propose a concept for a neutron detector based on the 10B-converter, stacked MWPCs with macrostructured cathodes and resistive wire readout.
The concept was extensively studied in GEANT4 and GARFIELD calculations which were validated in a several experimental campaigns with a small area test detector.
Design and construction of a realistic size demonstrator with 2 MWPCs (3 cathodes layers and 2 anode wire plans) almost completed. Intense testing planned with a moderated fission source and collimated neutron beam.
The proposed detector could be an option for instruments that require large-area, moderate position resolution detectors.

## Slide 16: Model simulations
tside = ttop
tside = ttop * sin(α/2)
tside = ttop * (x/h)
(forming after coating)
(coating after grooving, ideal)
(coating after grooving, realistic)
Possible scenarios for the conformality of the coating of a macrostructured plate
h

## Slide 17: Model simulations
h = 5 mm
h = 2.1 mm

## Slide 18: Model simulations
Calculations for 45° grooves, 2 mm deep) and  = 4.7 Å
Chart: Thickness coating top (µm)
| 0.2 | 0.5 | 1.0 | 1.5 | 2.0 | 2.6 | 3.0 | 5.0 | 8.0
Flat | 1.904 | 4.277 | 6.888 | 8.153 | 8.692 | 8.899 | 8.795 | 7.066 | 5.289
side=top | 3.275 | 7.183 | 11.765 | 13.877 | 14.335 | 14.346 | 13.92 | 9.537 | 5.242
side=top*sin(α/2) | 1.719 | 4.006 | 7.136 | 9.317 | 10.812 | 12.21 | 12.847 | 13.069 | 10.939
side=top*(x/h) | 1.959 | 4.302 | 7.357 | 9.594 | 10.92 | 11.78 | 12.031 | 11.428 | 9.14
Results used in a general sense to understand trends in the performance of macrostructured cathodes.

## Slide 19: The 10B-demonstrator
Stack with 3 MWPCs with macrostructure cathodes and
resistive wires
Demonstrator closed and ready to take data
Design by Ilario Defendi, FRM2