## Slide 1: Narrow Slit Smearing on Bonse Hart & vSANS at NIST
March 26th, 2021
John Barker
NIST Center for Neutron Research
Gaithersburg, Maryland
Notes: Fig. 1
Very Small Neutron Scattering Instrument

## Slide 2
Rectangular Slit Smearing: two slit case
For Kratky Camera and more general cases:
Hendricks & Schmidt (1967). Acta Physica Austriaca 26, 97-121
Simplifying assumptions:
Uniform source flux
Uniform detector efficiency
All rectangular: Source, Sample & Detector
Sample slit located at sample
“0D” Detector slit or column of 2D Detector
2D Weighting Function = “Source” * “Sample” * “Detector Element”
P(x,y) = PW(x) PL(y) = “Beam Profile” * “Detector Element”
{ Section along width PW(x) & length PL(y) }
Source
Sample
Beam
Detector
Column
2D Detector

## Slide 3
Kratky
Camera
Three Slit
Camera
More Complicated Cases
Same as two slit case, third slit
Does not define beam.
Sample is not at 2nd slit.

## Slide 4
Source Slit size: X1 x Y1 L1  source to sample distance
Sample Slit size: X2 x Y2 L2  sample to detector distance
Detector size: XD x YD
Beam Profile Trapezoidal
Width: ΔXU = | X1 L2/L1 – X2 (L1 + L2) /L1 |
ΔXL = X1 L2/L1 + X2 (L1 + L2) /L1
Height: ΔYU = | Y1 L2/L1 – Y2 (L1 + L2) /L1 |
ΔYL = Y1 L2/L1 + Y2 (L1 + L2) /L1
ΔYL
ΔXU
ΔXL
ΔYU
Convolution of source and sample slits
Projected slit sizes:
Source Sample Detector
L2/L1
(L1+L2)/L1

## Slide 5
Weighting
Function
qx
qy

## Slide 6
Smear sequentially: 1) wavelength, 2) height & 3) width
{ Wavelength }
{ Slit-Height }
{ Slit-Width }
Can also Combine into single “2D” Resolution Function

## Slide 7
VSANS Narrow slits Option
Source: 150 mm x 2.5 mm L1 = L2 = 24 m
Sample: 25 mm x 1.25 mm width ≈ Height / 100
Detector: 450 mm x 0.3 mm  see graph 
Spherical Particles:
5,000 Å radius
{ ignoring wavelength smearing }
Slit Smeared
Slit Height Function

## Slide 8
“White Beam Option”
• Beryllium filter cuts l < 4 Å
• Cut-off Mirror cuts l > 8 Å
Gain of factor 5 but with
Additional smearing …
“White Beam”
4 Å ≤ l ≤ 8 Å
NVS Dl/l = 12.5%
Graphite
Monochromator
Dl/l = 2 %

## Slide 10
Scattering from Teflon
Beamstop: 6 mm x 300 mm
Detector: 150 mm x 450 mm
Raw CCD image
{ Read bias and speckles not
Removed }

## Slide 11
Neutron Bonse-Hart Calculation (USANS Instrument)
Wavelength Spread Δƛ/ƛ = 6% { Ignore wavelength smearing…}
Width ≈ Height / 5,000 { Ignore width smearing … }
PH(qy) = 1/Δqy for -Δqy ≤ qy ≤ Δqy … Δqy = 0.117 Å-1
{ Detector dominates shape  “Box” }
2.5x10-5 Å-1
Width Function  Rocking curve

## Slide 12
Weighting Array Approximation: W(i,j), 1 ≤ i or j ≤ N
Evaluate slit height integral only at experimental qx,i
Use linear interpolation between qx,i
To cover q-range beyond last data point: q > qi,N
a) Use power law approximation: I(q) = aq-m
b) Leave as numerical integral
x x x x x x x x x x x x x
0 x x x x x x x x x x x x
0 0 x x x x x x x x x x x
0 0 0 x x x x x x x x x x
0 0 0 0 x x x x x x x x x
0 0 0 0 0 x x x x x x x x
0 0 0 0 0 0 x x x x x x x
0 0 0 0 0 0 0 x x x x x x
0 0 0 0 0 0 0 0 x x x x x
0 0 0 0 0 0 0 0 0 x x x x
0 0 0 0 0 0 0 0 0 0 x x x
0 0 0 0 0 0 0 0 0 0 0 x x
0 0 0 0 0 0 0 0 0 0 0 0 x

## Slide 13
Conclusions:
Narrow rectangular slit data smeared in three steps:
i) wavelength ii) height and iii) width
2) Meta Data Needed for Smearing Calculations:
i) Slit, detector and path length dimensions
ii) Pƛ(ƛ) Table for wavelength distribution
iii) PH(qy) Table for height function
iv) PW(qx) Table for width function