## Slide 1: SPL Wire Position Monitor System
Wojciech Żak
CERN/ESS Fellow
with the support of A.Vande Craen and V.Parma, CERN, TE-MSC
5th Open Collaboration Meeting on Superconducting Linacs for High Power Proton Beams
18-19 March 2015

## Slide 2: Outline
SPL Short Cyromodule
Application of the OWPM in the SPL Mock-up
Idea of the Optical Wire Position Monitor
Calibration at warm
Tests at liquid nitrogen
Intermediate conclusions
Fiber optics
Future steps

## Slide 3: SPL Short Cryomodule
DWT
DWT flange
IC support

## Slide 4
OWPM position on the SPL SCM
Stretched wire measuring system interface
Courtesy of CNRS

## Slide 5: SPL SCM Cavity position monitoring specs:
Stretched wire
Opto-coupler sensors
Last cavity support
Courtesy of V. Parma
Static position or slow movements: absolute movements (x,y,z) of each of 4 cavities during steady state operation and cool-down/warm-ups (300-2K)
Vertical range 0-2 mm
Precision < 0.05 mm
Resolution < 0.01 mm
Possibly vibration measures (0-1 kHz)

## Slide 6: How does it work?
wire

## Slide 7
LED and photo-transistor

## Slide 8: Tests setup for room temp
We expect to cover 0.5 mm range with one LED and phototransistor, to fullfill requirement our sensor will be equiped with 4 set
Box to shield the environment light noise,
setup with 2.5mm wire
Wire was move by high-precision, linear stages, providing an accuracy of 1 µm on 150 mm maximum displacement.
wire

## Slide 9: Results
Output voltage signal variations according to the movement of the 2.5mm wire
Courtesy of C. Petrone

## Slide 10: Test setup for LN

## Slide 11: Results of tests in LN (φ2.5 mm wire, no current adjusting: no reference LED and phototransistor)
1.5mm

## Slide 12: Results of tests in LN (φ1.5 mm wire, adjusting current: system equipped with reference LED and phototransistor)
Reference LED and phototransistor

## Slide 13: Voltage drop over time for phototransistors
Chart: temperature [K]
| 293.98500000000001 | 290.68700000000001 | 275.00099999999998 | 244.291 | 204.15799999999999 | 166.595 | 135.161 | 118.202 | 109.497 | 104.173 | 100.65600000000001 | 97.912999999999997 | 95.637 | 93.685000000000002 | 91.980999999999995 | 90.486999999999995 | 89.192999999999998
test 1 | 4.9051 | 4.8965 | 4.8108 | 4.5679 | 4.1777 | 3.6631 | 3.085 | 2.4098 | 1.7797 | 1.2994 | 0.94662 | 0.69209 | 0.50934 | 0.37237 | 0.26935 | 0.10837 | 0.079871
NO linear response and unreproducible behavior with temperature

## Slide 14: Voltage drop over time for LEDs
NO linear response and unreproducible behavior with temperature

## Slide 15: Intermediate conclusions
Despite quite good results at room temperature and steady state conditions in low temperature the are some disadvantages:
No linear reaction with temperature
We’re “blind” during transient period
Properties of phototransistors and LEDs strongly depends on the temperature
Receiver/LED has different behavior even from the same series
We need “Cinderella” to find enough pairs for our sensors (64 in total)
Sensors should be equipped with a temperature sensors
Final application: 64 sensors => 10 wires each => 640 cables
put as much as possible electronic outside the cryostat
temperature independent sensor

## Slide 16: Fibre optic
Couresty of Keyence Corporation
Commercially available product:
Out of the shelf
Relatively inexpensive
Really easy to use
Flexible in configuration
But never:
tested in cryogenic temperature
used for this application
Collimator
Most of the parts have been anodized in black optic

## Slide 17: Results (φ1mm wire, φ1.3mm fiber)
NO change after ~20 thermal cycle

## Slide 18: Results (φ1.3mm wire, fiber φ1.5mm)
0.65mm

## Slide 19: Voltage change with temperature for fibre optic
Chart: Temperature [K]
0.7mV
~1 nm
Temperature INDEPENDENT

## Slide 20: Future steps
Test bigger fibers (φ2 mm internal)
Build the feed thru
Check cross-talk
Procure more components
Try to make a test in helium
Cross-check with laser tracker
Test everything in the SPL mock-up

## Slide 21
Thank you for you attention!
Questions?