## Slide 2: Experience with O2iM operation at ESS¶
ODH Phase 2 PDR
PRESENTED BY Mattias Eriksson

## Slide 3: Agenda
Presentation of hardware and faults during operation 2018-2021
Maintenance perspective
Documentation
Lessons learned

## Slide 4: ODH monitors in operation at ESS
TS2
2 pcs - 02iM monitors, high flow (300ml/min)1port
ACC ODH phase 1
10 pcs -02iM monitors low flow 1 port (150ml/min)
12 pcs -02iM monitors 4 port high flow 150ml/min)
TS 2 and ACC ODH phase 1

## Slide 5: Issues experienced with O2iM monitors
  Monitor sampling port failure due to low pump pressure, 4 port monitors.
    02iM 5
    02iM 6
    02iM 2
    Calibration issues (a-alarm during calibration) monitor 20 (one port)
Summarised issues of the monitors.

## Slide 6: Issues experienced with O2iM monitors
02iM 2,5,6 - We get alerted by the cryo team that one monitor is making regular “beeps” (20s apart).
  When investigating the issue there seems to be a issue with one of the sampling ports that are experiencing a very low pump flow on one specific sampling port). We deactivate that channel and the monitor is behaving ok again.
  Work is scheduled for later and monitor is replaced and sent to Oxigraf for repairs.
  Replacement of solenoid valve/main circuit board and/or pump at Oxigraf.
02iM 2
- During yearly calibration we notice that during the calibration the monitor goes into A alarm when doing the low cal.
Investigation with Oxigraf, and we tried a reset procedure, this did not help. Monitor was replace and sent for maintenance. Main circuit board was updated and issue was resolved.
Explanation of the specific issues with the monitors.

## Slide 7: Maintenance during ACC ODH phase 1.
  Alarm units
  Some minor effect from UV on the units outside
  Yearly maintenance of cleaning is testing
-yearly maintenance of the entire system -6 months replacement of filters in G04 -yearly replacements of the rest of the filters
  Monitors
  Depending on the environment filters need to get exchanged with different intervals
  Yearly/6 months filter replacement
  Yearly maintenance of monitors (cleaning/testing functions/calibration)
  Cabinets
  Depending on the environment filters need to get exchanged with different intervals
  Yearly cabinet filter replacement
  Yearly maintenance of cabinets (cleaning/testing function)

## Slide 8: Maintenance during ACC ODH phase 1.
-yearly maintenance of the entire system -6 months replacement of filters in G04 -yearly replacements of the rest of the filters
  Monitors
  Depending on the environment filters need to get exchanged with different intervals
  Yearly or 6 months filter replacement, re-evaluated after each maintenance.
  Yearly maintenance of monitors (cleaning/testing functions/calibration)
  After each maintenance we are looking through the filters of the monitors and checking them visually how dirty they are and then we re-evaluate if we are doing a 6m or 12m maintenance interval.
  We are working to get a EAM solution for the reports but atm. we are using chess.

## Slide 9: Maintenance during ACC ODH.
Test/Verification/configuration of monitor according to ESS-0189148
Verification test according to ESS-0189148
Documentation process

## Slide 11: Filters
1 year intervall(left) from G04, vs New Filter (right)

## Slide 12: Calibration gases
Calibration gases used: High cal 99.99% O2 in N2 Tolerance +-2% rel. Shelf life 60 month's Low cal 20,70% O2 in N2 Tolerance +-2% rel. Shelf life 120 month's
Test and verification gas 15% O2 in N2 Shelf life 120 month's
Calibration set up and gases used
We are using a setup with a Douglas bag (the monitor draws the gas from the bag) Both for calibration and tests.
Estimated 2 year cost:
15%O2 test gas | 
 | Gasprodukter (rent cylinder)
Gas mixture percentage | 15%O2+N2
Price | 4800 SEK+1 year prepaid - 1750 SEK/cylinder/year alt. Daily - 6 SEK/cylinder/day (default if not agreed otherwise)
Shipping cost | 1230 SEK
Total price | 9530 SEK *Calculated for 2 years rental period *Responsibility of pressure test during rental period is with supplier
20.7%O2 calibration gas | 
 | Gasprodukter (rent cylinder)
Gas mixture percentage | 20.7%O2+N2
Price | 4800 SEK+1 year prepaid - 1750 SEK/cylinder/year alt. Daily - 6 SEK/cylinder/day (default if not agreed otherwise)
Shipping cost | 1230 SEK
Total price | 9530 SEK *Calculated for 2 years rental period *Responsibility of pressure test during rental period is with supplier
99.99%O2 calibration gas | 
 | Gasprodukter (rent cylinder)
Price | 5 450 SEK+ Rent: 1 year prepaid - 1750 SEK/cylinder/year alt. Daily - 6 SEK/cylinder/day (default if not agreed otherwise)
Shipping cost | 1230 SEK
Total price | 10180 SEK *Calculated for 2 years rental period *Responsibility of pressure test during rental period is with supplier
Price summary | 
Total price | 29240 SEK

## Slide 13: Manual calibration accuracy
Chart: 
| monitor 1 | monitor 2 | monitor 3 | monitor 4 | monitor 5 | monitor 6 | monitor 7 | monitor 8 | monitor 9 | monitor 10 | monitor 11 | monitor 12 | monitor 13 | monitor 14 | monitor 15 | monitor 16 | monitor 17 | monitor 18 | monitor 19 | monitor 20 | monitor 21 | monitor 22
Low cal 21,0 | 21.1 | 21.4 | 21.2 | 21.2 | 21.1 | 20.8 | 21.4 | 21 | 21.2 | 21.1 | 21.1 | 21.3 | 21.2 | 21.1 | 21.1 | 21.1 | 21.3 | 21.1 | 21 | 21 | 21
High cal 99,9% | 98.8 | 98.9 | 98.4 | 98.7 | 101.1 | 97.5 | 98.5 | 98.5 | 97.1 | 97.9 | 97.9 | 97.8 | 99.1 | 96.5 | 96.7 | 96.8 | 98 | 96.7 | 99.9 | 96.7 | 96.8
Low cal 21,02 | 20.8 | 20.9 | 21 | 21.1 | 21 | 21 | 20.5 | 21.2 | 21.4 | 21.4 | 21.4 | 21.1 | 21.1 | 20.7 | 21.2 | 21.1 | 20.8 | 20.6 | 20.4 | 20.9 | 21.1 | 21.2
High cal 99,9%2 | 98.7 | 99.1 | 99.2 | 98.1 | 99.5 | 98.6 | 98.4 | 99.1 | 99.9 | 99.5 | 99.5 | 97.2 | 99.7 | 98.4 | 99.6 | 99.7 | 103.4 | 99.5 | 97.5 | 98.7 | 99.8 | 98.7
2020-2021

## Slide 14: Overall experience of the Oxigraf 02iM monitors.
Overall:
Easy to maintain
Easy installation and configuration.
Fast and reliable support from supplier
We have had a few ”incidents” where the system have detected low O2, with very high accuracy and early when there has been a leak.
Reflection from the maintenance perspective:
More attention into filter placement is needed (you got to be able to reach them)
A long test period before you go ”live” with the system is good.
A ”maintenence with filterswap and calibration just befor you go live, so you have a good idea when this was done.

## Slide 15: Finish presentation