## Slide 1
Multi-Blade detector status
Francesco Piscitelli
2022/05/03

## Slide 2
wires X
Ti substrate
neutrons
θ
strips Y
10B4C
q = 5 degrees
A unit
cathode strips
anode wires
holding frame
10B4C substrate
32 Wires
A unit
Unit
front
neutrons
130mm
100mm
Unit
back
Multi-Blade
sample
neutrons
MWPC
Detector Principle

## Slide 3
260mm
130mm
Multi-Blade
AMOR 14 units
Estia 48 units
FREIA 32 units
ALL IDENTICAL
prototypes
Final unit

## Slide 4
Final unit
FRONT
BACK
cathode strips
anode wires
Ti blade coated with 10B4C
U-frame
260mm

## Slide 5
Multi-Blade: TOOLING is complete and tested on MB300L
cathode strips
parts
A unit
6-unit detector
Gluing machine KV-1
Wiring machine K-7
Oven clamp

## Slide 6
We built many prototypes that proved that MultiBlade is ready for science at the ESS. Until now, we have been using the electronics by CAEN, but on the ESS instruments, we will use the VMM because of number of channels.
Estia -> 4600ch
FREIA -> 3000ch
AMOR -> 1300ch
measured
calculated
Efficiency
Cassette shadowing
due to overlap
(not a dead area)
Shadows removed

## Slide 7
Neutron reflectometry measurements, scientific measurements (CRISP @ ISIS)
detector
sample
l1
l2
l3
q
q
Ir sample
Fe/Si sample
Si sample
0.5mm resolution MB detector
2mm resolution detector
conventional 0D detector
Ir sample: Dependance of the q-resolution on detector spatial resolution
Si sample: Collimated vs divergent mode
DIVERGENT
COLLIMATED
COLLIMATED: 120 min measurement
DIVERGENT: 14 min measurement
Fe/Si sample: Off-specular
(from SuperADAM at ILL)
Sample Horizon
B. Correlated roughness
domains from the layers
A. Spin-flip scattering
signal from the layers
G. Mauri et al., “Neutron reflectometry with the multi-blade 10b-based detector,” PRSA: Mathematical, Physical and Engineering Sciences, vol. 474, no. 2216, 2018.

## Slide 8
Angle 1: 1.4deg
Angle 2: 2.2 deg
Angle 3: 3.2deg
NiTi sample
q
q
DIVERGENT
Multi-Blade
Selene guide using the ReFocus principle
sample
detector
Combination of 3 angles
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
Divergence of the beam : 1.6 degrees
(ToF,y)
(q,l)
(q,l)
(q,l)
(q,l)

## Slide 9
Combination of 3 angles
NiTi Multilayer
November 2018 - AMOR reflectometer @ PSI
ESTIA Selene guide and High intensity mode
(q,l)
(q,l)
ESTIA Instrument Proposal
1
1
2
3
4
5
5
4
3
2
Measured
with Multi-Blade
Simulated
0.065Å-1
0.065Å-1

## Slide 10
Tower
MB18 prototype operational during entire 2021 for operations and commissioning.
AMOR decided to have a permanent Multi-Blade detector for the instrument.
MB18

## Slide 11
532mm
330mm
~340mm
This is the front end electronics, there is a box with fans to encapsulate them
neutrons
14 UNIT DETECTOR FOR AMOR, FINAL DETECTOR FOR THE INTRUMENT,
SAME UNITS AS ESTIA AND FREIA, SAME ELECTRONICS
BLADES
Installation Dec 2021

## Slide 12
14-unit detector for AMOR
(active area 260x140mm2)
Same units as Estia and FREIA
Multi-Blade @AMOR
AMOR reflectometer @ PSI
It’s conceptually identical to Estia

## Slide 13
Mechanics, problem with the manufacture of the Titanium blades
  The plan is to buy all the blades for all detectors at once not to end up in manufacturing issues later on
  We are working on this to solve the issue and there are currently 3 paths actively moving forward
  One option is to change from Titanium to Stainless Steel, it has been tested and no issues on neutronic side
  After this is solved the mechanics can be ordered for all instruments
Mechanics

## Slide 16
14 UNIT DETECTOR FOR AMOR

## Slide 17
  One option is to change from Titanium to Stainless Steel, it has been tested and no issues on neutronic side

## Slide 18
Electronics
December 2021, MB AMOR installed with final VMM electronics showed some stability issues
Electronics: the VMM review (25th March 2022) pointed out some issues in the way we connect the VMM electronics on the detector -> problems identified and working on the solutions
MB AMOR will be operational on AMOR (PSI) from 9th of May 2022 using the old CAEN electronics -> any solution about the blades and the electronics with VMMs will be fitted at a later stage once ready.

## Slide 19
Multi-Blade detector is compatible with both VMM electronics and CAEN test electronics :
1 TESTING ELECTRONICS: Athena single ended front-end based on FET and CAEN digitizers V1740D
2 VMM ELECTRONICS: Final ESS solution

## Slide 20
With Athena-CAEN electronics
With VMM electronics
Feed-through

## Slide 21: Readout chain
Detector
Wire and/or grid channels
Connector/adapter
RD51 hybrid
VMM
VMM
FPGA
Power
2V @ 2A
3V @ 200mA
Kintex 7
FMC to HDMI
card, max.
5 hybrids
FMC to SFP+
card
Frontend Assistor card (FEA)
FEA
FEA
LVDS - HDMI cable
2 x 400 Mbit/s (data/config)
Fibre 2 x 6 Gbit/s (data/config)
Backend
Master
module
Ring 1
Ring 12
DAQ software Event Formation Unit (EFU) at DMSC
2x 100 Gbit/s Ethernet (data)
ICS
Ethernet (config)
Master frontends
(MFE)
MFE
MFE
MFE
MFE
Ring n
Switch
1 to n EFUs
1 GB UDP
Slow control
Actual system architecture

## Slide 22
half blade
(contains only wires,
no strips)
half blade
(contains only strips,
no wires)
wires
substrate
=
strips
10B4C
A blade assembly (unit)
1 2 3 4 … 12 13 14 15
1 2 3 … 12 13 14
cassette
blade assembly
…
detector
wires
strips
Master back-end
(12 rings)
R0
R11
R10
R1
R2
Beam
Monitors
0
hybrids
1
2
13
12
FEA 0
3
…
FEA 1
FEA 6
…
Ring 0
Hybrid
VMM 0
VMM 1
AMOR example
N = 14
FEA = Front-End Assistor
Hybrid = one board
with 2 VMMs = 128ch
One cassette 96ch
(32 of hybrid unused)
HDMI
optical

## Slide 23
neutrons
VMMs hybrids
Electronics VMM: enclosure around the VMMs
blades
AMOR detector
14 cassettes = 14 VMMs hybrids
In two rows 7 odd and 7 even

## Slide 24
enclosed
blades
HDMIs
VMMs hybrids
AMOR detector
14 cassettes = 14 VMMs hybrids
In two rows 7 odd and 7 even
neutrons

## Slide 25
Master back-end
Front End Assiter crate (7 FEAs)
blades
HDMIs
HDMIs
HDMIs
optical
AMOR HDMI are 5 m
Tested fine up to 25m

## Slide 26
NOTES:
Direct beam on detector, missing pixels are due to changing noise over time.
wire
strip
wire
strips
No noise
noise

## Slide 27
Electronics: the VMM review (25th March 2022) pointed out some issues in the way we connect the VMM electronics on the detector -> problems identified and working on the solutions
In contact with ILL and ISIS to get inputs
MB AMOR will be operational on AMOR (PSI) from 9th of May 2022 using the old CAEN electronics -> any solution about the blades and the electronics with VMMs will be fitted at a later stage once ready.

## Slide 28
Estia vessel
Estia vessel and detector design ready, preparing CDR

## Slide 29
48-unit Estia detector
(active area 480x260mm2)
neutrons
Top view
Vessel designed and integrated with
Estia detector arm
It is designed to host 2 detectors including
the upgrade detector, but only one installed
on day 1

## Slide 30
70mm
262mm
262mm
El. connections
4 beams
BEAM 2
BEAM 1
TOP MB
UPGRADE TO
FULL SCOPE
BOTTOM MB
DAY 1
70MM
55MM
Design of blades adapted to reduce the dead area between two detectors below 55mm

## Slide 31
U-frame
MB_ESTIA full scope
U-frame
Slotted U-frame
U-frame
Top detector (upgrade)
Bottom detector
50mm
ARRAY PLATE
ARRAY PLATE
Some
fixture
MB ESTIA day 1
U-frame
Slotted U-frame
vessel
vessel
ARRAY PLATE
pins
Electronics connections
COMB
Bottom
detector
Empty space at top

## Slide 32
SUMMARY:
Mechanics, problem with the manufacture of the Titanium blades
  The plan is to buy all the blades for all detectors at once not to end up in manufacturing issues later on
  We are working on this to solve the issue and there are currently 3 paths actively moving forward
  One option is to change from Titanium to Stainless Steel, it has been tested and no issues on neutronic side
  After this is solved the mechanics can be ordered for all instruments
December 2021, MB AMOR installed with final VMM electronics showed some stability issues
Electronics: the VMM review (25th March 2022) pointed out some issues in the way we connect the VMM electronics on the detector -> problems identified and working on the solutions, In contact with ILL and ISIS to get inputs
MB AMOR will be operational on AMOR (PSI) from 9th of May 2022 using the old CAEN electronics -> any solution about the blades and the electronics with VMMs will be fitted at a later stage once ready.
Estia vessel and detector design ready, preparing CDR