## Slide 1: Multi-Blade detector status Oct 2022 Francesco Piscitelli

## Slide 2
Multi-Blade: TOOLING is complete and tested on MB300L
cathode strips
parts
A unit
6-unit detector
Gluing machine KV-1
Wiring machine K-7
Oven clamp

## Slide 3
OCTOBER 2021
Titanium
Previously used dental Titanium, always worked, last year dental Ti not available, switched to Gr2 Ti.

## Slide 4
OCTOBER 2021
Titanium

## Slide 5
14 UNIT DETECTOR FOR AMOR
OCTOBER 2021
Titanium
Either miss delivery or have
A wavy blade detector
We decided to continue and blades
can be replacved later at any time

## Slide 6: Multi-Blade for ESTIA, FREIA installed permanently at AMOR (PSI)
Selene guides
focusing reflectometry
(ESTIA ~ AMOR)
MB18 prototype operational on AMOR during entire 2021 for operations and commissioning.
AMOR (conceptually identical to ESTIA) decided to have a permanent Multi-Blade detector for serving the instrument.
14-unit detector for AMOR
(active area 260x140mm2)
Same units as ESTIA (48 units) and FREIA (32 units)
MB AMOR is operational on AMOR (PSI) from 9th May 2022 using the previously developed CAEN electronics -> any solution about the blades (solved!) and the electronics will be fitted at a later stage once ready.
Thanks to Jochen Stahn (AMOR instrument scientist)
and DG, ICS, and EC/DC groups for the support.
Direct beam 1.7MHz (saturation of CAEN
electronics), no problem up to ~1MHz over
10x10cm2 area
sample
detector
(q,l)
1
1
2
3
4
5
5
4
3
2
MEASURED with Multi-Blade NiTi sample
SIMULATED
0.065Å-1
0.065Å-1
Generally not
visible on
raw data!
Proof of the good
spatial res.
Y-ToF
EFU produced real time image
from ESTIA Instrument Proposal

## Slide 7
  One option is to change from Titanium to Stainless Steel, it has been tested and no issues on neutronic side, NO ACTIVATION ISSUE

## Slide 8
  Also tested in detector with neutrons
  We will have 2 batches of Stainless steel
Adhesion issue
solved
304L mirror
316 2B finish

## Slide 9
Yet another issue: Epoxy tape OUT OF PRODUCTION
Ti blade
Epoxy tape glue
PCB
20 cm rolling direction
26.5 cm
Epoxy tape glue
Ti blade
PCB
FOUND 2 ALTERNATIVES FROM 3M UNDER TESTING

## Slide 10
MECHANICAL ENGINEERING work to finalize MB detector:

## Slide 11
Schedule:
Strip PCB received 140pcs
Wire PCBs ordered, in production
Blades 140 pcs by end of year
Coating on 140pcs completed by Feb 2023
Wiring of first units can start by end of year, then store units
(same for all instruments)
ESTIA vessel manufacture next spring
Quantities:
48 units for ESTIA,
32 for FREIA,
10 TBL,
14 AMOR (replacement)
cathode strips
anode wires
holding frame
10B4C substrate
Ti-blade
Strip PCB
U-frame
4mm gas pipe
Top wire PCB
Bottom wire PCB

## Slide 12
Work on VMM electronics has restarted,
New EMC …
Filters…

## Slide 13
circuits modified according to VMM review recommendations

## Slide 14
1 blade circuit modified according to VMM review recommendations
And tested in a refurbished MB300L at STF