## Slide 1
Advanced Readout Technology in LoKI SANS Instrument at European Spallation Source
Andrea Abba
Davide Raspino
Fabio dos Santos Alves
Erik Schooneveld
Nigel Rhodes
Joseph Hindmarsh
Morten Jagd Christensen
Steven Alcock
Harry Walton
Judith Houston
Francesco Caponio
Luca Pastori
Alberto Cusimano
Yuri Venturini

## Slide 2: LoKI SANS @ ESS
Choppers
Beam
Neutron Guide
Collimator
Sample environment
Detector System

## Slide 3: Loki Readout pipeline
EFU
readout
event formation
Lund - CPH
detectors
'RM'
slits
choppers
samp. env
...
EPICS
FW
'EPICS'
NICOS
data
controls
instrument hardware
site infrastructure
DMSC
filewriter
storage

## Slide 4: LoKI Detector
4 layers of Al tubes
each containing 7 boron-coated straw detectors
Produced by Proportional Technologies INC.

## Slide 5: LoKI Detector
Efficiency: ~50%-60% at LoKI wavelength
Position Resolution: FWHM is ~6 mm up to 350 kHz
Rate capability: 15% rate lost at 2.3 MHz
Loki detector testing at ESS

## Slide 6: LoKI Straw Tube Readout
880 tubes x 7 straws x 256 pixels (binnig) = 1,576,960 pixels
This is a resistive divider network connected to each end of the tubes.
As the wires are also resistive, we get charge division in 2 dimensions; one along the wire, the other which straw tube has detected a neutron.
We need only 2 preamp and digitiser channels on each end of the tube, instead of 7 channels (one for each straw tube).
We need in total 4 preamp channels per tube (7 straw tubes).

## Slide 7: LoKI Straw Tube Readout

## Slide 8: LoKI Preamplfier
-3dB frequency is 32MHz
Gain 6mV/uA (transimpedance)
power dissipation is 25mW/channel (5mA/channel).

## Slide 9: LoKI Detector Banks

## Slide 10: Loki Detector tank

## Slide 11: Loki ADC – R5560

## Slide 12: Loki ADC – R5560
128 Analog Input Channels
R5560:
  Differential Input on low cost RJ45 connector
125 MSPS, 14 bit simultaneous sampling ADC.
4 x 32 LVDS pair or 4x64 single ended I/O (256 I/O)
4x FPGA based on Xilinx Zynq SOC 7035 Xilinx Zynq for a total logic up to 1.4 million of LUT
Embedded Linux installed in SoC SoC
4x 1 Gbps Ethernet interface connected to Zynq PS for fast data readout
8x 10 Gbps optical link (80 Gbps) connected to Zynq PL for extreme fast data readout.
Clock recovery capability from optical Link
Programmable in SciCompiler

## Slide 13: R5560 - DAQ
SFP
2
SFP
1
GBPS ETH
64 DIGITAL IO
8xADC
ZYNQ-7035
Open FPGA & Open SoC
H
LNK
2
H
LNK
1
DDR 3 – 1 Gbyte PS
QSPI
8GB MMC
8xADC
8xADC
8xADC
CLOCK PPLs
4 x 10Gbps
Link
JTAG / DEBUG
4x8channels simultaneous sampling ADC 125 Msps 14 bit
4x10 Gbps link (2 SPF + 2 Internal) distributed on 2 FPGA bank
PLLs for clock recovery
Programmable arbitrary frequency clock generator plugin module
Open source FPGA framework for custom firmware integration & open source SoC software

## Slide 14: R5560 – Internal Interconnections
MODULE 1
32 Channel 128 MSPS
14 bit
Analog Front End
Inputs
Gbps
Eth
SFP 1
SFP 2
32 x LVDS
USB DEBUG
MODULE 2
32 Channel 128 MSPS
14 bit
Analog Front End
Inputs
Gbps
Eth
SFP 1
SFP 2
32 x LVDS
USB DEBUG
MODULE 3
32 Channel 128 MSPS
14 bit
Analog Front End
Inputs
Gbps
Eth
SFP 1
SFP 2
32 x LVDS
USB DEBUG
MODULE 4
32 Channel 128 MSPS
14 bit
Analog Front End
Inputs
Gbps
Eth
SFP 1
SFP 2
32 x LVDS
USB DEBUG
HORIZONTAL LINK
HORIZONTAL LINK
HORIZONTAL LINK

## Slide 15: R5560 – LOKI READOUT BANKS AND RINGS
Readout Backend
R0
R1
R2
R3
R4
R5
R6
R7
R8
R9
R10
R11
BM

## Slide 16: ESS READOUT INFRASTRUCTURE
Ring Topology
COTS Switch
100G
Backend
To DMSC
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
Backend
P-to-P Topology
Ring interface
ICS
Ctrl
ICS Ctrl
&
Timing
data
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
ctrl + data
failsafe ring
Backend
ctrl
Detector Readout for ESS Instruments
Scott Kolya, IKON-13, Sept.2017
Ring interface
x12
x16

## Slide 17: ESS TIMING DISTRIBUTION
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
FE
ASIC/ADC
failsafe ring
Backend
‘time now’
Master FE
Backend
Master FE
FE
FE
Recovered clock is sent through jitter-cleaner IC before forwarding
Acquire accurate timestamp.
Collect digitized (timestamped) neutron data, and downstream it to the BE.
Separate FPGAs must be clocked synchronously for local timestamps to ‘tick’ equally.
Master embeds its clock as 8b10b bitstream. Each slave recovers and re-transmits the clock.

## Slide 18: R5560 – CLOCK RECOVERY AND DISTRIBUTION
SFP1
SFP2
H-LINK
RIGHT
H-LINK
LEFT
7035 + 7045
7030
CLK1
CLK2
CLK1
CLK2
TX+RX (1)
TX+RX (2)
TX
RX
RX
TX+RX (3)
TX+RX (4)
CLK
DISTR
ADC
ADC
ADC
ADC
CLK OUT
MUX_SEL
CLK
DISTR
125 MHz from BASEBOARD
RX+TX
CLOCK PROPAGATION
CLOCK PROPAGATION
JITTER
CLEANER
CLK
DISTR
156.25
ARBITRARY CLOCK MEZZANINE

## Slide 19: R5560 – CLOCK RECOVERY AND DISTRIBUTION
SFP1
SFP2
7045
CLK1
CLK2
CLK1
CLK2
TX+RX (1)
TX
RX
TX
RX
CLK
DISTR
ADC
ADC
ADC
ADC
CLK OUT
MUX_SEL
CLK
DISTR
125 MHz from BASEBOARD
CLOCK PROPAGATION
JITTER
CLEANER
CLK
DISTR
156.25
ARBITRARY CLOCK MEZZANINE

## Slide 20: ESS TIMING DISTRIBUTION
FE
FE
FE
Backend
Master FE
Based on local timestamp, FE’s emit a pulse-per-millisec
Stable, t ~ 10ns, timestamp synchronisation.
50ns worst case with 16 devices in the chain
(Gahl et al. Hardware Aspects, Modularity and Integration of an Event Mode Data Acquisition and Instrument Control for the European Spallation Source (ESS), 2014, arXiv: 1507.01838v1)

## Slide 21: R5560 – LOKI POSITION SENSING FIRMWARE
A and D signals on the same Ethernet cable: twisted pair 1 and 2 or 3 and 4
B and C signals on the same Ethernet cable: twisted pair 1 and 2 or 3 and 4
Σ
Self trigger
Peak
Finder on the Σ
Amplitudes @ peak of Σ
2xIntegration
2xIntegration
1xDifferentiation
1xDifferentiation
A
D
B
C
Aamp
Bamp
Camp
Damp

## Slide 22: R5560 – LOKI POSITION SENSING FIRMWARE
EMMA
1m long straws
HT=1100V
at y=0.3m from the centre
Rate ~400 kHz/1” tube
Raw signals
SIGNAL
DERIVATIVE

## Slide 23: R5560 – LOKI POSITION SENSING FIRMWARE
A+B+C+D
Derivative
2nd Der.
Peak Found
A+B+C+D
Derivative
2nd Der.
Peak Found
1
2
3

## Slide 24: R5560 – LOKI POSITION SENSING FIRMWARE
CHANNEL1
CHANNEL 2
2xIntegration
1xDifferentiation
AFTER INPUT PROCESSING
DELAYED TO CENTER THE PEAK
PEAK CAPTURED
TRIGGER SUM
TRIGGER DIFFERENTIATOR
TRIGGER SIGNAL
DATA PACKET TRANSMISSION

## Slide 25: R5560 – LOKI POSITION SENSE FIRMWARE
LoKI Firmware fully designed by Dr. Davide Raspino @ ISIS, STFC

## Slide 26: R5560 – LOKI POSITION SENSE FIRMWARE
LoKI Firmware fully designed by Dr. Davide Raspino @ ISIS, STFC

## Slide 27: R5560 – LOKI MONITOR AND DEBUG
Real time oscilloscope, spectra and counts monitoring

## Slide 28: R5560 ESS INTERFACE – DATA READOUT
Neutron Data
User
(Master)
FE Firmware
(Slave)
ACLK
TDATA[32]
TVALID
TREADY
Data to backend
A ESS Customized version of SciCompiler Custom Packet is made available to all ESS experiments in order to allow direct data integration to the ESS FEA
LoKI, Miracle, Bifrost, Cspec, Vespa

## Slide 29: R5560 ESS INTERFACE – DATA READOUT
A ESS Customized version of SciCompiler Custom Packet is made available to all ESS experiments in order to allow direct data integration to the ESS FEA
Data Format and packet can be customized to integrate others experiments
BIFROST
DATA FORMAT

## Slide 30: R5560 ESS INTERFACE - CONTROL
User Firmware
FEA Firmware
AXI
Interconnect
Axi-Lite
Timer 1
SciCompiler
Register File
Configuration
registers
0x50000000 –
0x5F420000
Configuration
registers
0x41C10000 –
0x41C1FFFF
W 0x41C0....
Ctrl
( Read returns )
Front End
Backend
INTEG
RATOR
GAIN

## Slide 31: R5560 ESS INTERFACE - CONTROL
Every Memory Mapped Component placed in the SciCompiler design can be access using FEA slow control module

## Slide 32: R5560 ESS INTERFACE - TIMESTAMP
User Firmware
FEA Firmware
Timestamp Counter
AXI-S TIMESTAMP
SciCompiler
TimeStamp Module
PROCESSING
FIRMWARE
TIMESTAMP 56 BIT
TRIGGER
DATA
PACKET
CREATOR
(ESS FORMAT)
AXI-S NEUTRON

## Slide 33: ESS MASTER MODULE
Main requirements:
Data acquisition system common across all ESS instruments.
Integrate ICS timing, ICS EPICS slow control and Data Egress downstream to DMSC.
Output bandwidth 2x100Gb/s.
Data rings to front-end: 12 bidirectional rings (24 fibers).
Ring connection done through daughter boards interfacing with fibre via SFP connectors.
The readout data is sent over Ethernet as UDP payload which can be as long as 8972 bytes long.
The connectivity between the Readout Master and the FENs is achieved through 8B/10B encoded MGT-driven (Multi-Gigabit Transceiver) optical fibres links arranged in a ring topology.
For the slow control interface, the Readout Master instantiates a lightweight UDP stack which communicates with the EPICS IOC (Input Output Controller) using a standard Gigabit Ethernet link.
Based around VCU118 FPGA (UltraScale+ Architecture) hardware from Xilinx.

## Slide 34: R5560 READOUT RACK
ESS Detector Group
Development Setup
One of the 5 Rack with R5560 front-end electronics installed at ESS for LOKI
All R5560 readout module for LOKI Experiment has been provided to ESS and installation in the facility is in progress

## Slide 35: LOKI TEST BEAM
First test beam at ISIS on EMMA
Successful beamtime at ISIS was the first true end-to-end test of the full DAQ chain:
  Straw tubes
  Amplifier
  Readout Electronics and firmware
  FEA firmware
  Master Module
Provides confidence that the ESS Readout System can be used for configuration, time distribution, and data readout of a real Instrument.

## Slide 36: LOKI TEST BEAM
Collected calibration data on the LoKI rear detector using the full ESS software stack on LARMOR (ISIS) SANS

## Slide 37: LOKI TEST BEAM
LARMOR Test Beam Results
Sample measured:
Cd stripped mask
Silver behenate
SDS powder
empty beam
blocked beam
ISIS standard polymer
Silica particles
Vanadium
NeXus file displayed in scipp

## Slide 38: LOKI TEST BEAM
LARMOR Test Beam Results
Sample measured:
Cd stripped mask
Silver behenate
SDS powder
empty beam
blocked beam
ISIS standard polymer
Silica particles
Vanadium
Notes: Here are some pretty pictures of the processed data that we recorded. This is what the scientists had expected, so that are very happy with the performance of the detector system.

## Slide 39: Conclusion
An innovative digitizer has been developed by the collaboration of ESS/ISIS with Nuclear Instruments and Caen
The new digitizer takes the advantage to be reconfigurable with Sci-Compiler and its firmware can be customized for several experiment in ESS/ISIS (LoKI, Miracle, Bifrost, Cspec, Vespa) and other facilities
The firmware can be easy maintained directly by the scientists involved in the development of the detectors
The ESS readout is integrated in Sci-Compiler and used by several detectors now in developing and future ones.
The Loki electronics readout has been commissioned and a scaled model of the full detector up to the data analyses has been validated, including all readout elements from analog frontend to ESS readout interfacing
Notes: Here are some pretty pictures of the processed data that we recorded. This is what the scientists had expected, so that are very happy with the performance of the detector system.