## Slide 2: ESS Target Station Project Update
PRESENTED BY Rikard Linander

## Slide 3: Outline
1 Preamble – Background and Context for the ESS Target Project
2 ESS Target Station Project Current Status
3 Challenges and Issues
4 Achievements and Successes
5 Concluding Remarks

## Slide 4
Preamble Background and Context for the ESS Target Project
1

## Slide 5: Preamble
Ambitions
  Aiming for unprecedented performance in terms of proton beam power, neutron flux, etc
  Exploring new scientific capabilities
  Creating provisions for future scientific ideas
Important conceptual decisions and largely unique features
  5 MW proton beam power capacity
  Long-pulse (2.86 ms) with 14 Hz repetition rate
  Helium-cooled rotating tungsten target
  Low profile moderators (cold and thermal)
  Maximised access to the neutron source for scientific instruments
The contextual background for the ESS target station project
Building on existing knowledge
  Understanding issues and failures encountered for existing peer facilities
  Reusing well-performing and proven technology

## Slide 6: Preamble
proton beam
proton beam window
proton beam instrumentation plug
moderator and reflector
target wheel
neutron beam extraction port
monolith vessel
target monitoring plug
Target Systems
Tungsten target material inside a rotating, disc-shaped pressure vessel
Helium cooling, 3 kg/s, 11 bar, 40 °C inlet/240 °C outlet
Moderators
Provisional locations of moderators above and beneath the target wheel
1st MR exploits the upper space:
  Cold, 30 mm high, liquid H2, 17 K
  Thermal, 30 mm high, H2O, 300 K
Target Safety System
Monitors five critical operating parameters and prohibits proton beam if outside limits determined by the radiation safety analyses
Monolith Systems
Evacuated Monolith Vessel for maintaining inert and neutron transparent conditions
Shielding structures outside and inside the vessel
Access openings for neutron beam extraction and for service/replacement of internals
The realisation of the ESS target station project
Remote handling systems
Large active cells for safe storage and processing of spent radioactive target components
Seven shielded casks for transfer of spent components from monolith to active cells

## Slide 7
2
ESS Target Station Project Current Status

## Slide 8: ESS Target Station Project Current Status
The ESS Target Station is close to be ready to operate with proton beam
  By 2025-11 all target station systems for neutron beam production were fully installed, commissioned, and functionally verified
    Integral testing and commissioning had commenced and are still on-going
    The project were on track for start of operation 2026-03
    However, recent events led to damage our Moderator/Reflector Plug which prohibits operation with proton beam until its replacement component becomes available
  Early 2026 the license to operate with proton beam on target was received
  Meanwhile, during 2026, works continue
    Integral target station operation to gain experience and to train the organisation
    Addressing findings from the testing and functional integration
    Resolving remaining punch list items important for reliable future operation
    Finalising the remote handling systems

## Slide 9: Monolith systems group
Photographs
All cooling systems feed and return pipes connected to their monolith vessel penetrations
Water-cooling systems

## Slide 10: Moderator systems group
Photographs
Installation of the Moderator/Reflector Plug
Cryogenic Moderator System Junction Box
Cryogenic Moderator System Cold Box

## Slide 11: Neutron beam extraction systems group
Photographs
Last Neutron Beam Port Insert being installed together with in-kind partner FZJ

## Slide 12: Target systems group
Photographs
Installation of the Target Wheel
Helium screw compressor
Target drive unit, including the helium rotary union, and helium feed and return lines installed and connected

## Slide 13: Target Safety System
Photographs
TSS venturi tube flow measurement
TSS field control cabinet
TSS control room cabinet
TSS target wheel rotation sensor
TSS helium temperature sensors

## Slide 14: Remote handling systems
Photographs of casks bodies
#1
#2
#3
#4
#5
#7
#9
#3 with hoist

## Slide 15
Challenges and Issues
3

## Slide 16: Weld Quality and Cleanliness
Monolith Inner Shielding
Upon receipt of several monolith inner shielding blocks, issues with both cleanliness and weld quality were found
  Already extensive Factory Acceptance Tests and inspection programmes proved insufficient. Our lesson learned was that even more intensified quality control were needed earlier in the process. (And pandemics do not help)
  The necessary remedial measures were technically straight-forward, while the schedule impact were concerning.

## Slide 17: Helium Circulator Failure
The helium system flowchart

## Slide 18: Helium Circulator Failure
Damages to both circulator assemblies
Damaged lower part of the journal bearing of the rotor in the circulator #1
Damaged upper part of the motor for the circulator #2
  The most likely direct cause of the failure was debris in the test loop coupled with a non-robust design
  Our two-fold remedial action was (A) to recover the failed turbo machines by implementing new motors and bearings of different type and (B) to procure and install an alternative industrial, robust screw compressor technology
  For mitigating future risks, both alternatives were pursued in parallel

## Slide 19: HRU issues
The HRU Function
  Convey helium between stationary piping and rotating target
Challenges
  Complex load pattern
    Mechanical, static/dynamic/cyclic
    Thermal, expansion/high temperatures
  Make possible mounting/dismounting
  Leak-tightness
    Minimising radioactive emissions and He loss
    Minimising internal by-pass
The challenges with the Helium Rotary Union (HRU)
For leak-tightness the HRU has an internal sealing cartridge
  Rubber o-rings
    between parts without relative rotation
  Graphite rings
    between parts with relative rotation
For mounting purpose
  Internal sleeve

## Slide 20: HRU issues
The recovery of the Helium Rotary Union (HRU)
Mitigative actions for the HRU challenges are
  HRU version 1 is pursued for primary use at BoT
    Extensive experience with the component has been gained during cold commissioning
    Spare parts are available
    Technicians have trained refurbishment and can get support from manufacturer
  An alternative HRU version 2 is under development
    Includes lessons learned from HRU#1 through important design modifications
    Can be fitted sealing cartridges with several different technologies, by modular design
    Aim is to have this alternative available as backup at time of BoT

## Slide 21: Damage of Moderator/Reflector Plug
  Cryogenic moderator hydrogen loop was operating normally
  Thermal moderator/pre-moderator water loop was operating normally
  Power transient caused water pump to trip and restart was delayed
  When pump restarted flow very low likely through bypass loop
  Data indicated loss of vacuum in both monolith and cryogenic pipe insulation
  After extraction of the MRP#1 a few days later the damage was obvious and deemed not repairable
  Direct cause is ice formation due to stagnant water coupled with “tight” design with the cryogen H2
  The remedial action is to install the MRP#2, which was already in production, during second half of 2026
2025-11-25 – Target systems were configured for integrated system testing

## Slide 22
Achievements and Successes
4

## Slide 23: Monolith vacuum performance
Hardware configuration
  Vacuum boundary
    Monolith and PBW vessels
    Connection ring
    Neutron beam port tubes
    Neutron beam windows
  Internals
    Neutron beam port inserts and plugs
    All permanent inner shielding
    Feed pipes for water-cooled shielding
Successful achievement of monolith vacuum performance
  Leak tightness confirmed for the vacuum boundary
  No significant contamination of the monolith was found, i.e. sufficient cleanliness
  Excellent repeatability
Monolith cleanliness and vacuum compatibility

## Slide 24: Precision of the neutron beam extraction
As-installed precision for neutron beam extraction
  The precision of the neutron beam extraction relies on the manufacturing and installation results of many components, some large and heavy, others very delicate.
  The ultimate requirement
    Maximum misalignment between each neutron extraction optics assembly and their corresponding optimal moderator surface.
    With reference to each beam line’s theoretical centre line the misalignment should be within ±2 mm (vertical), ±3 mm (lateral) and ±5 mm (axial).
  The successful end-result
    All individual structures were built and installed within their specified local geometrical tolerances.
    Positioning deviation of the neutron beam port inserts, as-installed, were confirmed within a few hundred micrometres, relative to the theoretical moderator position.
    Together with the expected repeatability precision for replacement of the moderator/reflector plug, i.e., ±1 mm (vertical) and ±2 mm (lateral), the required precision will be accomplished!

## Slide 25: Precision of the neutron beam extraction
Maximum deviation from nominal position
Baseplate 0.6 mm (vertical)
Monolith vessel 0.5 mm
Port tubes after welding 2.9 mm (lateral)
Moderator/reflector seat 0.05 mm (vertical) 0.14 mm (horizontal)
Note that average deviation values are generally significantly better.

## Slide 26: Precision of the neutron beam extraction
Maximum deviation from nominal position
Shielding top interface with remote handling system (casks) 0.6 mm
Moderator/reflector plug positioning Expected first time install <0.5 mm
Expected replacement repeatability precision, by remote handling, <1 mm (vertical) and <2 mm (horizontal)
Positioning of the Neutron Beam Port Inserts 0.2 mm (vertical) 0.3 mm (lateral & axial)
Note that average deviation values are generally significantly better.

## Slide 27
Concluding Remarks
5

## Slide 28: ESS Target Station Project Current Status
The ESS Target Station is close to be ready to operate with proton beam
  By 2025-11 all target station systems for neutron beam production were fully installed, commissioned, and functionally verified
    Integral testing and commissioning had commenced and are still on-going
    The project were on track for start of operation 2026-03, but now forecast early 2027
    Due to the damage of the MRP#1, BoT relies on delivery of the MRP#2 which is currently forecast 2026-10
    System modifications for mitigating similar damage to MRP#2 are in work
  Early 2026 the license to operate with proton beam on target was received
  Meanwhile, during 2026, works continue
    Integral target station operation to gain experience and to train the organisation
    Addressing findings from the testing and functional integration
    Resolving remaining punch list items important for reliable future operation
    Finalising the remote handling systems

## Slide 29: Remarks on organisational readiness to operate
  Getting ready to operate stably and efficiently
  Engineered radiation safety features, performance and capacity requirements are verified and validated
  However, full understanding of the neutron factory behaviour requires extended usage and operation experience
  Training and learning on the job will be instrumental for mastering the art of operating the ESS Target Station
Collaboration will be more important than ever before
  Within ESS; with internal functions, interfaces and stakeholders
  Continued support by and exchange with our project partners
  With the worldwide community and peer facilities
Pending installation of MRP#2 we will be ready to start operation with proton beam, but learning to operate stably, efficiently and with expected availability will take time

## Slide 30: Thank you for your attention