## Slide 1
Zhiduo Li
On behalf of CSNS RH&SH group
Institute of High Energy Physics of Chinese Academy of Sciences
Report of Dongguan Branch
Progress on CSNS Remote Handling Upgrade and
Post-Irradiation Examination (PIE) Platform
2026 ICANS XXV in Malmö

## Slide 2
Outline
CSNS Remote Handling(RH) Overview
Current Status of CSNS RH
Progress of CSNS-II RH
R&D on key technology/equipment
Progress of irradiation & PIE platform construction
Summary & Outlook
2

## Slide 3: CSNS overview
The CSNS facility in Dongguan, Guangdong, is an accelerator driven neutron scattering facility for materials research
CSNS operates presently at 180 kW@25Hz proton beam power incident on a solid tungsten target
Equipped with 20 neutron beamline shutters in the target station
Presently 11 neutron instruments are open for operation
Target & Instrument building layout
BL#01
BL#02
BL#16
BL#18
Target building
BL#14
BL#13
3

## Slide 4: Remote Handling(RH) System overview
The main task of remote handling system at CSNS:
  Operation and maintenance of shielding (shutter ) system
  Remote handling for maintenance objects at Neutron&Muon Target Station
CSNS Remote handling mainly includes:
  Major key comp. replacement (e.g.Target, PBW, MR, Shutter Insert, NBW, Ion exchanger and filter, etc.) & process development
  Waste handling operations (Cask loading, shipment, waste volume reduction, waste storage and decommissioning at CSNS campus)
  R&D on key technology/equipment and remote handling for potential weak spot maintenance needs of the facility
  Emergency handling safeguard for abnormal working conditions(e.g. sudden key component failures, etc.)
  PIE application and operations development
  Advanced shielding materials and comp. development
Supporting
Safeguard
Service
3S
4

## Slide 5: Overview- Basic capability & direction of RH
Target station
5
Better Operational Capability
(Load Capacity/Shear Force/Reliability/Durability, etc.)
Faster Diagnostic Capability
(Quick faults diagnostic and response/Leak Monitoring/Material Failure Analysis, etc.)
Finer Observational Capability
(Optical/Electronic/neutron, etc.)
Basic capabilities of CSNS RH:
To Ensure the normal operation and remote maintenance work
To support the development of disciplines such as PIE
To cope with abnormal operating conditions
Development directions of RH: High reliability, high automation, and high intelligence

## Slide 6: Overview-RH goals & general principle
Refer to ALARA /time-distance-shielding principle and CSNS Safety Management Regulations
Safe, reliable, reasonable and effective completion of replacement/disassembly, volume reduction, transfer, storage and decommissioning of radioactive maintenance objects
Target & trolley system: Horizontal handling
MR/PBW/NBW/Shutter insert: Vertical handling
Target station
6

## Slide 7: Overview-RH major equipment & test platforms
PIE test room
adjacent to Hot-cell
Hot cell mockup
at maintenance building
Large comp. handling mockup
at target building (MR,PBW,NBW)
Shutter Insert/MR/PBW/NBW replace device at Highbay
7

## Slide 8: Overview-Key comp. RH replacement history
Targets:3+2
Shutter Inserts:3+8+3
Ion Exchangers:3+3+1
PBW:1
Since CSNS formal operation began in 2018, a significant amount of remote operations have been successfully completed
8
Green: the replacements completed in 2025

## Slide 9: Current Status of CSNS RH-Shutter system
9
Shutter System Operational Summary:
Count: 14 neutron beamline shutters in operation
Runtime: >5,000h of continuous operation in each cycle
Cycles: >6,000 opening/closing cycles in total each year
Status: All interfaces (PPS, spectrometer safety interlocks) and status signals are stable and nominal
Durability testing of the water hydraulic cylinders
Lifetime Assessment of Water Hydraulic Cylinder(WHC): • An durability test platform for the hydraulic cylinders has been built • Successful leak-free operation for over 90,000 cycles has been demonstrated • Testing continues to define service life and prepare for proactive replacement
With a current max. 1,500 cycles/year, The WHCs are expected to withstand an equivalent of full-life based on accelerated wear testing
The shutter system maintains highly efficient and stable operation
Self-developed pure Water Hydraulic Cylinders with radiation resistant work very well and has not been replaced yet
6 imported pure water pressure pumps have been replaced
Replaced 2 rotation motors
Lifetime Assessment of Key Comp. for Shutter Driving System

## Slide 10: Current Status of CSNS RH-Shutter system
10
Remote replacement of the shutter inserts
There are 3 shutter inserts start running at 2018 and currently 14 shutter inserts have been operated
Smoothly completed the replacement of 3 sets of shutter inserts for CSNS-II Instruments in 2025
The last 5 shutter inserts will be replacement this summer
The maintenance and replacement process for shutter insert in radioactive environment has been standardized!

## Slide 11: Current Status of CSNS RH-Target maintenance
Successful replacement of the target modules is accomplished using only remote handling tooling and procedures in Hot cell
So far, CSNS has completed the remote replacement of 5 targets
  Replaced the first Target #1 (No thermocouple) as maintenance panning in 2020
  Urgently replaced the 2nd Target #2 for the seal leakage of target module in May 2022
  Urgently replaced the 3rd Target #3 for the seal leakage of target module in Dec. 2023
  Preventatively replaced the 4th Target #4 by a new torque electric wrench in Feb. 2025
  Urgently replaced the 5th Target #5 due to the inlet leakage of target trolley in Jun 2025
  Typical target replacement and maintenance requires: ~5 days
11
T5 removal
T6 installation & recover
the inlet leakage handling of target trolley

## Slide 12: DN50 Pipe Leakage Maintenance in Target Trolley
The water inlet pipe joint of the target trolley was leaked due to defects in the sealing ring.
Replacing the seal ring and performing a hydraulic test to complete an emergency repair takes 26 days-one week ahead of schedule
Follow-up plan : In 2027, during the long shutdown period of CSNS-II upgrade, the trolley pipe will be reformed, and the existing inlet and outlet pipeline joints will be changed to welding methods to improve reliability.
12
water inlet pipe leakage
Replacement process of sealing ring of inlet pipe joint
Observation after replacement of outlet pipe sealing ring

## Slide 13: DN50 Pipe Leakage Maintenance in Target Trolley
  Main maintenance processes for emergency repairs of DN50 inlet pipe joint leakage
  Addressed a water loss alarm with an emergency shutdown, preliminary diagnosis, dose measurement, contingency planning, fabrication of temporary shields/tools, target disassembly, replacement of two sealing rings, successful testing, and system restoration, achieving a full recovery.
13
Remote observation in hot cell

## Slide 14: Summary of DN50 Pipe Leakage Maintenance
Lessons and experience
The better design of remote handling is very important at the beginning, as it determines the success, failure, and quality of the maintenance
Quick fault diagnosis is very important when a failure occurs
Quickly develop short-term and long-term maintenance plans based on the existing reserve of remote handling technology and capabilities
Key maintenance procedures require extensive simulation exercises and rapid optimization validation in the mockup
High efficient collaboration of various internal and external resources
Strengthen the reserve of preventive maintenance technology for weak points
Difficulties and challenges
Extreme narrow limited space, high dose rate
Only manual operation
Tight project schedule
The integrity of the joints and threads is uncertain
After shielding
Seal ring replacement test in mockup
14

## Slide 15: Ion Exchanger Replacement for Cooling Water
15
Main maintenance experience
Safely and efficiently completed 7 sets of ion exchangers replacement ( 2020&2024&2025)
Strengthen gas tritium protection and monitoring. When the metal hose is disassembled, a short time (about 10s) of gas tritium produces 400-500Bq/L
Maximum radiation dose at the operation position: 14.5 μSv/h; Tritium-in-air monitoring: <460 Bq/L

## Slide 16: Progress of CSNS-II Remote Handling System
CSNS-II RH overall goal:
Comprehensively strengthen the building and improvement of remote handling capabilities!
Completed the RHS detail & engineering design and carried out the upgrade implementation as system CPM plan
The RH system upgrade mainly includes:
Hot cell upgrade
RH design and upgrade of key comp. in target station
  TMR/PBW/NBW +muon target station RHS construction
  Shutter system, water cooling purification equipment, etc.
Waste storage/basement space upgrade & construction
  Storage of retired components such as TMR and PBWs
  Decommissioned ion exchanger reduce volume, storage, etc.
Upgrade and improve the handling capabilities under abnormal & emergency working conditions for target station maintenance objects
R&D on Key technology and equipment
PIE capacity building for retired core components
16

## Slide 17: Hot Cell (HC) upgrade
All the stainless steel liner upgrade of HC
In order to better improve the sealing level and decontamination capabilities of the hot cell for irradiated sample preparation & PIE
Completed upgrade construction of the remaining stainless steel liner on the inner wall of hot cell within 60 days in summer maintenance period of 2025
17
After the hot cell upgrade
Before the hot cell upgrade
Upgrade construction of the remaining stainless steel liner
High reliable crane upgrade
Upgraded to a new 10-ton capacity crane with dual-redundant drive and emergency towing functions, aimed at enhancing remote operation reliability, operational precision and surface decontamination capability
Completed the HC crane process from design finalization to factory acceptance within 45 days
After the new crane upgrade
Power Manipulator reliability upgrade
Safely and efficiently completed the relocation task of the power manipulator in HC, providing the necessary conditions for the HC liner upgrade
upgrade of the manipulator's redundant drive system of CSNS-II will be completed this summer to enhance the reliability of HC maintenance
Power manipulator upgrade
X/Y/Z axis redundant drive upgrade
Comprehensively improve the reliability of remote handling and its capability to cope with abnormal operating conditions!

## Slide 18: MR replacement process development
MR plug is large, complex and limited-lifetime(10y@CSNS-I, 5y@CSNS-II) key component of the target station
Main RH design features: Overall installation + Sectional disassembly
MR plug split into four segments and its pipeline split into three segments, each segment will be disposed separately
Use lifting tools and transfer vessels to transport plugin and pipeline to storage vessels in the hot cell, and then lift them into the basement to store
High radioactive MR replacement is full of challenging: High reliable handling equip. & processes, high dose & gas tritium protection, etc.
The optimization of MR maintenance processes and experimental testing are ongoing
18
MR replacement plan: April 2028
MR maintenance test in mockup

## Slide 19: MR replacement process development
Completed the design & development of main process and device for MR remote replacement
MR replacement process mainly consists of ten steps
Except for the lifting and transportation of plugins, all other steps have been tested on the testing platform
19
RH testing platform

## Slide 20: MR maintenance equipment for CSNS-II
Completed the engineering design and manufacturing review of the CSNS-II MR maintenance equip.
Factory acceptance will be completed by the end of 2026
MR Hoisting and
Transfer Device
Lower Section MR
Storage Vessel
Lower Part of the Middle Section MR Storage Vessel
Upper Part of the Middle Section MR Storage Vessel
MR Pipeline Storage Vessel
Vessel Lifting Tool
High-Torque Teleoperated Wrench
Low-Torque Teleoperated Wrench
20

## Slide 21: Proton Beam Window(PBW) Replacement
PBW component is designed for vertical maintenance, including PBW module, inflatable bellows sealing, wire target & cable, main frame, cooling water pipeline, vacuum pipeline, shielding iron and quick sealing structure
Propose the replacement task and confirm the interface at the target station in 2022
Completed maintenance related tooling / mock-up tests and operational exercise in 2023
Successfully completed the first PBW 1# replacement during the summer maintenance of 2024
Old PBW 1# @A5083-O
Designed by 100kW
Single layer window & edge water-cooled
New PBW 2# @A5083-O
Designed by 500kW
Double-layer window & center water-cooled
PBW module
21

## Slide 22: Proton Beam Window(PBW) Replacement
Remote replacement of PBW module:
Handling difficulties: Shearing/cut of utility pipes and cables
Key Points: Ensure the PBW sealing surfaces intact and the vacuum degree at the accelerator side < 10-4Pa
Standardized remote replacement process will be achieved after further improvement
The monitoring radiation dose level is equivalent to the theoretical calculation
The protection and monitoring of gas tritium is very important
Gas tritium in inflatable seal bellows: Max.3520Bq/L
22

## Slide 23: Basement waste room upgrade
The ion exchanger and filter are stored separately to improve the utilization of storage space
Develop a special crane lifting system to increase lifting load and accuracy for remote operation
Completed design and manufacturing review of this crane upgrade system
During this summer maintenance period, the on-site installation of the new crane will be completed
Upgrade content: Improvement of waste maintenance and storage capabilities in TS basement
Special lifting device
23
Target, The bottom section of MR plug, and PBW
Solid waste storage room for high-level waste:

## Slide 24: Muon system remote handling
Based on the operation status of the device and radiation dose, remote maintenance includes two parts: regular maintenance and fault maintenance.
The muon target has a fixed service life and is a subject that requires regular maintenance.
Other equipment is of long-lasting design and can be remotely maintained in case of failure.
The Muon target is designed with a fixed storage area that can accommodate 8 targets. For other remote maintenance equipment in the case of failures, we will transfer it with storage container to the solid pollutant storage area of CSNS.
Remote maintenance work includes pipe cutting, remote hoisting, transportation and storage.
24
On-site storage pits for old targets

## Slide 25: Muon system remote handling
Main Process of Maintenance the Muon Target
Drain the cooling water and replacement of gas in the sealed container, remove the top shielding body;
Remote control operation to lift the shielding body into the storage container;
Remote operation of cutting water pipes and gas pipes;
Set the transfer container, the two-position displacement table and the remote observation system;
Remote-controlled threaded gripper lifted the muon target into the transfer container;
Transfer the muon target to the fixed storage area and move it into the storage container.
25

## Slide 26: R&D on key equip. for special working conditions
Development of automatic cutting and welding pipe device for confined spaces of DN50/DN32 pipe joints in target trolley
Technical difficulties:
The cutting and wire filling welding process generates the solid waste, making collection a necessary consideration
Repair space is constrained within a 30-80mm range, which severely limits accessibility
Complete the weld in one pass with 100% quality, avoiding reworking
Consider the synergy of parameters such as observation, positioning & operating time during the process
The development of the automatic cutting groove and welding tools has been initially completed, and the prototype test results are OK!
It will provide key technical support & guidance for the target base & pipe replacement in target trolley during the long shutdown in 2027.
26
To be repaired 5 SS pipes: 2*DN50+3 *DN32
Solution progress:
A minimum radial space of 40mm, adaptable for automatic cutting& welding of DN50/DN32 pipes, and the cutting & welding tools had pass test for DN50/DN32 pipe.
Manual cutting and welding tests for narrow space are also carried out in parallel.

## Slide 27: R&D on key equipment for abnormal conditions
Development of cleaning and polishing device for sealing surface of Target Base and PBW
The device's factory acceptance and the verification of key process tests have been completed.
Finished the cleaning test of surface by remote
To explore adaptive surface polishing materials & the process for high-precision mirror surface
Cleaning and polishing device for PBW mirror surface flange
Cleaning and polishing device for sealing surface of Target Base
Three-wheel pneumatic rotation
27

## Slide 28: R&D on key equipment for abnormal conditions
Helium Leak Detection Trolley
Narrow Space-Operational Area
Remote-Operated Welding
Development of a remote-operated helium leak detection system
Addresses the challenge of remote operation within the confined spaces of outer reflector welding seam area in the target trolley's channels
Developed a rail-mounted helium leak detection system and corresponding welding technology with a master-slave operation mode
Solves the challenges of welding tool accessibility and reliability in restricted high radioactive areas
The repair welding system is now ready for formal prototype development
trolley's channels
Outer reflector
fillet welds area
Challenge: Remote operation leak detection + remote operation repair welding in high-level radioactive restricted space
28
durability testing in mockup hot-cell
Testing of the helium detection system in hot cell will be carried out this summer

## Slide 29: R&D on key equipment
Large-scale self-developed water pressure
heavy-duty mobile manipulator
Medium-scale detection - decontamination robot
Small-scale detection - decontamination robot
Development of a series of manipulators & robots to improve RH capabilities
Complete these developments by the end of this year
Force-Controlled Servo Master-Slave Manipulator
Small Sample Sorting System
Focus on: Radiation resistance, high reliability, high precision, and intelligence.
29

## Slide 30: R&D on key technologies
Development of Efficient Intelligent Operation and Maintenance System Integrating New Technologies
AI-Based Remote Operation Assistance System for Hot Cell
VR-Based Hot Cell Simulation Training Platform
Intelligent Health Management System for Shutter HPS
30

## Slide 31: Progress of irradiation & PIE platform construction
SNIF(Spallation Neutron Irradiation Facility)-I
SNIF-II
A - High Irradiation Dose Area
(1013-1014) n/cm2/s
D-E Low Irradiation Dose Area
(1010-1011) n/cm2/s
The annual cumulative fluence of high-energy neutrons: 1020/cm2@>1MeV
the irradiation temperature: 30~450oC @A area;30 ~70oC @D-E area
Samples were first placed in the D‑E zone for irradiation in Feb. 2024
A zone on the T7 target first began irradiation of samples in Oct. 2025
Based on CSNS operating target
Based on CSNS-II muon target station beam stop
Proton beam
Irradiation plug-in and sample chamber
Neutron flux: ~3×1012n/cm2/s
Irradiation temperature: High & low environment
The preliminary design is currently underway
Sample irradiation will be carried out in May 2028
31

## Slide 32
Progress of PIE Platform
The existing hot cell of CSNS has the basic sample preparation function for PIE samples and can also conduct macroscopic mechanical property and thermal conductivity tests.
The PIE platform will be completed by the end of 2026 and the environmental assessment qualification will be obtained
Microstructure Characterization
(Dedicated FIB‑SEM)
Automatic sample changer for neutron diffraction of radioactive samples
PIE Sample Testing & Characterization Construction
Automatic optical microscope
Mechanical property test platform
DIC test system
Laser thermal conductivity analyzer
Existing PIE equipment
New PIE equipment added in 2026
Sample preparation equip. in hot cell
PIE lab. in Low-Level Waste Hall of CSNS-II
CSNS Hot-cell & PIE Sample Preparation
FIB-SEM
Impact testing machine, microhardness tester
Automatic Neutron diffraction & small-angle experimental device for irradiated samples
Configuration of test environment equipment

## Slide 33: Summary & Outlook
RH system maintained efficient and stable operation, successfully completed the maintenance and replacement of some key radioactive components including emergency repairs, and accumulated valuable remote handling experience.
Currently, The CSNS-II RH upgrade work, as well as the development of key technology/equipment and the construction of PIE platform, are proceeding smoothly according to the plan.
Hope to strength international cooperation and jointly enhance the remote handling and PIE capabilities of the international spallation neutron source, such as by establishing special seminars or cooperative organizations.
33

## Slide 34
Thank you for your attention!
CSNS REMOTE HANDLING TEAM
34