## Slide 1: Future 8 MeV Upgrade for the LvB Compact Accelerator-driven Neutron Source
Mina Akhyani, Ha Shuai, Ferenc Mezei

## Slide 2: AGENDA
LvB project history and now
LvB upgrade plan
Performance estimate in the upgrade
Brightify implementation
Conclusions

## Slide 3: LvB compact source project
Started at 2018
Purpose of the LvB project
  Meet the neutron demand of Mirrotron Ltd for supermirror and other instrument testing
  Provide neutron beam time for other industrial companies
TMR
ECR
RFQ 3 m
70o beam line
2.5 MeV
35 keV
BNCT
DTL 5.5 m
8 MeV
Upgrade area

## Slide 4: LvB Accelerator
Accelerator Parameters |  | 
Ion Source | Mode | Pulsed beam
 | Pulse width | 1.25 ms
 | Output energy | 35 keV
 | Pulse frequency | 40 Hz
 | Peak current | 20 mA
RFQ | Output energy | 2.5 MeV
 | RF frequency | 200 MHz
 | Duty factor | 5%

## Slide 5: Lvb current tmr assembly
TMR specifications |  | 
Target | Material | Lithium
 | Geometry | V-shaped 15 deg wrt beam
 | Dimensions | 3,5*3,5 cm2
 | Thickness | 50 µm
Pre- Moderator | Material | Water (300 K)
 | Wall Thickness | 1.5 cm
 | Length | 21 cm
 | External Cross-section | 8,2*9,9 cm2
Cold Moderator | Material | Para-H2 (20 K)
 | Cross-section dimensions | 1,5*3 cm2
 | Length | 15 cm
Reflector | Material | Lead
 | External Dimensions | 60*60*60 cm3
Side view
top view

## Slide 6: Upgrade plan: which parts of TMR must/could be modified?
Target (MUST)
Replace existing solid Li target
Goal: enable high neutron yield and long target lifetime at 8 MeV proton energy (blistering, heat load)
Moderator (COULD)
Adjust dimensions
Optimize location
Reflector (COULD)
Adjust dimensions

## Slide 7: Target optimization
Heat load requirement:
~500 W/cm² -> Cross-section: 4 × 4 cm²
*IAEA, 2023. Advances in boron neutron capture therapy. IAEAL 23-01601.
~530 µm

## Slide 8: TMR performance
Brightness: Figure-of-Merit
  Phase-space density
Brightness is calculated by T-cross, using angular current within 2 degrees.
Void cell is added to the exit surface of the moderator (9.9 cm * 8.2 cm)
Energy range : 2-80 meV (1-6 Å)
T-cross tally
Exit surface of the moderator

## Slide 9: TMR performance
Target | Proton energy [MeV] | Be layer thickness [µm] | Moderator length [cm] | Reflector length [cm] | Brightness*1e-8 [1/cm2/sr/source] | Relative error
Li | 2.5 | - | 15 | 60 | 0.4 | 0.04
Be-Li | 8 | 400 | 15 | 60 | 3.8 | 0.03
Proton energy: 2.5  8 MeV
Li target  Be-Li target
9X better brightness

## Slide 10: Upgrade plan: which parts of TMR must/could be modified?
Target (MUST)
Moderator (COULD)
Adjust dimensions
Reflector (COULD)
Adjust dimensions
Reflector Length
Moderator Length

## Slide 11: Reflector optimization
Highest brightness: 2*2*2 m3 lead
BUT:
Weight: 100t  Too heavy for CANS
Optimal economical choice:
1,2*1,2*1,2 m3
max
opt

## Slide 12: Moderator optimization
Due to cryogenic costs, smaller moderator would be a better choice.
3X more cryostat power is required compared to the current TMR.
Optimal moderator length choice:
15-16 cm

## Slide 13: TMR performance
Target | Proton energy [MeV] | Be layer thickness [µm] | Moderator length [cm] | Reflector length [cm] | Brightness*1e-8 [1/cm2/sr/source] | Relative error
Li | 2.5 | - | 15 | 60 | 0.4 | 0.04
Be-Li | 8 | 400 | 15 | 60 | 3.8 | 0.03
Be-Li | 8 | 400 | 15 | 120 | 5.2 | 0.03
Proton energy: 2.5  8 MeV
Li target  Be-Li target
Reflector size increased by 8X
13X better brightness

## Slide 14: Is brightness calculation accurate?

## Slide 15: Is brightness calculation accurate?
Current Method
  Integrates over the entire moderator exit surface
  Uses same phase-space volume for both instruments
  Assumes normal direction for brightness calculation
Caveat
  Guides accept only part of the moderator exit surface
  Each instrument has its own requirements
  Both instruments cannot simultaneously access the normal direction
9 degrees

## Slide 16: Improving the Brightness Calculation: Brightify*
  Enables flexible selection of phase space volume
  Allows brightness evaluation for arbitrary emission directions, not limited to the surface normal
  Introduces an additional dimension –direction- into the brightness map.
Impact:
Supports optimized guide design and placement
* https://doi.org/10.1016/j.nima.2025.171037

## Slide 17: Brightify Implementation: step-by-step
  1. Phase space definition
    Calculate the phase space volume on the moderator surface analytically (using COFSI*)
    This defines the resolution of the brightness map
  2. Directional optimization
    Use brightify to identify optimal emission direction for cold brightness (highly directional)
  3. Consistent direction application
    Apply the same optimized direction to: Thermal and bi-spectral brightness maps
*https://doi.org/10.1016/j.nima.2023.168643

## Slide 18: Instrument-specific brightness
Instrument | Phase-space volume | (1) Cold Brightness 2-20 meV [n/cm2/source/sr] *1e-8 | (2) Thermal 20-80 meV [n/cm2/source/sr] *1e-8 | (3) Bispectral 2-80 meV [n/cm2/source/sr] *1e-8
Reflectometer | 5.4*6.1 cm2 2 degrees | 4.7 | 3.9 | 8.5
Diffractometer | 4.4*6.3 cm2 2 degrees | 5.45 | 4.0 | 9.5
1
2
3

## Slide 19: What if the instrument is not in the normal direction to the moderator surface?
Top view
Moderator exit surface
guides

## Slide 20: Final tips & takeaways
LvB upgrade
~13X performance gain by:
  Proton energy increased:
  2.5 MeV  8 MeV
  Target design improved:
    Li  Li-Be
  Reflector volume increased by a factor of 8
Improved brightness evaluation
Brightify implemented for more realistic, instrument-relevant brightness estimation.
Deviation from optimal direction results in:
  ~15% loss at 4.5°
  Up to 50% loss at 15°
How can we better arrange the guides around a single moderator to without sacrificing this hard-won brightness?

## Slide 21: THANK YOU
Mina Akhyani
m.Akhyani@fz-Juelich.de
https://github.com/BrightnessTools/Brightify