## Slide 1
Ping Wang
NTDS Team
wangping@ihep.ac.cn
The first time-of-flight thermal and cold neutron test beamline based on a spallation neutron source in China
― Neutron Technology Development Station (NTDS)

## Slide 2
Outline
Overview
Design
Installation and commissioning
Summary

## Slide 3: 1 Overview
Beam time allocation
5000 Hours/Year
Application classification
Requirements of a new beamline for China spallation neutron source (CSNS)
1. Numerous neutron device’s neutron beam testing tasks of the CSNS II project
2. Serve as a key platform to advance neutron technology development (such as materials, devices, methodologies, and metrology) for CSNS

## Slide 4: 1.1 Design highlights
Terminal 3 Multi-function
Terminal 2
Neutron reflectivity
Terminal 1
Single-crystal orientation
1. Versatile functions: three terminals
2. Lower background: Bender + T0 chopper, long flight path
3. Higher beamtime utilization efficiency: automatic technology
Wavelength: 0.5~20 Å (0.2~300 meV)
Bandwidth of TOF: 5.4 Å
Beam Size: 25mm(W) × 50mm(H)

## Slide 5: 1.2 Schedule and organization

## Slide 6: 2 Design
Main components:
Choppers: 3 sets T0+Single disk T1+Double disk T2
Supermirror neutron guide Bender+ Straight
Second shutter
Neutron slits: 4 sets
Single-crystal orientation terminal Rotating sample stage, CCD Detector
The reflectivity terminal Rotating sample stage, 8mm 3He Detector
Shielding：Carbon steel, Boron-containing polyethylene, Concrete

## Slide 7: 2.1 Neutron guide system
Beam
Neutron wavelength spectrum on guide exit
Neutron spacial homogeneity
Wavelength
Flux
Bender: to separate from Beam line 8 and remove fast neutrons and gamma rays background
Straight guide: to efficiently transport neutrons

## Slide 8: 2.2 Neutron chopper system
T1(7.5m): to reduce frame overlap background
T0(13.7m): to block the prompt high energy particles
T2(14.7m): to select the desired bandwidth neutrons
The prompt fast neutron background is suppressed by two orders of magnitude
Neutron chopping time and space diagram
Wavelength (Å)
Degree
Bandwidth (5.4Å)

## Slide 9: 2.3 Shielding
Reconstruction based on the BL8 shielding
Now
The dose rate is less than 2.5 μSv/h@500kW
Neutron shielding calculations

## Slide 10: 2.4 Terminal 1: Single-crystal orientation
Neutron wavelength and bandwidth: 0.85~6.25Å，5.4 Å
Beam size: 0.1×0.1mm2~25×50mm2
Alignment precision: <0.5°
Sample size: ≥ 0.5mm3
Detector: CCD camera (200×200mm2),0.2mm each pixel
Neutron flux at sample: 7.3×106n/s/cm2@160kW
Incident beam
Slit
Detector
Beam stop
Scattering beam
Sample
Requirements from instruments
Specific inverted lattice faces and directions
Quickly check the quality of single crystal samples
Limited by the detectors, can not be carried out on the spectrometer
By Yuhua Ma

## Slide 11: 2.5 Terminal 2: Neutron reflectivity
Neutron wavelength and bandwidth: 0.85~6.25Å，5.4 Å
Beam size: 0.1×0.1mm2~25×50mm2
Q range: 0.005~0.22Å-1（m=0.2~10）
Best Resolution Q/Q：~1%
Detector: 4*He3 tube,8mm×300mm
Neutron flux at sample: 3.8×106n/s/cm2@160kW
Incident beam
Slit1
Detector
Reflecting beam
Sample
Slit2
Testing platform for neutron supermirror development of neutron optics
By Qing Zhang

## Slide 12: 3 Installation and Commissioning
Other components (choppers, second shutter, reflectivity terminal, etc. ) were installed in Oct. 2025
The neutron guides and shielding were installed in Sep. 2024
Operation area
Test Station
Shielding
Neutron Guide

## Slide 13: 3.1 The first neutron beam
NTDS successfully achieved its first neutron beam on March 5, 2026
The neutron flux: 4.2×106 n/(cm2*s) at 0.5 ~ 5.9 Å by choppers

## Slide 14: 3.2 Neutron beam performance
Time (ms)
Flux
High energy particles were suppressed by T0 chopper
1~6.4Å
Wavelength (Å)
Flux
Comparison of overlapping neutron wavelength frames by bandwidth choppers
Beam spatial distribution by the neutron beam monitor

## Slide 15: 3.3 Neutron reflectivity commissioning
Standard nickel film sample
The difference between the experimental and theoretical Qc values is 5% due to the thin film density being lower than that of the bulk material
Standard m=3 supermirror sample
It seems better to use a large angle of reflection with long wavelength neutrons for better consistency in the reflectivity data, and the commissioning is still going on

## Slide 16: 4. Summary
NTDS has completed most of its project construction tasks, and it is the first beamline to achieve neutrons in the CSNS II project
NTDS is on commissioning and operation, which will strongly support the numerous neutron beam testing tasks of the CSNS II project
NTDS will play an important role for the neutron technology development of CSNS in the future

## Slide 17
Thanks for your attention!