## Slide 1: Supporting ESS operations and commissioning with "spallation chemistry"
Monika Hartl, Katrin Michel, Patric Lindqvist Reis, Hossein Sina, Harald Schneider, Nicholas Weisend, Michael Wohlmuther, Günter Muhrer
European Spallation Source Eric
ICANS XXV, Malmö, April 13-17th, 2026

## Slide 2: How everything started in 2022
Spallation Chemistry Team
SMEs from Spallation Physics Group and the Chemistry and Life Science Group formed 2022-2025 the ESS-wide “Spallation Chemistry Team” collaboration
  4 chemists, 2 physicists, 1 material scientist, 1 engineer – 50% of team was experienced from previous spallation/neutron sources with the goal to:
  avoid duplicated requests for assistance & measurements
  avoid requests asking for a sledgehammer to crack a nut
  create searchable knowledge base to be able to understand what is going on
Main Topics in 2022:
(1) Concrete analysis (shielding for accelerator tunnel, bunker):
  assuring no unwanted trace elements in the material -> activation
(2) Accelerator water analysis:
  huge water body, lots of debris and blocked filters
Shielding material
analysis
Water and filter
analysis

## Slide 3
Part of the Chemistry&Lifescience Support (CLS) group
PRESENTED BY Monika Hartl
1
Water chemistry
Cooling water, DI water
Accelerator (klystron, RFQs)/Target
=> pH, conductivity, oxygen content => blocked filters, corrosion

## Slide 4: Blocking filters in the accelerator cooling water system
CWH cooling system/Radio frequency power station – frequently blocked filters:
Deposit: Cu(red),Cu2O(black),CuO(red)
Cooling water system - no heat exchangers between sections
Upstream: reddish deposit -> downstream: black
2 HNO3 + CuO + 6 H2O [Cu(H2O)6]2+ + 2 (NO3)- + H2O
14 HNO3 + 3Cu2O+ 6 H2O 6[Cu(H2O)6]2++ 12 (NO3)- + 7H2O + 2NO
XRD pattern –
phase identification
HNO3
Cu
rust
Crystalline Copper oxide species are blocking the filters

## Slide 5: Blocking filters in the accelerator cooling water system
CWH
CWH / CWM system (up to summer 2025)
Loop | CWM | CWH
Water temperature | 24-30C | 45-55C
Filtration CW mains | 1000 μ | 
Filtration LB, MB, HB | 100 μ | 
conductivity | 0.23 S/cm (5 M) | 
pH | 6.5-7.5 | 
accelerator
no heat exchangers between sections (500 m3 water)
Large ion exchanger (bypass) removes Cu+, Cu2+ (if dissolved in water!)
ICP-OES: limits for Cu-detection 5 ppb
100-500 ppm O2 in loop
Summary of copper release rates in function of oxygen level and pH. [R.Svoboda, H.G.Seipp: Power Plant Chemistry, 6 (2004) 1, pg 7-142]
=> Need to reduce O2 in water

## Slide 6: Blocking filters in the accelerator cooling water system
Are we continuously eroding Cu? Or is it an installation problem?
Less oxygen to not dissolve Cu => no blockage
Origin of copper
Metallic Cu
Metallic Cu
Inside of filter
Cu pieces (red) get stuck in the filter (At installation or later? Erosion?)
Over time, Cu pieces in filter are slowly oxidized due to O2 concentration in the water.
Cu2O (black) crystallizes on downstream side of filter (less turbulence) first to Cu2O then to CuO (red).
Outside of filter - downstream
Inside of filter – upstream
Reddish color
Black color
Why do we find copper oxide downstream instead of upstream?
Is oxygen in the water causing formation of Cu2O/CuO?

## Slide 7: CWH status today
Separation of cooling water loops since Oct. 2025: each subsection is separate
Much improved, but still Cu2O deposits: historical or new?
O2 dissolved in water monitored and much reduced: continue to eliminate O2 sources
Adjustment of pH?
Other corrosion (Fe,..)
Ongoing Work:
Rutambhara Yogi
Björn Eldvall
Morten Rostrup Forup Jensen
Walther Borg
Olle Lagerblad
Emmanouil Trachanas
Hadi Sadeghzadeh
Mats Nilson
-> Will be needed during operations
100
50
150
200
250
300
Sep
2025
Nov
2025
Jan
2026
Juli
2025
May
2025
March
2026
Water flow
Flow l/min
CWH
x
x
x

## Slide 8
2
Chemical analysis/ quality checks
90 deg
45 deg

## Slide 9: Composition of ESS shielding and construction materials
Materials in the (future) radiation field
  Materials should not activate (longterm) isotopes (e.g. Eu, Co,…).
  Material composition is important for MCNP calculation.
ESS expertise & equipment has developed:
(1) Sample preparation (ball mill, fusion furnace, microwave)
(2) Density measurements (pygnometer, gas pygnometer)
(3) X-ray fluorescence measurements (XRF) and Laser-induced-breakdown Spectroscopy
(5) ICP-Optical Emission Spectroscopy (ICP-OES)
(6) Scanning Electron Microscopy (SEM)
Commercial analysis labs are expensive and often work according to standardized operating procedures – does not always fit the problem.
various materials: beam port inserts, shutters, concrete walls, beam dump,….and lots of quality gate measurements for neutron scattering instruments.

## Slide 10: Composition of ESS shielding materials
Boron containing polyethylene / Gd2O3 containing kapton
Borated PE (target)
Borated HDPE
for CSPEC
Gd2O3 Collimator for CSPEC
Kaptonsheet not homogenously covered: thickness of Gd2O3 layer is 15 µm at the bottom while in the middle and top is >20 µm.
Borated PE for
BIFROST
Fernando Yamil Moreira
Rosa Camilleri Lledó
Piotr Slawik
Sara Ghatnekar Nilsson
Weight change upon heating: B4C or B2O3 remains
Phase transition around 100 -150 C allows differentiation between LDPE and HDPE.

## Slide 11: Active Cell – remote diamond wire cutting
Shaft Cutting Station - Moderator Reflector Plug
  Super Duplex SS mock-up
Dry cutting trials

## Slide 12: Active Cell – remote diamond wire cutting
Testing dust/swarf distribution before active materials are present
Diamond-string saw used at varying speed.
Chart: Active Cell Dust – Particle Size
| 45 | 64 | 125 | 250 | 500 | 1000 | 2000 | 4000 | 10000
 | 8.8599 | 8.1492 | 22.695 | 29.004 | 26.1318 | 4.245 | 0.69127 | 0.11683 | 0.1071
90 
45
Small particle with sharp edges are produced
Various sampling points at different angles
90 deg
45 deg

## Slide 13: Active Cell – remote diamond wire cutting (ONGOING)
Shaft Cutting Station - Moderator Reflector Plug
  Super Duplex SS mock-up
  5hours (2022) => 27 hours of cutting (2025)
  More tests coming with ventilation in place
Dry cutting trials (2022 and 2025)
Carwyn Jones,
Magnus Göhran
Gumer Verdu
Amina Chemlal
Donna Tam
Harry Robinson
James Kent
Jonathan Rose
James Davies
William Blyth

## Slide 14
3
Trouble-shooting

## Slide 15: Unknown substances in the wrong places
(1) Yellow substance found in monolith, beam ports
assumed chemicals: “crack-testing liquid” used for inspection
FTIR not convincing, SEM only light elements found
-
Goo on the target area cooling water blocks
(2) Liquid (yellowish) found on target wheel
When sampling, noticed foaming:
Gave idea to look at cleaning process for the components
Check pH: acidic
Chemical test for
phosphate
-
sample
Krister Blomberg
Mats Segegerup
Alice Corani

## Slide 16: Unknown substances in the wrong places
Dust and debris in accelerator and target
Transport and storage of unpainted lead bricks in the area of the delay tank bunker for the cooling water => dust containing lead –> method development done.
Lead Dust exposure –delay tank bunker
Formation of PbO on lead shielding on a klystron.
Carbonformation on
tetrodes
Alice Corani
OHS group
And a lot of all other kinds of materials in unexpected places
- pump oil, glue, lubricants, graphite/steel shavings,…
Often a simple and quick identification helps solve the issue
and prevent delays.

## Slide 17: ESS Project Support – trouble shooting
Wire grid for beam rastering – assembly challenges
Grid
beam on target: beam amplitude, position, edges, projected horizontal and vertical profiles
wires are held by ferrule within a sleeve.
Assembly issue:
  components oily – slip of wire does not allow assembly
  cleaning of components in lab with hexane - still slipping
Measurements via optical microscope (10x):
  123 m  4 (specified was 78 m)
  wire diameter 100 m
Ferrule in sleeve shows gaps
Wire is not held in place.
Cyrille Thomas
Mateusz Dominiak-Sumila
Antoine Lepine
Rasmus Johansson
Gabriel Calcisca
Ferrule was dimensioned incorrectly

## Slide 18: Development Work with Fluids & Source Technologies
Smurf Oil for He-compressor (target cooling)
Lubricant for He compressor was heated to
1.5x maximum temperature under 2.5 bar H2.
Lubricant turned blue after 8 weeks.
Test to see if compressor oil is inert to H2 (as a chemical model for tritium):
Properties and main chemical components of the oil seem to be unaffected
Oil composition unknown: Usually Zinc dialkyldithiophosphate (ZDDP) is added as anti-oxidation, anti-wear, and anti-corrosion additive. => forms protective layer on top of steel (FeS) but not confirmed.
Naja de la Cour
Jia-Fei Poon
Jaroslaw Frydrych
Jaime Arriagada
Other He compressor needs – method
development:
Oil quality check:
  rheology/viscosity (Rheometer)
  water in oil (Karl-Fisher titration)
  metals (ICP-OES)
GC-MS determining (NO !) oil in He gas

## Slide 19: Advantages of having in-house support
….over hiring external analytical support
Commercial analysis labs are expensive and often do not deliver what is needed but what they have been asked for.
ESS stakeholders know what they want but often do not know what analysis/support to ask for.
In-house support understands “spallation source” issues better (experience) and can help more quickly; easy to talk to.
Analysis starts by correct sampling
In-house experts can develop methods to analyze materials as needed and then teach technical staff in the respective groups
Sending activated components outside of ESS might become difficult (or costly).
Chemistry can be fun!
We gladly support you in taking care of things yourself.

## Slide 20: Thank you
Spallation Chemistry
ChemistrySpallationTeam@ess.eu
2022-2025

## Slide 22
5
Capabilities

## Slide 23: Chemistry User Laboratories in E04
Used to support ESS construction and commissioning
Life science Lab
X-ray diffraction
Detector lab
Motion Control lab
storage
SEM
Cutting/ polishing
Physical Characterization
Spectroscopy room
Cold
room (later)
Basic Chemistry lab
storage
SULF Offices
Level 100
Level 110

## Slide 24: Radioactive Materials Laboratory
Test cases:
  Testing of oil in He-loop
  tensile testing of proton beam window
Radiation monitor
Lab to check (slightly) radioactive components
RML
Changing
area
Exp. halls

## Slide 25: Phase Characterization/Identification
Powder Diffraction, Optical microscopy, SEM
RIGAKU Smartlab SE Powder X-ray machine (Cu tube): sample changer
Leica Optical microscope with indenter and camera
FEI Environmental Scanning electron microscope (E-SEM):
Owned by Target, operated by Hossein Sina

## Slide 26: TG, BET, pycnometry and electrochemistry
Density, thermal decomposition, porosity/surface area, electrochemistry
Netzsch Jupiter STA: DTA/DSC
T=30-1600 °C, sample size 1-30 mg
BET surface area determination
nitrogen
Gas pycnometer (helium):
measuring the density of a solid material
Autolab potentiostat:
cyclic voltammetry
Hg-drop polarography
voltammetry:

## Slide 27: Elemental Analysis
Operational: XRF, Fusion furnace, ICP-OES, Digestion Microwave
Microwave digestions instrument:
Up to 20x50 ml Teflon cups, T<100 °C
Automatic fusion machine:
Up to 2 samples;
Pt crucibles, approx. 1200 °C
X-ray Fluorescence:
permanent loan from STFC
Ion-coupled-Plasma-Optical Emission Spectroscopy (ICP-OES): ppm/ppb range, up to a few mg
Elemental Analysis CHNO:
Up to 80 samples;

## Slide 28: Spectroscopy equipment
UV/VIS Spectrometer range: 200-850 nm
  D2/W-halogen lamp: 230nm-2.5m
Operational: UV/VIS, FT-IR, RAMAN, DLS
FT-IR Spectrometer range: 450-4000 cm-1
  -ATR for powders
  -hATR for liquids
Dynamic light scattering:
  - R= 0.18 nm - 2500 nm
  - 300 Da to 1MDa
Raman Spectrometer range: 20 - 4000 cm-1 , Transmission and reflectance; can be performed at low temp (cryo – 10K)

## Slide 29: Active Cell – remote diamond wire cutting
Testing dust/swarf distribution before active materials are present
Shaft Cutting Station: Diamond-string saw July 2022 – fast, dry cut (5 h), super duplex SS
90 deg
45 deg
CO2
Large HEPA filter from the air handling units clearly show weigh difference after cutting => what part is debris?

## Slide 30: ESS Project Support – trouble shooting
Various small projects
Pin of flowmeter
(oxidation ->H2WO4)
Pin material in failed flow meter: WC (tungsten carbide). Are cracks in surface normal aging?
Intact galvanized layer
Slightly damaged galvanized layer
Rust bloom on screw thread
Corroded screw after transport of target wheel shaft
drive from manufacturer
SEM/EDX: Cl- found
Salty, moist climate caused corrosion
Prolonged exposure during transport on ship
SEM/EDX: Shows cracks – logbook showed usage of flow meter that exceeded its specifications.

## Slide 31: S
3
Development work

## Slide 32: Development work with Beam Diagnostics Group
Flamespray of target wheel - March 2022. Production of 3x 1 kg luminescent powder
Cr-doped Al2O3 (sintered)
Ce-doped YAG
Mixture of both
We can use commercial chemicals. We can optimize the coating. We are independent.
Collaboration with SNS to test new materials in the radiation field (on the SNS target).
Cyrille Thomas
Emelie Wiklund
Anton Järild
Shrikant Joshi
Håvard Gjersdal
Stefan Bjoreklund
Magnus Sandburg