## Slide 1: ESS Target Monitoring System
ICANS Presentation
PRESENTED BY Evan Foy and Kristoffer Sjögreen

## Slide 2: Agenda
1 ESS Target Wheel monitoring functions (Evan)
2 Instrumentation and control technology (Evan)
3 Examples of mechanical issues that demand monitoring (Kristoffer)
4 Future work to improve remote monitoring (Kristoffer)

## Slide 3: Target System
Logical view
Physical view

## Slide 4: Target System
Bearing vibrations
Bearing and motor temperatures
Rotor position (x-y)
Target wheel position (z)
HRU displacement (x-y-z)
Rotor-to-HRU torque transfer
HRU helium outlet pipe torsional strain
Wheel temperature monitoring (individual sectors)
..
Mechanical monitoring functions

## Slide 5
Layer architecture of control system
ESS Main Control Room
Phoebus GUIs
IOCs VM
Timing
Distribution System
Timing
IOC
EVR
VIBR
IOC
VIBR
PLC
LUBE
PLC
LUBE
IOC
TDU
PLC
TDU
IOC
TMS
PLC
TMS
IOC
other human interfaces
CS Infra IDMZ
Read-Only
EPICS Gateway
Phoebus
Alarms
Phoebus
Data Browser
Alarm Server
Archiver
Appliance
Target Wheel

4x PLCs
1x MTCA crate
TDU
HRU
Ethercat PLCs, interlinked with Ethercat

## Slide 6: Vibration Monitoring
~90 kEUR
~6 kEUR
Instrumentation and Control
Beckhoff ELM3602-0002 AI modules
24bit, 50ksps, BNC
Beckhoff C6030 IPC
4 cores

## Slide 7: Vibration Monitoring
Beckhoff Twincat – TF3600 / Condition Monitoring Function
  Helpful documentation from Beckhoff!
  Sample projects
Vibration assessment
Crest Factor
Magnitude spectrum
Envelope spectrum

## Slide 8: Vibration Monitoring
  Function is verified using a handheld B&K Calibration Exciter.
  A spare accelerometer and cable is used to connect the exciter to each analog input channel of the Beckhoff data acquisition system.
  When the exciter is switched on we expect the following results:
    10 m/s2 Acceleration RMS
    10 mm/s Velocity RMS
    10 um Displacement RMS
    Vibration peak at 159.2 Hz on the FFT
FAT / SAT

## Slide 9: Vibration Monitoring
EPICS Integration

## Slide 10: Vibration Monitoring
EPICS Integration

## Slide 11: Vibration Monitoring
EPICS Integration

## Slide 12: Vibration Monitoring
EPICS Integration

## Slide 13: Vibration Monitoring
EPICS Integration

## Slide 14: Temperature Monitoring
Instrumentation and Control
  Sensors are interfaced to Beckhoff EL3214 AI modules
  PLC logic monitors each measurement and determines a “rotation permit” condition.
  Status is monitored by separate drive control PLC, which interlocks rotation in case any “rotation permit” condition is False.
TE-023,
TE-024
TE-021,
TE-022
TE-011,
TE-012
Beckhoff PLC
Beckhoff EL3214 AI terminals
RTD/PT100 (3-wire), 4 ch.
6x PT100s

## Slide 15: Temperature Monitoring
EPICS Integration

## Slide 16: Temperature Monitoring
Behaviour

## Slide 17: Rotor XY Position Monitoring
Instrumentation and Control
  Rotor position measurement is based on two Precitec Chromatic confocal optical heads mounted in the Target Monitoring Plug (TMP) inside the Monolith vacuum vessel. The optical heads are connected to an external controller via optic fibers.
  The instrumentation measures raw distances to the rotor surface. The control system calculates the x,y coordinates of the rotor center, and correlates the values with the rotor angle in real time.
Beckhoff PLC
Precitec Chromatic confocal sensor technology
Precitec CHRocodile

## Slide 18: Rotor XY Position Monitoring
EPICS Integration

## Slide 19: Wheel Z Position Monitoring
Instrumentation and Control
  Target Wheel vertical position (z) is measured using a Microepsilon laser time-of-flight instrument.
  The instrument is mounted outside the Monolith vacuum vessel. It directs a laser downwards through a sightglass, through the Target Monitoring Plug via mirrors to the surface of the Target Wheel.
Beckhoff PLC
Microepsilon ILR Laser time-of-flight distance instrument
Profibus

## Slide 20: Wheel Z Position Monitoring
EPICS Integration

## Slide 21: Wheel Z Position Monitoring
FAT / SAT tests

## Slide 22: Wheel Z Position Monitoring
FAT / SAT tests

## Slide 23: HRU X-Y-Z Displacement Monitoring
Instrumentation and Control
  Three high-sensitivity eddy current sensors from Telemess are mounted at various positions around the HRU to measure the displacement in x, y and z directions.
  A PLC monitors each measurement and determines a “rotation permit” condition. Status is monitored by separate PLC for drive control, which interlocks rotation in case any “rotation permit” condition is False.
Beckhoff PLC
EL3174 AI modules
Telemess high sensitivity eddy current sensors
Telemess signal conditioning units

## Slide 24: HRU Helium outlet pipe torsional strain
Instrumentation and Control
  A high-sensitivity eddy current sensor is used to measure the gap between a pair of brackets mounted on the helium outlet pipe. The gap indicates the torsional strain experienced by the pipe.
  A graphical user interface displays the value and indicates MINOR / MAJOR alarms when the value exceeds predefined limits.
Beckhoff PLC
Beckhoff EL3174 AI module
Telemess high sensitivity eddy current sensor
Signal conditioning units
Pipe torsion
Mechanical
bellow
Helium outlet pipe

## Slide 25: Rotor-to-HRU torque monitor
Instrumentation and Control
  A high-sensitivity eddy current sensor is used to measure the gap between a pair of brackets mounted on the helium outlet pipe. The gap indicates the torsional strain experienced by the pipe.
  A graphical user interface displays the value and indicates MINOR / MAJOR alarms when the value exceeds predefined limits.
Beckhoff PLC
Beckhoff EL3174 AI module
Telemess strain gauge + calibrated signal conditioning unit
+/-30V

## Slide 26: HRU monitoring functions
EPICS integration

## Slide 27: HRU monitoring functions
EPICS integration

## Slide 28: Target Sector Temperature Monitoring
Instrumentation and Control
  Target Sector Temperature Monitoring is based on two pyrometer measurements. Temperature is measured using infrared (IR) sensor technology. The pyrometer heads are mounted inside the Target Monitoring Plug (via the IR-insert). IR radiates from the wheel surface to the pyrometer heads via mirrors.
  The pyrometer devices are interfaced to the TMS PLC. The Twincat project for the TMS PLC is “TcProj_TMS”. The Sector Temperature Monitoring function is handled by “PRG4_SectorTemps”.
2x Mergenthaler
2x Pyrometers
”Cup”
”Shroud”
TU-301
TU-302
TMS PLC

## Slide 29
Sound and vibration examples and approach
2

## Slide 30: Introduction
The Target System was installed in the Monolith in May 2024.
Up to now, it has been accessible most of the time, and the monitoring of the machinery has been done through inspections in combination with monitoring.
Now we approach a point where we will need to rely solely on the monitoring functions, since the machinery will not be accessible.
In these few slides, we will go through some experiences with the current monitoring system, and end with some suggestions of how to continue the work.
Sub-headline to strengthen the headline above

## Slide 31: Sound
Examples from the testing
Contact in a labyrinth
Too low load on a bearing
Misaligned encoder

## Slide 32: Sound
Examples from the testing
Internal movement
(Shaft sleeve)
Internal movement
(cartridge)
Microscopy of shims beneath the cartridge

## Slide 33
The shaft and the sleeve are constrained to the same rotational velocity, and a force is applied on the sleeve, forcing it into contact with the shaft. Since it is not possible to have different rotating speeds, the shaft will try to slow down the sleeve
Gradual internal movement

## Slide 34: Vibration monitoring- online
Acceleration, velocity and displacement as well as Crest factor is monitored for all bearings
Vibration vs flow rate, for the Vacuum feedthrough
Vibrations in drive unit (BA-011y) as well as the HRU (BA-021 and BA-022) during last run before HRU failure
Periodic vibration, corresponding to the sound

## Slide 35: Vibration monitoring- offline
Bearing frequencies
It is challenging to try to keep track of bearing damage for a very slow rotor like the target.
To enable offline analysis, periods of 30 seconds of data are logged in EPICS with an interval of three hours
Various techniques commonly used

## Slide 36: Vibration monitoring- offline
Test magnitude spectra for one of the drive unit main bearings
It is possible to listen to the accelerometers by converting raw data to sound files

## Slide 37: Helium rotary union instrumentation
Measurements on resulting moment from the pipe on the HRU
Measurements of the horizontal and vertical displacement
Measurement of the rotating torque

## Slide 38: Helium rotary union instrumentation
Measurements on resulting moment from the pipe on the HRU
Measurements of the horizontal and vertical displacement
Measurement of the rotating torque
 | Before start | With rotation (23.3 rpm) | Buffer pressure (11 bar g) | Vessel pressure (8 bar g) | Stop | Lifting due to pressurization of the vessel
Absolute vertical 1 | 22.40 mm | 22.61 mm | 22.67 mm | 23.45 mm | 23.85 mm | 0.78 mm
Absolute vertical 2 | - | 14.34 mm | 14.39 mm | 15.00 mm | 15.40 mm | 0.61 mm
Relative vertical | 3.79 mm | 3.94 mm | 3.93 mm | 4.56 mm | 4.64 mm | 0.63 mm
X- direction absolute | 3.85 mm | 3.84 mm | 3.85 mm | 3.88 mm | 3.81 mm | 0.03 mm
Y- direction absolute | 6.68 mm | 5.70 mm | 5.69 mm | 5.75 mm | 5.73 mm | 0.06 mm
Relative horizontal | - |  |  |  |  | 
Torque arm | 8 N/2 Nm | 451 N/121 Nm | 658 N/ 177 Nm | 629 N/169 Nm | 516 N/ 139 Nm | -

## Slide 39: Conclusions
The vibrations in the Target systems are dominated by broadband vibrations from the helium flow. In addition there are deterministic vibrations from the rotation.
The current online vibration system works well in predicting the average vibration level and protecting the system from damage.
The soundfile conversion of accelerometer data helps in understanding sound phenomena.
It is recommended to continue developing the analysis toolbox for vibration, to be able to develop the current system towards a proven condition monitoring system, especially to characterize damage frequencies of the rotor.
The additional sensors installed to monitor the movement of the HRU, the torque and the forces from the piping system have been proven useful to validate calculations.

## Slide 40: Finish presentation