## Slide 1
Zhijia Sun（zhjsun@ihep.ac.cn） Detector&Electronics Group
China Spallation Neutron Source, IHEP, CAS
April 16 2026，Malmö
Neutron Imaging Detectors @ CSNS
The 25th International Collaboration on Advanced Neutron Sources

## Slide 2
Summary and Outlook
Contents
Outline
Resolution, FOV, Compact Neutron Imaging
Background
Energy Resolved Neutron Imaging Based on Timepix4
III
I
II
IV

## Slide 3: China Spallation Neutron Source(CSNS)
Aerial View of CSNS
 | Phase I | Phase II
Proton Beam Power (kW) | 100 | 500
Pulse Frequency (Hz) | 25 | 25
Mean Intensity of Proton Beam (A) | 62.5 | 312.5
Energy of Protons (GeV) | 1.6 | 1.6
Proton Energy at Injected Port of RCS (MeV) | 80 | 300
Number of Instruments | 3 | 9
Properties of CSNS
Layout of CSNS
Instruments Layout
3
01 Small-Angle Neutron Scattering
02 Multi-Purpose Reflectometer
03 Liquid Reflectometer
04 Cold Neutron Inelastic Spectrometer
05 High-Energy Direct-Geometry Inelastic Spectrometer
06 Indirect-Geometry Molecular Vibrational Spectrometer
07 Reserved
08A Neutron Technology Development
08B Engineering Material Neutron Diffractometer
09 High-Resolution Neutron Diffractometer
10 Neutron Backscattering Spectrometer
11 Atmospheric Neutron Irradiation Spectrometer
12 Neutron Physics and Application Spectrometer
13 Energy-Resolved Neutron Imaging Instrument (ERNI)
14 Very Small-Angle Neutron Scattering
15 High-Pressure Neutron Diffractometer
16 Multi-Physics Instrument
17 Elastic Diffuse Scattering Neutron Diffractometer
18 General-Purpose Powder Diffractometer
19 Macromolecular Single-Crystal Neutron Diffractometer
20 Direct-Geometry Polarized Inelastic Spectrometer
3 phase-I instruments
8 cooperative neutron instruments
9 phase-II instruments

## Slide 4: Road-map
Science
Neutron imaging physics,
instrument requirements
Team
Detector & Electronics
Group, CSNS/IHEP
Technology
Scintillators, Timepix4,
in-house electronics
CSNS
Neutron Imaging Detectors
Instruments
Materials Science,
Energy, Archaeology
Archaeology and Cultural Heritage
Environmental Science
Fuel cell
Renewable Energy

## Slide 5: Neutorn Imaging Dtector Family
2025
High spatial and Medium field detectors
Ceramic scintillator
Energy-resolved detector (Timepix3)
2017
2023
Imaging detectors commissioned at ERNI
Camera based on Timepix4
……
2021
Large field detectors
2024
Neutron imaging microscope, spatial resolution 6.6 μm
Energy-resolved detector (Timepix4)
Ceramic scintillator thickness 20 μm spatial resolution 12 μm
2018
spatial resolution 18 μm achieved
Compact imaging detector
Neutron Imaging Detector family
High spatial resolution detector
GOS scintillator (powder, ceramic) and hBN scintillator
Spatial resolution <10 μm
Energy-resolved detector
High speed camera (Timepixe3，TimePixel4)
Max FOV ≥100mm × 100mm，spatial resolution <14 μm
Large field detector
Large area 6LiF/ZnS scintillator
Max FOV ≥200 mm × 200 mm
Compact detector
sCMOS camera, medium area 6LiF/ZnS scintillator
FOV 100 mm × 100 mm, spatial resolution 122 μm

## Slide 6
Summary and Outlook
Contents
Outline
Resolution, FOV, Compact Neutron Imaging
Background and Motivation
Energy Resolved Neutron Imaging Based on Timepix4
III
I
II
IV

## Slide 7: High Spatial Resolution Imaging Detector
Ultra-thin isotopically enriched gadolinium oxysulfide (¹⁵⁷Gd₂O₂S:Tb) scintillator
Thickness 5 μm , 88% abundance of 157Gd, particle size ranges from several hundred nanometers to approximately 2 μm.
Scintillator simulation
FWHM 6μm
Light Spot
Scintillator fabrication
（Technical Institute of Physics and Chemistry，CAS）
GOS particles
XRD of Gd isotopes
Fluorescence spectrum
SEM ( low magnification）
SEM( high magnification）
SEM(cross-section）

## Slide 8: High Spatial Resolution Imaging Detector
The micron-level spatial-resolution neutron imaging detector
First ‘Neutron imaging microscope’ at CSNS, spatial resolution 6.6 μm，field of 6.4 mm×6.4 mm
Principle of the detector
Design of the detector
Experiment layout at ERNI
Spatial resolution performance

## Slide 9: High Spatial Resolution Imaging Detector
Research on alternative scintillator for GOS powder scintillator
GOS ceramic scintillator：beset resolution 12 μm at 50 μm thickness
GOS Ceramic
12 μm
17 μm
14 μm
15 μm
Notes: 有疑问，问

## Slide 10: Large field imaging detector at ERNI
Large sensor CCD camera and large area 6LiF/ZnS scintillator,
FOV > 200 mm×200 mm, spatial resolution 75 μm.
Detector structure
Commissioned at ERNI
FOV 220mm × 210mm
Resolution 120μm
High resolution large FOV detector

## Slide 11: Compact imaging detector
Development of compact imaging detector for fast imaging
Detector structure
Detector parameters
122μm@5%MTF
PCO sCMOS camera and medium area 6LiF/ZnS scintillator, FOV ~ 100 mm × 100 mm, spatial resolution 122 μm.
GPPD, KB Lens Testing
HD, Neutron Beam Testing
VSANS, Slits Testing

## Slide 12
Summary and Outlook
Contents
Outline
Resolution, FOV, Compact Neutron Imaging
Background and Motivation
Energy Resolved Neutron Imaging Based on Timepix4
III
I
II
IV

## Slide 13: Energy-resolved imaging detector based on Timepix3
High spatial resolution and high time resolution imaging detector was developed based on Timepix3 chip
Principle of the detector
Detector engineering design
Fast connection interface
Key components
pixels :256×256,55μm×55μm | 
time res.1.56 ns | 80 Mhits/s
Neutron imaging and energy spectrum, Single event (X,Y,TOA,TOT), Variable field

## Slide 14: Energy-resolved imaging detector based on Timepix3
ERNI（ Energy Resolution Neutron Imaging instrument ）
Field of view: 100mm×100mm, spatial resolution: 14 μm, relative wavelength resolution: 0.16% @2Å
GOS Ceramic, magnification 2.74X
Operated at ERNI instrument
Performance of the detector
304 steel Bragg edge
Mean 15.43ms
Sigma 0.024ms
Time(wavelength) resolution
GOS Powder, magnification 1.9X
Spatial resolution
Maximum field
The detector has been operating stably since 2023

## Slide 15: Energy-resolved imaging detector based on Timepix4
Key Advances
Medipix4 Collaboration (since 2022)
CSNS joined the Medipix4 Collaboration (CERN), gaining access to the latest Timepix4 ASIC and collaborative R&D resources for next-generation pixel detectors
Timepix4 + In-house Electronics
Developed Timepix4-based detector with fully in-house designed readout electronics, achieving independent control of detector system and data acquisition chain
Real-time Data Pre-processing
In-house electronics enables real-time data pre-processing, replacing commercial ASI TPX3 camera with significantly reduced cost and improved flexibility
CSNS In-house Timepix4 System
✓ Timepix4 chip (512x448 pixels, 55 μm pitch)
✓ Fully in-house designed readout electronics
✓ Real-time on-board data pre-processing
✓ Member of Medipix4 Collaboration (CERN)
✓ Customizable firmware & DAQ pipeline
✓ Spatial resolution: 28 μm (GOS scintillator, 2.1x)

## Slide 16: Energy-resolved imaging detector based Timepix4
Joined Medipix4 Collaboration in 2022
Oct 2022
Joined
Medipix4
Collaboration
(CERN)
2023
Circuit
Development
In-house readout electronics design
2024
Prototype
Development
System assembly & bench testing
2025
Neutron Beam
Experiments
TOF tests & calibration
2026+
Application
Exploration
Materials, energy, archaeology...
Welcome international collaborations on developments on Timepix4 based neutron imaging detectors!
Brazil LNLS detector team visit CSNS

## Slide 17: Energy-resolved imaging detector
CSNS CTPX1 High-Speed Event-Driven Camera
Based on Timepix4
Full-speed readout of single module (16 x 5.12Gbps links)
Based on Xilinx Zynq Ultrascale+ ZU15EG FPGA
Multi-stage buffer, max count rate 1.2Gcps
Integrated HV generation/monitoring and semiconductor temperature control
Timepix4 Chip Board and Readout Board
PCIe Readout Card
Readout links
80Gbps/ 1.25Gcps/32GB buffer
64Gbps/ 800Mcps/32GB buffer
PCIe Connection
32Gbps/400Mcps/1T buffer

## Slide 18: Energy-resolved imaging detector based on Timepix4
Image intensifier and camera
Top view of the detector
Isotope Light Source and USAF Test Pattern
Testing by Laser
The detector structure

## Slide 19: Energy-resolved imaging detector based on Timepix4
neutron beam testing — TOF and counting rate
TOF
Neutron TOF spectrum from TPX4 camera is consistent with LND 3He PSD.
counting rate
Img. Int. Gain | Duration | File Size | Data Rate | Average Counting Rate | Maximum Counting Rate
Gain3400 | 5min | 4169MB | 13.9MB/s | 0.9 Mhits/s | 3.4M hits/s
Gain3600 | 5min | 13903MB | 46.3MB/s | 2.89 Mhits/s | 11.2M hits/s
Gain3800 | 5min | 29525MB | 98.4MB/s | 6.1 Mhits/s | 23.5M hits/s
Gain3900 | 5min | 37853MB | 126.2MB/s | 7.8 Mhits/s | 30.1M hits/s
Gain4000 | 5min | 44168MB | 147.2MB/s | 9.2 Mhits/s | 35.5M hits/s
Maximum event rate ：35.5M cps
Limited by flux of neutron beam

## Slide 20: Energy-resolved imaging detector based on Timepix4
Spatial resolution ：28 μm
Measurement setup
Spatial resolution
Scintillator: GOS, Mag.: 2.1x
Scintillator: 6LiF/ZnS, Mag.: 1x
Scintillator: 6LiF/ZnS, Mag.: 0.348x
Scintillator: 6LiF/ZnS, Mag.: 0.1375x
FOV Φ8mm
FOV Φ17mm
FOV Φ48mm
FOV 100mm×100mm

## Slide 21: Energy-resolved imaging detector based on Timepix4
Samples measurement
Bragg-edge Imaging
Energy-selected Imaging
Austenitic stainless steel
Imaging result
Co/Mo/W/Fe/Ni/Cu/Brass
Imaging result
W
Ni
W
Ni
W and Ni
The relative strength of W and Ni has reversed.

## Slide 22
Summary and Outlook
Contents
Outline
Resolution, FOV, Compact Neutron Imaging
Background and Motivation
Energy Resolved Neutron Imaging Based on Timepix4
III
I
II
IV

## Slide 23: Summary and Outlook
Summary
High Spatial Resolution
• Neutron imaging microscope: 6.6 μm
• GOS ceramic scintillator: 12 μm @ 50 μm
• ¹⁵⁷Gd₂O₂S:Tb, 5 μm, 88% enriched
Large Field Detector
• FOV: 200 mm × 200 mm
• Spatial resolution: 75 μm
• Commissioned at ERNI
Compact Detector
• FOV: 100 mm × 100 mm
• Spatial resolution: 122 μm
• with sCMOS camera
• Size: 18 x 37 cm x 22 cm
• Deployed at GPPD, HD, VSANS
Timepix4 Energy-Resolved
• Spatial resolution: 28 μm
• Max rate: 35.5M hits/s at CSNS
• Bragg-edge & energy-selected imaging
Outlook
Neutron Imaging Detector R&D
Continue pushing the performance frontier across high-resolution, large-FOV, compact, and energy-resolved neutron imaging detectors
International
Active member of Medipix4 Collaboration (CERN) since 2022; further Timepix4 application development and deeper participation in the global pixel detector community
Welcome Collaborations!
Warmly welcome interested teams worldwide to collaborate on neutron imaging detector R&D — in scintillators, electronics, imaging algorithms, and applications sunzj@ihep.ac.cn

## Slide 24
Thanks for your attention！
Notes: 感谢CEAN合作团队，国内团队，funding