## Slide 1
Bin Tang (tangb@ihep.ac.cn), Detector&Electronics Group
China Spallation Neutron Source, IHEP, CAS
April 16 2026，ICANS XXV，Malmö
Progress of Scintillator Neutron Detectors at CSNS

## Slide 2
Summary and outlook
Contents
Outline
Scintillator Detector Family at CSNS
Motivation and detector demands at CSNS
SHiNE: a high-resolution detector for future instruments
3
1
2
4

## Slide 3
China Spallation Neutron Source
The energy of the superconducting linear accelerator will be increased to 300 MeV
Increase the beam power to 500kW
Total of 20 instruments need to be installed
High-power target station upgrad
CSNS I（3 instruments ）
CSNS timeline
Cooperation（9 instruments ）
CSNS II（11 instruments ）
2011~2018
2018~2024
2024~2029

## Slide 4: Detector Requirements for CSNS Instruments
NO | Spectrometer | Detector type | Detector area | Main paremeters
1 | GPPD (BL18) | Scintillator detector | 6 m2 | Efficiency≥40% @ 2Å, Spatial resolution ≤ 4mm×4mm
2 | SANS (BL01) | 3He Tube Array Detector | 1m2 | Efficiency≥80% @ 2Å, Spatial resolution ≤ 8mm (FWHM)
3 | MR (BL02) | 3He_MWPC | 200mm×200mm | Efficiency≥60% @ 2Å, Spatial resolution ≤ 2mm×2mm
4 | MPI (BL16) | 3He Tube Array Detector | 3m2 | Efficiency≥90 % @ 2Å, Spatial resolution≤8mm (FWHM)
5 | ANIS (BL11) | 3He Tube Array Detector | 3m2 | Efficiency≥ 90% @ 2Å, Spatial resolution ≤8mm (FWHM)
6 | EMD (BL08) | Scintillator detector | 3m2 | Efficiency≥60 % @ 2 Å, Spatial resolution ≤ 3mm×180mm
7 | ERNI (BL13) | Scintillator detector | 4 m2 | Efficiency≥60 % @ 2 Å, Spatial resolution ≤ 3mm×50mm
8 | VSANS (BL14) | 3He Tube Array Detector | 3.6 m2 | Efficiency≥ 90% @ 2Å, Spatial resolution ≤ 8mm (FWHM)
9 | HD (BL05) | 3He Tube Array Detector | 20 m2 | Efficiency≥ 90% @ 2Å, Spatial resolution ≤ 25mm
10 | HPND (BL15) | 3He Tube Array Detector | 5.7 m2 | Efficiency≥80 % @ 2Å, Spatial resolution ≤ 8mm (FWHM)
11 | HRD (BL09) | 3He Tube Array Detector | 8.8 m2 | Efficiency≥70% @ 2Å, Spatial resolution ≤ 8mm (FWHM)
12 | NTDS (BL08A) | 3He Tube Array Detector | 32mm×300mm | Efficiency≥80% @ 2Å, Spatial resolution≤ 8mm (FWHM)
13 | NPAI (BL12) | 10B GEM | 50mm×50mm | Efficiency≥40% @ 6Å, Spatial resolution≤ 2mm (FWHM)
14 | IMVS (BL06) | 3He Tube Array Detector | 0.36 m2 | Efficiency≥80% @ 1.8Å, Spatial resolution≤8mm (FWHM)
15 | NUBS (BL10) | 3He Tube Array Detector | 0.44 m2 | Efficiency≥85% @ 6.3Å, Spatial resolution≤12.7mm (FWHM)
16 | EDS (BL17) | 3He Tube Array Detector | 7m2 | 3He:Efficiency≥85% @ 1.8Å, Spatial resolution≤25.4mm (FWHM)
17 | CNIS (BL04) | 3He Tube Array Detector | 15.2 m2 | 3He:Efficiency≥85% @ 1.8Å, Spatial resolution≤25.4mm (FWHM)
18 | PACS (BL20) | 3He Tube Array Detector | 12.2 m2 | Efficiency≥85% @ 1.8Å, Spatial resolution≤25.4mm (FWHM)
19 | SCND (BL19) | Scintillator detector | 2.3 m2 | Efficiency ≥ 50% @ 2Å, Spatial resolution ≤1mm×1mm
20 | LNR (BL03) | 3He_GEM | 600mm×1000mm | Efficiency≥70% @ 6Å, Spatial resolution 2mm×2mm
CSNS I
Cooperation
CSNS II
Notes: Detector performance has been improving steadily .

## Slide 5: scintillator Detectors for CSNS Instruments
    Requirements of the neutron scattering spectrometer
Large area coverage
High detection efficiency
Good spatial resolution
n/γ discrimination
Engineering scalability
Notes: Why 4 instruments choose scintillator detectors as the main detector
Because it provide the most suitable balance for these specific applications.
our detector can fulfill these specific requirements.
Large area coverage for wide angular acceptance.
High detection efficiency for weak signals and time-resolved measurements.
Good spatial resolution to support instrument resolution.
n/γ discrimination for background suppression.
Engineering scalability is essential for multi-square-meter detector systems.

## Slide 6: Scintillator Detector Family at CSNS
2010~2018 1st generation FiRST
Chart: Scintillator detector family performance
| Detection efficiency | Spatial resolution | Count-rate/chanel | Count-rate/moudle | Dead-zone fraction
1st generation | 50 | 20 | 40 | 20 | 60
2nd generation | 85 | 40 | 40 | 50 | 80
3rd generation | 75 | 90 | 80 | 90 | 100
2018~2024 2nd generation LoGHT
2023~2029 3rd generation SHiNE
4mm×4mm
43.6%
28%
2×105cps
3mm×50mm
61.2%
18%
2×105cps
5×105cps
106cps
5×105cps
0.35mm×0.35mm
2%
50.3%
Notes: Detector performance has been improving steadily

## Slide 7
FiRST (Fiber-readout Scintillator Tracker)
  Neutrons
MA-PMT
Double layer 6LiF/ZnS:Ag Scintillator
Detection efficiency：43.6％ @ 2Å
Spatial Resolution： 4mm(H) × 4mm(V)
Max Count Rate：2×105cps/moudle
before
after
Ma-PMT Readout
Self-design ASIC chips
Improved neutron detection efficiency uniformity for each channel
Structure
Key technolgy
Multi-channel
High gain
Matched to WLS fiber arrays
Compact integrated
Adjustable discrimination threshold for each channel
channel
relative gain(a.u.)
Notes: The gain difference between --- needs to be addressed.
compact

## Slide 8
Detector array in GPPD (6m2)
Improved detection-efficiency uniformity gives a more uniform diffraction ring and improves the reliability of peak-intensity measurement.
Supports GPPD（General-Purpose Powder Diffractometer）good data quality measurements in 2018~2026
Angle ( ° )
resolution
1.4752 Å wavelength
Simulation
Measurement
The diffraction peaks of the Si powder sample
FiRST (Fiber-readout Scintillator Tracker)
Resolution

## Slide 9: LoGHT（Louver Geometry High-efficiency Thermal-neutron Detector ）
Cured detector shape: keeps the neutron incidence angle constant over the whole detector area
Louver Geometry: increase the neutron incident length
Simulation results of detection efficiency for different tilt angle
Linear range for Single-channel: 10~2×105cps
Beam test @CSNS BL20
Structure
Bank1
Bank2
Bank3
Bank5
Bank6
Detection efficency: 61.2% @2Å
Notes: SiPM readout provides: compact size, low operating voltage
The key idea is the louver geometry. By tilting the scintillator screens, the neutron path length inside the converter is increased, so the detection efficiency can be significantly improved compared with the flat-panel configuration. As shown in the simulation, the efficiency increases with the tilt angle, and for our optimized geometry it reaches more than 60% at 2 angstrom.
The second important point is the curved detector shape. This keeps the neutron incidence angle nearly constant over the full detector area. This helps maintain a more uniform detector response across the bank.
LiGHT combines louver scintillator geometry, WLS-fiber light collection, SiPM readout, and customized ASIC-based electronics, so it is not only a high-efficiency concept, but also a detector architecture suitable for large-area instrument deployment.
Finally, the beam test at CSNS BL20 confirms the detector performance. We achieved more than 40% efficiency at 1 angstrom, more than 60% at 2 angstrom, and a single-channel linear count-rate range from about 10 to 2 × 10^5 cps.

## Slide 10: LoGHT（Louver Geometry High-efficiency Thermal-neutron Detector ）
Key technology: SiPM readout and customized ASIC electronics
SiPM: good gain uniformity, compact size, low operating voltage, temperature-drift compensation
ASIC chip: independent threshold adjustment, supports up to 2×105 cps/pixel and 5×105 cps/moudle
FPGA
PA
CMP
Customized ASIC
SOC FPGA
Data Processing
Data Aggregation
Kafka DAQ Network
SiPM
SOC Data Aggregator
FE Readout Board
Amplifying & Discrimination
Off-detector
On-detector
Under detector crate
Notes: another key technology of LiGHT
Compared with conventional PMT-based solutions, SiPMs offer several advantages for instrument-scale deployment, including good gain uniformity, compact size, low operating voltage, and the possibility of temperature-drift compensation.
On the electronics side, we developed a customized ASIC-based readout chain. It supports signal amplification, discrimination, on-detector data processing, and data aggregation.
In particular, the threshold of each signal channel can be adjusted independently, which is very important for channel-by-channel optimization and uniform detector response.
With this readout architecture, LiGHT achieves both high integration and a single-pixel count-rate capability up to

## Slide 11: LoGHT（Louver Geometry High-efficiency Thermal-neutron Detector ）
Start operation from 2023~2026
~3 m², 3 mm × 180 mm pixel
The variation of the residual stress of the full-size high-speed train wheels along the depth direction obtained through EMD
EMD（Engineering Materials Neutron Diffractometer）
The uniaxial lattice stress-strain curve of the entropy alloy CrCoNi
Ti-2Al-2.5Zr biaxial lattice stress-strain curve
Notes: LoGHT, has been deployed on EMD and ERNI.
EMD has been used to measure the residual-stress variation along the depth direction in full-size high-speed train wheels. It has also supported in-situ mechanical studies, such as the lattice stress–strain behavior of advanced alloys.

## Slide 12: LoGHT（Louver Geometry High-efficiency Thermal-neutron Detector ）
Start operation from 2023~2026
~4 m², 3 mm × 50 mm pixel
Bank1
Bank2
Bank3
Bank5
Bank6
Research on Lithium Battery Performance (Imaging + Diffraction)
ERNI (Energy Resolved Neutron Imaging Instrument)
Notes: and LiGHT serves as its diffraction detector system.
The total detector area is about 4 square meters, with a pixel size of 3 mm by 50 mm
the detector is already supporting real scientific applications. As shown here, ERNI can be used for lithium-battery studies, where imaging and diffraction information are combined to follow structural evolution during charge and discharge.

## Slide 13: SHiNE ( SiPM-based High-resolution Neutron Detector)
Tne Anger Camera in SNS-TOPAZ instrument
SHiNE concept -- a kind of Anger Camera
neutron
Scintillator screen
optial guide
light
SiPMs Array
Structure
Notes: SHiNE stands for SiPM-based High-resolution Neutron Detector.
The basic idea is inspired by the Anger-camera concept, which has already been successfully applied on instruments such as TOPAZ at SNS for single-crystal neutron diffraction.
Our detector uses a SiPM array instead of a conventional PMT-based solution, together with a scintillator screen, a light guide, and compact readout electronics.
From the detector structure shown on the right, a neutron is captured in the scintillator screen, the scintillation light spreads through the optical guide, and the light distribution is then sampled by the SiPM array. Based on that light-sharing information, the hit position can be reconstructed with high precision.
The design is aimed at future single-crystal neutron instruments.

## Slide 14
C D
New Scintillator: high detection efficiency
neutron signal
Performance：
Pulse Timing​​: Rise time <50 ns, Fall time <500 ns；
​​Light Yield​​: Up to 4.7 times that of standard GS20 glass and 75% of commercial ZnS:Ag/LiF
Detect efficiency：50%>2Å
n/gamm rejection：~10-4
S. Lv, B. Tang, Z. Sun, S. Zhou et al, Nat. Commun. 15, 6746 (2024).
Degui Tu, Dazhao Wang, Xunpiao Liu, Shichao Lv, Bin Tang*, et al, Journal of the American Ceramic Society, 107, August 2024
Centroid Algorithm: sub-mm positon resolution
Detection mesurement result
Key technology
SHiNE ( SiPM-based High-resolution Neutron Detector)
Multi-channel Independent Readout Electronics
The threshold value, pulse width, and Time-of-Flight (ToF) are recorded
Notes: This slide shows the key technologies of SHiNE. channel-multiplexed readout electronics
The first is the centroid-based position reconstruction algorithm, which uses light sharing among multiple SiPM channels to achieve sub-millimeter spatial resolution.
The second is the new high-performance scintillator, which provides fast timing, high light yield, and promising neutron detection efficiency.
The third is the multi-channel independent-readout electronics, which records channel-level timing and ToT information for more accurate event reconstruction.
So SHiNE combines algorithm, material, and electronics optimization to achieve high-resolution neutron detection.

## Slide 15: Stucture: Minimal dead-zone design
Scintillator
Optial guide & Radiation shield
SiPM Array
neutron incident
Splitting and intensity distortion caused by module gaps
The SiPM array can be tiled over a large area with constant pixel gap across modules
Large-area Optical-guide coupling scintillator together with a small gap below 0.1 mm
Cross-module event merging based on temporal and spatial clustering
Integrity
Diffraction-peak obtained by detector
Notes: For tiled large-area detectors, the gap between adjacent modules is a critical issue because it can cause peak splitting, intensity distortion, and data loss near the module boundary.
To solve this problem, we use ultra-tight scintillator tiling with a stitching gap below 0.1 millimeter, large-area tiling of SiPM arrays with a uniform pixel pitch, and cross-module event merging based on temporal and spatial correlation.
In this way, the detector can be scaled to a larger area while preserving diffraction-peak integrity 衍射峰完整性
Ultra-tight scintillator tiling + tiled SiPM arrays + cross-module event reconstruction

## Slide 16
Two-step: temporal clustering followed by spatial clustering
Temporal clustering：T0-based pre-segmentation of the signal stream
Spatial clustering：Position-window-based event selection
The spaical resolution is better than 0.35 mm (averaged FWHM )
Clustering-based position reconstruction algorithm
Neutron Beam test
Notes: In this way, we can reconstruct neutron events more accurately from distributed SiPM signals. makes the reconstruction more efficient
The projected peaks are then fitted, giving an averaged FWHM of about 0.35 mm.
we first determine which signals are likely to belong to the same neutron event in time.
use position windows to identify which channels belong to the same physical event in space.

## Slide 17: SHiNE @ SCND(Single Crystal Neutron Diffractometer)
Accurate Bragg-peak localization
Intensity
Multi-wavelength "white LoGHT"
Laue diffraction
Single crystal
Bragg scattering:
Sharp & high intensity
Average structure of atoms in real space
Long-range order
Advantages of SHiNE for SCND
Sub-mm spatial resolution
Detector pixel Angle 0.0477°(0.5 mm )
High detection efficiency
Minimal-dead-zone
Large-area coverage

## Slide 18
Summary and outlook
03
focus on higher efficiency, minimal dead zones, scalable large-area assembly, and real-time reconstruction.
Future work
02
Novel scintillators, MA-PMT and SiPM readout, customized ASIC chips, and advanced position-reconstruction methods.
Key technologies developed
01
Scintillator detector family at CSNS address instrument needs from large-area deployment to future high-resolution applications.
Scintillator detector family
04
International colleagues for valuable discussions and exchanges, domestic collaborators suport.
ISIS group for detailed technical discussions and constructive suggestions.
Acknowledgements
CSNS Detector & Electronics Group
Notes: The next-generation detector should also provide:
small dead zones
compact module structure
vacuum compatibility
scalable large-area assembly

## Slide 19: Thanks
Welcome and Acknowledgements
Be curious and go deep into matter, like a neutron.
Be accelerated like a proton in a synchrotron.