## Slide 1
SNS Luminescence Test Results
Beamtime 6.8-24.8.2012

## Slide 2: Idea
Use the photons emitted from the interaction between the beam particles and the residual gas to measure the beam profile.
Photon yield depending on
  Amount of protons per time
  Residual gas pressure
  Specific cross section
Beam
Lens
Detector

## Slide 3: Experiment Objectives
Residual gas composition measurements are not widely available, especially composition between cryostatic sections
The idea was to measure the photon yield, and estimate the gas composition from that
Due to preparation time only single photo anode Photomultiplier (PMT) were used

## Slide 4: Setup
Two PMT looking into the chamber
One is blinded
  This should give information about background signals

## Slide 5: Setup II
Warm section (replace-ment), where the PMT was attached to.
PMT

## Slide 6: Location
Last Cryomodule installed
PMT installed
Laser Wire

## Slide 7: RF Influence
To calibrate the ionization chambers, every 10 s a pulse is empty (everything on, no particles)
This was captured to have a reference of RF influence
No major influence could be noticed

## Slide 8: Results I
The initial idea of measure between Cryomodules could not be achieved
No information about the residual gas composition there was elsewhere available

## Slide 9: SNS Beam Structure Reminder
Variable
Macro-
Mini-
Micro-

## Slide 10: Measurement I
Use of a hybrid to analog sub-tract signals of open and closed PMT to reduce background
200 averages

## Slide 11: Theoretical predictions
Cross sections for proton beams measured at different labs
Calculation results in ~30 photons / 10 s for SNS

## Slide 12: Measurements II
Using a PMT counter module to count events
Countrate dependent on discriminator threshold
Much more counts than expected (~30 in 10 sec).
10 sec (=600 pulses)
1 pulse

## Slide 13: Theshold Determination
Counting behavior as expected
Threshold was chosen to high
Countrate even higher than measured

## Slide 14: Results II
The ministructure of the beam could be measured
The countrates achieved were much to high compared with theoretical predictions

## Slide 15: Conclusions
Countrate much higher then expected
As the empty pulse showed no significant counts, resulting assumption: Signal is beam related
Losses should be seen in blind PMT as well
Possible explanations:
  Electrons accelerated by the beam (H--Beam) induce additional luminescence in residual gas
  Inner-beam effects (under investigation)
  Losses / electrons induce luminescence elsewhere (window, residual gas)

## Slide 16: Possible Next Steps
Check if losses induce scintillation in window
  Setup ready, waiting for production beam
Cross-check results with proton beam
  Looking for opportunity using SNS proton beam
Setup position sensitive PMT & optics at the window and measure beam / background luminescence distribution within beam pipe
  Possibly cross-checking with H- / proton beam necessary

## Slide 17: Special Thanks to
V. Kamerdzhiev
A. Aleksandrov
W. Blockland

## Slide 19: Cable Quality

## Slide 20: PMT