## Slide 1: Non-invasive Profile Monitors for the ESS Linac
Cyrille Thomas (ESS)
Jacques Marroncle (CEA)
www.europeanspallationsource.se
10th February 2016

## Slide 2: ESS Linac and NPMs
BIF
IPM
BIF
Baseline:
BIF: 1st choice, everywhere it is possible
IPM: where BIF cannot deliver expected performance
Commissioning tool: profile and measured size for matching beam and lattice parameters to the model predicted nominal values, emittance measurement and measure emittance growth
Operation: monitor beam nominal parameters
Performance:
Profile measurement per pulse (14Hz)
Intra-pulse measurement bonus (highly desired)
NPM:
Transverse profile measurement up to full power
  Tuning and operation diagnostic

## Slide 3
BIF
Imaging system
Beam long. axis
Beam trans. axis
Residual gas – beam interaction: fluorescence
Residual gas – beam interaction: ionisation
IPM
Profile readout system
Beam long. axis
Beam trans. axis
Transverse beam profile
E (V/m)

## Slide 4: BIF: Design Studies
Source: selection of the gas specie(s), and spectral bandwidth
  COSY experiment and collaboration
Readout profile: selection of the imaging system
  Design of the optical system
  Selection of the single photon counting 1D/2D electronics
    PMT array, EMCCD camera, Intensified camera, etc.
    Coupling to fibre bundle for image transportation to camera
  Design calibration: monitor degradation
Radiation environment:
  Analysis of the perturbation to the BIF
  Maintenance plan
    Lifetime of the system in operation
      Can eliminate a initially favoured system
      operation cost
Radio-spectrometry of gas-proton interaction: measure of cross sections and fluorescence yields over UV-Vis-IR

## Slide 5: BIF: Source Studies
n=1
2
3
s
4
f
d
p
∞
n=1
2
3
s
4
f
d
p
∞
Lyman: n to 1: l=121nm to 91nm
Balmer: n to 2: l=656nm to 364nm
Paschen: n to 3: l=1875nm to 820nm
Case of hydrogen
P + H2 -> P + H2* + g

## Slide 6: ESS gas composition
Typical gas composition for ESS Linac:
  H2 (~<80%)
  CO (~10%)
  CO2 (~10%)
  N2 (~1%)
  Other
Note: cross section: CO > 7 H2  10% CO > 70% H2
Gas jet injection: potential scheme
  Studies to specify requirements for the vacuum chamber pumping performance and contamination of the cryo-cavities
  Studies to design a gas jet injector within the Diagnostics vacuum chamber fixed geometry
Measure spectra for each species
Measure fluorescence yield for each rays in the UV-Vis-IR spectra

## Slide 7: Optical Assembly
Conceptual design: reflective assembly
  No chromatic aberration
  Large spectral band for operation
  Radiation hard
1D profile assembly
  Field of view: 20mm: large sensor camera
  Fibre bundle must be tapered
    Could be preferred solution as it reduces the cost of the readout unit
  Or additional system must be used to reduce magnification to focus on ≈1mm diameter
    Solution to be compared with the cost of the fibre taper (performance and manufacturing cost)
Source
Optics
readout
Fibre Bundle
Camera, etc.
Spherical mirrors
M=1
NA=0.18
Camera, etc.

## Slide 8: Readout image unit
Single photon sensing unit:
  EMCCD camera: cheaper than Intensified cameras
  Cooled camera with ≈1e noise floor: CMOS camera, CCD camera designed for spectroscopy
    Choice based on dedicated tests with cameras to demonstrate sensitivity and acquisition schemes (single shot 3ms, 30 shots burst mode over 3ms, etc.)
  PMT Array: almost excluded, but not yet
    Radiation hard
    Highly sensitive to radiation
    Imposes a higher than 1 magnification: smaller field of view, less photons / pixel
    Inhomogeneity larger than 10% for each channel

## Slide 9: Management of Radiation Environment and Maintenance
Cameras, PMTs, most electronic readout image units will not live long or may not operate properly in the tunnel
  Solution: provide shielding:
    In the tunnel: >>1m3 steel is necessary for a camera to live and operate one year (average rate for 1W/m loss): this solution is not practical so it is likely to be abandoned
    In the stubs: enough shield for cameras to live and operate properly (based on shield calculation in stubs): this solution may work, but temperature and humidity condition must be investigated. It imposes almost the usage of a fibre bundle to transport image from optics image plane to the camera
    In the Klystron Gallery: best shield, but as above requires usage of a fibre bundle. In this case, the fibre must be radiation hard for low maintenance
tunnel
Stub
Gallery
linac
>105 /cm2 /s
<102 /cm2 /s
<1 /cm2 /s
Lifetime: 20’
Lifetime: >1 year
Lifetime: years

## Slide 10: LEBT Design, Production, Integration, Commissioning (2 NPMs)
Team work:
Requirements (Benjamin, Aurélien)
Conceptual design (Cyrille)
Mechanical assembly design (Thomas)
Vertical integration
Motor control (Klemen, Alex, Ander)
Camera control (Klemen, Hinko)
Software design (Klemen, Hinko)
Vacuum (Marcelo, Simone): gas composition

## Slide 11: MEBT/A2T Design
Interface viewport defined
Optical system: first possible design done
Readout system: selection of several potential cameras
Looking into fibre bundles and tapered optical coupler
Remaining: mechanical assembly
Calibration system: for monitoring lifetime of the system

## Slide 12: Cold Linac NPM
IK Collaboration (Jacques Marroncle – CEA Saclay)
Baseline: IPM
Feasibility Studies:
  Space Charge issues
    Development of numerical model of an IPM and benchmark collaboration: J. Storey, B. Dehning (CERN), K. Sato (J-Parc)
  Radiation environment issues
    Selection of the readout profile detector:
      Timepix3: 10MGrey, time and energy of particles
      Pure glass Fibre coupler + fibre bundle + single photon sensitive camera
      Etc.
  High Voltage cage design
    Constraint due to allocated space in LWU
    Constraint of the high voltage imposed by IPM model
Interface documentation
Maintenance plan

## Slide 13: IPM
Space Charge:
  3D IPM model (Matlab – Cyrille Thomas)
    Benchmarking (collaboration CERN, J-PARC, CEA, GSI)
    Study ESS case with ions/electrons
3D field analytical expression
Moving frame at bunch speed
Periodic bunch
Screen for stopping calculation
Equation of motion solve with Runge-Kutta non-linear solver
Set of particles with initial condition
Bunches motion
Q0Qb < 0 : focussing
Q0Qb > 0 : defocussing
Sign(x)

## Slide 14: IPM
ESS bunches:
2 GeV
sxyz = 2, 1.6, 2.1 (mm)
F = 25 kV / 0.1m
Uniform Zero speed at start time
No Magnetic Field taken into account
No differential cross-section
ESS bunches:
2 GeV
sxyz = 0.5, 0.4, 2.1 (mm)
F = 25 kV / 0.1m
Uniform Zero speed at start time
No Magnetic Field taken into account
No differential cross-section
Preliminary results from the IPM model

## Slide 15: IPM
Detector and readout
  Timepix3:
    Silicon pixel detector (2D image)
    Position, arrival time and energy discrimination
    10MGy radiation hard
    Readout ADC (1MGy): copper cable to transport digital signal to the readout electronic
    All pixels individually addressed, can be interrogated, and calibrated / tested
    50mm pixel size or less depending on model
    Integration into EPICS control requires a full driver development
  High sensitive Camera based solution:
    Radiation hard and luminescent glass fibre bundle (as used in streak cameras), fibre bundle, and high sensitive camera (photon counting)
    Radiation hard, but some of the elements should be demonstrated to be
    High sensitivity (single particle to produce > 500 photons)
    High resolution (fibre diameter)
    Integration into EPICS straightforward
  MCP:
    Radiation tolerant, but known to decay rapidly while operated under radiation
    High sensitivity
    High resolution
    Analogue signal to be transported outside the tunnel
    Requires expensive electronics (many channel ADC) and large effort in control integration
  Metallic strips
    Radiation hard
    Possibly not sensitive enough (single particle?)
    Expensive electronic similar to MCP
    Analogue signal to be transported outside the tunnel
    Large effort in control integration

## Slide 16: Last words
NPM project is half-way
  Conceptual designs in progress and soon to be achieved
  Detail designs production will follow, with the objective to deliver all the systems for the commissioning and 1MW on target milestone
In-kind collaboration starts in April with a challenging project for the Cold Linac NPM:
  Feasibility studies and full conceptual design within a year
  Full design the following year for matching the project milestone and high power commissioning

## Slide 17
Thank you to all our partners, collaborators and colleagues, for great support, good spirit and attitude facing the many difficult situations and challenges such a large scale project can generates.

## Slide 18: Additional slides

## Slide 19: Optical Assembly
2D profiles
  Advantage of real 2D image as opposed to 2x 1D profiles
  Divide the cost by almost 2, since only one imaging system is requires
  Feasibility and Reliability must be proven and demonstrated
Beam

## Slide 20: Transverse profile monitors in “high-energy” sections for Beam Commissioning and Operation
Spoke: 1 WS/NPM + 2 WS
  Monitors are located as close to the DTL-spoke interface.
  The aim is to match the beam from DTL to the periodic lattice structure of the spoke section.
  3 locations measurement method to measure emittance, beta, and alpha.
MB: 3 WS/NPM
  Monitors are located close to the spoke-MB interface.
  The aim is to match the beam from the spoke section to the periodic lattice structure of the MB sections.
  3 locations measurement method to measure emittance, beta, and alpha. (Quad scan is also possible, but want to avoid if possible to avoid losses in SC cavities.)
  The spoke-MB interface is the location of the frequency jump (352.21 MH => 704.42 MHz) and one of the most delicate location within the linac to preserve beam quality.
HB: 1 WS/NPM (placeholder)
  MB and HB have the same lattice structure so, ideally, no need of matching the beam. Thus, only 1 placeholder at the beginning in case problems are identified in MB.
  If we have monitors here, we rely on the quad-scan method.
Contingency: 3 WS
  Monitors are located close to the dogleg.
  The contingency section has the same lattice structure as MB and HB so, ideally, no need to match the beam at the HB-contingency interface.
  The aim is to send the matched beam to the dogleg and A2T.
  Dogleg: none
A2T: 1 WS/NPM
  Located in the middle of the raster magnets.
  The aim is to check the profile at this location as well as the cross-over point (180 deg separated).
Courtesy of Ryoichi Myamoto
Notes: I’m asking questions to Heine about the argument of having these 3 WS so I ‘ll send you more once I’ve got a reply from him.

## Slide 21: Table of RMS beam sizes at the locations of the transverse profile monitors
 | Sigx [mm] | Sigy [mm] | Sigz [mm]
SPK#1 | 2.4667 | 1.8439 | 2.3014
SPK#2 | 2.7995 | 2.1518 | 1.8631
SPK#3 | 2.1961 | 2.6880 | 2.0237
MB#1 | 2.6798 | 3.3621 | 1.4715
MB#2 | 2.4696 | 2.7221 | 1.3750
MB#3 | 2.7216 | 2.4701 | 1.5072
HB#1 | 2.2712 | 1.7594 | 1.2034
Contingency#1 | 1.3037 | 1.7387 | 3.2097
Contingency#2 | 1.1619 | 1.1730 | 4.5839
Contingency#3 | 1.1944 | 1.6647 | 6.0959
A2T#1 | 0.8658 | 0.6494 | 11.5544
From the 2015 lattice baseline