## Slide 1
INFN- Legnaro National Lab. Activity
on ESS DTL Diagnostics: BPM & BCM
Beam position monitor prototype development
Test bench Beam position monitor simulation
Beam Current monitor description
Conclusions
Marco Poggi INFN-LNL

## Slide 2
Pictures from SNS project
20 mm
30 mm
64 mm
DTL - BPM
Beam direction

## Slide 3
FIRST BPM PROTOTYPE

## Slide 4
SMA female

## Slide 5
SMP male
feedthrough
NEW BPM PROTOTYPE

## Slide 6
30 mm
54 mm
SMP female connector
SiO2 cable-F=0.09”

## Slide 8
How do make BPM simulation at best with wire test bench.
Identify working frequency: 704 MHz
Simulate gaussian bunch signal with wanted charge inside: 140 pC
Simulate periodic signal with 1/RFfreq using PSPICE: 2.84 ns
FFT of the periodic signal with PSPICE: 75 mA
@ 704 MHz
5) Measure of the wire resistance: 20 W
6) RF Signal generator to produce 75 mA on 20W @ 704 MHz: 1.5 Vrms

## Slide 9
Test bench
HP sampling osc.
RF generator

## Slide 10
Chart: 
| -7.5 | -6.25 | -5 | -3.75 | -2.5 | -1.25 | 0 | 1.25 | 2.5 | 3.75 | 5 | 6.25 | 7.5
out port 4 | 2.15 | 2.12 | 2.23 | 2.46 | 2.77 | 3.15 | 3.67 | 4.34 | 4.86 | 5.81 | 6.92 | 8.22 | 9.84
out port 2 | 11.69 | 9.3 | 7.54 | 6.2 | 5.11 | 4.23 | 3.55 | 3.03 | 2.42 | 1.98 | 1.64 | 1.34 | 1.14
mm
Input signal: 10 dBm @ 704 MHz

## Slide 11
Problems in simulation (1)
Reflection due to a mismatching between
signal generator and wire
Solution:
Resistance network to match characteristic
impedances between signal output and wire
Input (not tested yet)
signa
signal generator
output impedance
50 W
impedance matching network
39 W
27 W
0.1 mm wire + BPM under test (F =20mm)
20 W
318W
Impedance wire + bpm
Resistance wire
out signal
input wire
input wire

## Slide 12
Problems in simulation (2)
Wire signal has b = 1 (er = 1), but beam has 0.08 < b < 0.4
Solution:
Not really a solution but some hints from Shafer make us
thinking positive!!!
Chart: Proton Energy (MeV)
| 4.7300000000000004 | 10.739999999999998 | 30.77 | 63.349999999999994 | 85.47
704 MHz | 3.25 | 2.9 | 2.71 | 2.66 | 2.65
352MHz | 2.78 | 2.68 | 2.63 | 2.61 | 2.62
*dB/mm = 20log¦V/Vopp.¦/D(mm) D= wire diplacement from center

## Slide 13
BPM resolution request: at least 100 mm
Signal 10 dBm  7 mA @ 704 MHz
Pos. 0 port 4 = 3.63 --- port 2 = 3.63
Pos. 50 mm port 4 = 3.65 --- port 2 = 3.60
Pos. 100 mm port 4 = 3.66 --- port 2 = 3.56
Results in mV, variation in tens on microvolt (red digits) at 10 dBm, but increase in
resolution if signal increase.
Remember: nominal bunch charge 140 pC if current about 70 mA @704 MHz
…..and don’t forget Shafer gift!!
Signal 20 dBm  20 mA @ 704 MHz
Pos. 0 port 4 = 10.34 --- port 2 = 11.88
Pos. 50 mm port 4 = 10.43 --- port 2 = 11.79
Pos. 100 mm port 4 = 10.53 --- port 2 = 11.64
BPM position resolution

## Slide 14
Beam current monitor purchases
In vacuum sensor:
22 mm height, ID=30 mm, OD=102 mm
Front-end electronics
Cable
Low cutoff frequency: 3Hz High cutoff frequency: 1MHz Droop: <2%/ms Full scale output voltage: +/-10V Full scale range: 80 mA

## Slide 15
Conclusions
Some problem with first BPM prototype: feedthrough welding not so easy
New prototype design, with different feedthrough and different connection (no pin soldering)
First BPM simulations using wire test bench
List of purchases about the beam current monitor
Method to simulate better the beam characteristics: not a tutorial but only my ideas . IT WOULD BE NICE TO HAVE OTHER SUGGESTIONS ON THIS SPECIFIC TOPIC.
Thank you