## Slide 1: Present status of accelerator decommissioning project in Japan
High Energy Accelerator Research Organization
Kazuyoshi MASUMOTO
ARIA 2017/05/22-25 Lund

## Slide 2
In Japan, the clearance system which contains (1) the clearance level and (2) the control procedures for the radioisotope handling facilities and accelerator facilities was provided.
In this regulation, the definition and the handling rules for activated materials were also defined. This system is very difficult, expensive and time consuming for small facilities to perform.

## Slide 3
Number of accelerators in Japan (as of March 31, 2015)
Category | Total | Hospitals &Clinics | Educational Institutions | Reseach Institutions | Private Companies | Other Organizations
Total | 1,595 | 1,180 | 66 | 165 | 147 | 37
(%) | 100% | 74 | 4.1 | 10.3 | 9.2 | 2.3
Cyclotrons | 212(13.3) | 143 | 4 | 23 | 39 | 3
Synchrotrons | 40(2.5) | 9 | 3 | 23 | 4 | 1
Synchrocyclotrons | - | - | - | - | - | -
Linear Accelerators | 1,202(75.4) | 1024 | 25 | 55 | 65 | 33
Betatrons | 3(0.2) | - | 1 | 2 | - | -
Van de Graaff Accelerators | 38(2.4) | - | 14 | 23 | 1 | -
Cockcroft-Walton Accelerators | 73(4.6) | - | 16 | 27 | 30 | -
Transformer-type Accelerators | 17(1.1) | - | - | 9 | 8 | -
Microtrons | 9(0.6) | 4 | 3 | 2 | - | -
Plasma Generators | 1(0.1) | - | - | 1 | - | -

## Slide 4
How to keep the medical use.
-> Activated area should be defined for the decommissioning work in hospitals. But, it is very difficult to do in hospitals.
Then, we performed the neutron monitoring using activation detector, TLD, CR-39 in hospitals.
Calculation of activation has been also done by the coupling use of Monte Calro methods and Dchain-SP.

## Slide 5: Electron Accelerators in Hospitals
Purcased year
Survey results of 2008
Exchange rate is about 10 years interval.
!00 mashines were exchanged per 1 year in Japan.
Notes: This is the answerer from hospitals in 2008. It was found that exchange rate is 10 years interval.
!00 mashines were exchanged per 1 year.

## Slide 6: X-ray maximum energy
Energy (MeV)
Survey results of 2008

## Slide 7: Neutron measurement
Dose
  TLD；UD-813PQ4
  Track detector : CR-39
  conversion to fluence
Neutron fluence
  Au-foils (20mm in thickness)
  + covered with and without
  Cd-foils(0.5mm in thickness)

## Slide 8
0
45
90
Varian Clinac iX (15 MeV)
Dose monitoring of neutrons
X-ray beam -> 0 dgree
Neutrons -> isometric
135

## Slide 9: Comparison C/E
E decreased more rapidly than C (MCNP)
C/E<2
Bolt(SUS304) of floor
Co-60: 0.016 Bq/g
10y、400hr/y operation
C=0.029Bq/g
(C/E=1.8)
Neutron dose (15MeV, 0-degree)
Distance (cm)
Dose (mSv/h)

## Slide 10: Neutron monitoring at 18 MeV
Beam exit
Phantom water
Varian Clinac 2300 C/D 18MeV
1 min. irradiation
Run-1:electron(without target)
Run-2:Ｘ-ray（window:40cm×40cm）
Run-3:Ｘ-ray（with Multi lleef）
Run-4:Ｘ-ray（with steel filter）
Run-5:Ｘ-ray（with Pb filter）
Run-6:Ｘ-ray（window :0.5cm×0.5cm）

## Slide 11: Measurement position
Front viw
Side viw
27
29
30
28
Head beneath

## Slide 12: Thermal neutron
← Head parts →
High fluence at No. 22, 24 and 26
Electron irradiation -> No neutros

## Slide 13: Neutron distribution
Au、TLD、CR-39 were set for 1d.
Total operation time=11.1min

## Slide 14: Thermal neutron distribution
Chart: 
| 1.0 | 2.0 | 3.0 | 4.0 | 5.0 | 6.0 | 7.0 | 8.0 | 9.0 | 10.0 | 11.0 | 12.0 | 13.0 | 14.0 | 15.0 | 16.0 | 17.0 | 18.0 | 19.0
熱中性子 | 1082.3257 | 371.876 | 1010.1706 | 1293.2404 | 7898.2023 | 11983.288 | 36693.6161 | 58939.5718 | 52345.7107 | 61026.5178 | 27441.119 | 32597.4296 | 20919.413 | 62941.4017 | 69057.9294 | 63307.7273 | 74858.085 | 86136.4738 | 72476.9685
熱外中性子 | 80.4123 | 136.3513 | 106.2841 | 158.0277 | 284.9392 | 471.9853 | 1284.4991 | 3579.3963 | 3120.347 | 5712.0705 | 1569.0888 | 1722.2218 | 1121.2272 | 5378.5342 | 6091.4068 | 4984.5139 | 4437.3606 | 16963.8493 | 11145.8444
Maze
ceiling
floor
Irradiation room
0.95 m hight
cm-2s-1

## Slide 15: Our results
X-ray energy(MeV) | Neutron flux
6 | ND
10 | Thermal:1X103cm-2s-1、Epithermal:1X102cm-2s-1 Thermal 106 cm-2 (d)、10 10 cm-2 (10y)
15 | Thermal:2X104cm-2s-1、Epithermal:2X103cm-2s-1 Thermal 107-8cm-2(d)、10 11 cm-2(10y)
18 | Thermal:5X104cm-2s-1、Epithermal:4X103cm-2s-1 Thermal 108 cm-(d) 、10 11 cm-2(10y)

## Slide 16: Activated area assigned by NSA
6 MeV :No activation
10 MeV or less : target and collimators
15 MeV or less : target, collimators and shielding materials
Non activated region -> Surroundings and Building

## Slide 17: PET cyclotron
Peak: 2005
After 2007, delively system was prevailed.
Now, decommisiong case was increased.
Purchsed Year
Number

## Slide 18: Cyclotron
(a) Proton
(b) H- I ion
without self-shield
Self-shield

## Slide 19: RI production and reaction
Produced RI | Reaction | Site number
18F | 18O(p,n)18F 20Ne（d,α）18F 18O(p,n)18F, 20Ne(d,α)18F両方 | 82 1 4
11C | 14N(p, α)11C | 29
15O | 14N(d,n)15O 15N(p,n)15O | 17 5
13N | 16O(p,α)13N | 15
Others | 64Ni(p,n)64Cu | 1
18F is major RI
11C(33%)、15O(25%)、13N(17%)

## Slide 20: Source of neutrons of PET-cyclotron
Neutrons are mainly produced by the 18O(p,n)18F reaction.
Neutron yield from the 14N(p,α)11C reaction was 8.9% of 18F-production
Therefore, neutron fluence is mainly dependent on the produced amount of 18F.

## Slide 21: Measurement of neutron fluence
Target:water（O-18）
Production：F-18
H- ion ,18MeV
Current：21μA
Irradiation：63min
SHI HM-18

## Slide 22: Neutron distribution
cm-2s-1×105
Near target：106 cm-2s-1
Opposite side：2×105 cm-2s-1
Inside the target box：107 cm-2s-1
Bolt(Brass,Cu:Zn=65:35) was sampled from the wall near the target
Averaged neutron flux of operation
Cu-64 : 4.5×105 cm-2s-1 (2days)
Zn-65 : 6.3×105 cm-2s-1(2years)

## Slide 23: Results
Near target: 107cm-2s-1
Inside the vault: 105-6 cm-2s-1
Estimation of induced activity in concrete
  Co :10ppm
  Operation:500h/year
  Neutron flux= 4.6×105
Specific activity of Co-60 became 0.1Bq/g.

## Slide 24
Concrete core sampling
Core sample（diameter: 5cm, length50cm）
Ge-detector (ORTECGMX-20195)

## Slide 25
Comparison of C/E of Floor concrete
2001/8～2012/12
2012/12～2013/12

## Slide 26
2001/8～2012/12
2012/12～2013/12
Comparison of C/E of wall concrete data
2001/8～2012/12

## Slide 27: Activation of concrete
Major RI：H-3, Co-60, Eu-152, Sc-46, Cs-134
Products of Thermal neutron capture reaction
Maximum activity observed at 10cm in depth.
At 30cm in depth, activity decreased about one order.
Mn-54, Na-22 (Fast neutron reaction products) are observed in concrete near the target. Build up effect was not observed.
Activity of Co-60 sampled from surface was below 0.1Bq/g.

## Slide 28: Outside of self-shield
CR-39, TLD were set for one month
Thermal neutron dose rate on operation：several μSv/h
Thermal neutron flux on operation： under 102cm-2•s-1
Activation by neutrons is negligibly small.

## Slide 29: Metal samples
Metals used for more than 5 years were collected.
Fe, SUS, Brass, Al
Several g to several 100 ｇ
Measurement for several 104 second (0.5 d) by Ge-detector: No activity was detected

## Slide 30: Zoning of PET-cyclotrons
Without self-shield
  Inside the vault is activated area
Self-shield type
  No activated area : Outside the self-shield
  Activated area : Inside the self-shield
Toyoda -> Activation of cyclotron
Matsumura -> Activation of concrete

## Slide 31: Conclusion
Neutron measurement suring operation is very effective to estimate activation of surrounding matereials.
To estimate induced activity, operation record is also important.
We will tried to define the activation zone in the facilities of SR, Tandem, Particle therapy.

## Slide 33: Estimation of activity
Neutron fluence measurement under operation
Operational record
Material analysis
↓
Estimation of activity
And/or
Monte Carlo calculation by neutron transport + Induced activity

## Slide 34: Measurement for decommissioning
Sampling -> Activity measurement
In-situ measurement using Ge-detector
->activity evaluation
Surface dose rate monitoring
We obtain the relationship between activity and dose
Decommissioning work can be performed by using survey-meters