## Slide 1: In-situ determination of residual specific activity in activated concrete walls of a PET-cyclotron room
Hiroshi Matsumura, Akihiro Toyoda, Kazuyoshi Masumoto
Toshiyuki Yagishita, Takayuki Nakabayashi, Hiroyuki Sasaki, Kazuhiro Matsumura, Yoshiyuki Yamaya
Japan Environment Research Co., Ltd., JAPAN
High Energy Accelerator Research Organization (KEK), JAPAN
concrete walls
In-situ determination
Yoshiharu Miyazaki
The Medical and Pharmacological Research Center Foundation, JAPAN

## Slide 2: Typical image of a PET-cyclotron room
PET-cyclotron body
Concrete walls
of PET-cyclotron room
Target for RI production

## Slide 3: Typical image of a PET-cyclotron room
PET-cyclotron body
Target for RI production
Activated by neutrons
For accelerator decommissioning, specific activity in the concrete walls has to be determined in order to identify the contaminated part.
Concrete walls
of PET-cyclotron room
eg. Eu-152: 0.1 Bq/g
Co-60: 0.1 Bq/g
Purpose of this study is to establish methodology of specific activity determination in concrete walls.
Clearance
limit

## Slide 4: Methods of specific activity determination
1. Off-site measurement with a Ge detector after core boring
2. In-situ measurement with a Ge detector
3. In-situ measurement with a NaI survey meter
Concrete wall
Concrete wall
Concrete wall
Depth profile
Activation mapping
Easier method

## Slide 5: Examination site
~15 years ~2 years ago
Accelerator:
Purpose of use:
Operation duration:
Room size:
Measurement:
The Medical and Pharmacological Research Center Foundation in Hakui, Ishikawa, Japan
12-MeV (0.1 mA) proton cyclotron (NKK/Oxford)
Production of radiopharmaceuticals for PET
Facility:
Operation end:
5.6 m x 5.2 m x 2.7 mH
January 2017 at empty room

## Slide 6: Methods of specific activity determination
2. In-situ measurement with a Ge detector
3. In-situ measurement with a NaI survey meter
Concrete wall
Concrete wall
Concrete wall
1. Off-site measurement with a Ge detector after core boring
Depth profile

## Slide 7: Core boring locations
CW1
CW5
CW2
CW4
Shield
door
South wall
North wall
East wall
West wall
Target
Cyclotron

## Slide 8: Depth profile of the specific activity
CW1
Constant specific activity
Turned to be decrease
FIG. Typical example of depth profile of specific activity (CW1).
Attainable region of gamma rays agreed with the constant region of specific activity.
Good for direct measurement at concrete surface
Independent of nuclide
Concrete depth (cm)
Specific activity (Bq/g)
Efficiency was calculated by ISOCS.

## Slide 9: Methods of specific activity determination
2. In-situ measurement with a Ge detector
3. In-situ measurement with a NaI survey meter
Concrete wall
Concrete wall
Concrete wall
1. Off-site measurement with a Ge detector after core boring
Merit
Measurable
at many locations
Demerit
Measurable
at a few locations

## Slide 10: Measured locations
S1
S3
S2
S5
S6
W5
W3
W2
W1
E1
E3
E4
E5
Shield
door
Target
South wall
East wall
West wall
Cyclotron
North wall
E2
W4
S7(high)
S4
S8(low)
N2(high)
N1
20 locations
6.5-cm-thick Pb shield

## Slide 11: Specific activity distribution
Eu-152, Co-60, Cs-134, Eu-154, and Mn-54 were determined.
Major nuclides are Eu-152 and Co-60.
We could obtain detailed distributions of the specific activity.
Specific activity (Bq/g)
FIG. Distributions of specific activity
Location name
Eu-152
Co-60
Cs-134
Eu-154
Mn-54

## Slide 12: Relative specific activity
Specific activity ratio
of each nuclide for Co-60
Eu-152/Co-60 = 1.59
Cs-134/Co-60 = 0.239
Eu-154/Co-60 = 0.165
Mn-54/Co-60 = 0.138
Co-60/Co-60 = 1.00
Although specific activity of Eu-152 was higher than Co-60, dose rate from Co-60 became higher than Eu-152.
The activity ratio was independent of location.
Specific activity ratio of each nuclide for Co-60 was obtained.
Location name
Eu-152/Co-60
Co-60/Co-60
Cs-134/Co-60
Eu-154/Co-60
Mn-54/Co-60

## Slide 13: Methods of specific activity determination
2. In-situ measurement with a Ge detector
3. In-situ measurement with a NaI survey meter
Concrete wall
Concrete wall
Concrete wall
1. Off-site measurement with a Ge detector after core boring
Require great care
Demerit
Merit
Easier

## Slide 14: Measured locations
S1
S3
S2
S5
S6
W5
W3
W2
W1
E1
E3
E4
E5
Shield
door
Target
South wall
East wall
West wall
Cyclotron
North wall
E2
W4
S7(high)
S4
S8(low)
N2(high)
N1

## Slide 15: Contact dose rate vs specific activity
Contact dose rate (μSv/h) by NaI survey meter
Co-60 specific activity (Bq/g)
by Ge detector
Co-60 specific activity is correlated with contact dose rate by NaI survey meter.
FIG. Correlation curve between dose rate and specific activity
Eu-152/Co-60 = 1.59
Cs-134/Co-60 = 0.239
Eu-154/Co-60 = 0.165
Mn-54/Co-60 = 0.138
Measured dose rate
Specific
activity
of Co-60
Specific activity
of Eu-152, Cs-134,
Eu-154, and Mn-54

## Slide 16: Reproducibility of specific activity from dose rate
Specific activity (Bq/g)
FIG. Distributions of specific activity
Location name
Eu-152
Co-60
Cs-134
Eu-154
Mn-54
TOTAL
Results determined by using Ge detector

## Slide 17: Reproducibility of specific activity from dose rate
Specific activity (Bq/g)
FIG. Distributions of specific activity
Location name
Eu-152
Co-60
Cs-134
Eu-154
Mn-54
TOTAL
Put Co-60 results estimated from dose rates

## Slide 18: Reproducibility of specific activity from dose rate
Specific activity (Bq/g)
FIG. Distributions of specific activity
Location name
Put all results estimated from dose rates
reproduced well from dose rate!
Eu-152
Co-60
Cs-134
Eu-154
Mn-54
TOTAL

## Slide 19: Conclusion
We obtained an efficient method
for specific activity determination
in concrete wall of PET cyclotron.

## Slide 21: Contribution from surrounding wall
No shield
6.5-cm thick Pb shield
~94%
~50%
1
3
2
4
6
5
6.5-cm thick Pb shield is important for avoiding contribution from other walls.
Contribution ratio from each wall in gamma-ray counting.

## Slide 23
石膏ボード（約1.5cm厚）切断
石膏ボードはがし
モルタルはがし
測定場所マーキング

## Slide 24: Gamma-ray spectrum
Eu-152
Co-60
Eu-154
Cs-134
Na-22
Mn-54