## Slide 1: The Bifrost spectrometer
The three options – Budget and scope
Rasmus Toft-Petersen
Technical University of Denmark

## Slide 2: Outline
Technical requirements and the choices made
Option 1 – Cost Category B – 12 M€
Option 2 – world leader – 14.1 M€
Option 3 – primary scope – 16.4 M€
Summary and our position

## Slide 3: Main requirements
1.7 Å bandwidth
Polychromatic flux > 1010 n/s/cm2
PSC: Pulse duration 0.1-3.1 ms
Beam size: 2x2 cm
Divergence < 0.75 degrees
Flexible E-resolution, 30-170 μeV
Angular resolution 0.7 degrees
10 analysers pr Q-channel
Order sorting chopper
120 degree angular coverage
Moveable evacuated tank
Translateable filter+collimator
Temperatures below 50 mK
Magnetic fields above 15 T
Pressures above 2 GPa
Primary spectrometer
Secondary spectrometer
Sample environment

## Slide 4: Main requirements
1.7 Å bandwidth
Polychromatic flux > 1010 n/s/cm2
PSC: Pulse duration 0.1-3.1 ms
Beam size: 2x2 cm
Divergence < 0.75 degrees
Flexible E-resolution, 30-170 μeV
Angular resolution 0.7 degrees
10 analysers pr Q-channel
Order sorting chopper
120 degree angular coverage
Moveable evacuated tank
Translateable filter+collimator
Temperatures below 50 mK
Magnetic fields above 15 T
Pressures above 2 GPa
Primary spectrometer
Secondary spectrometer
Sample environment

## Slide 5: STAP recommendations
Take activation into account – consider sample changer. 
Cover 90 degrees scattering angle from Day 1 
Divergence jaws for small samples 
Investigate saturation from Bragg peaks
Upgrade priority: 1. High field magnet, 2. PA, 3. Order sorting chopper. 
Avoid building a tank to be disposed of 

## Slide 6: Changed requirements for all options
Flexible E-resolution, 30-170 μeV
Angular resolution 0.7 degrees
6 analysers pr Q-channel
90 degree angular coverage
Moveable evacuated tank
Translateable filter+collimator
Temperatures below 50 mK
Magnetic fields above 15 T
Pressures above 2 GPa
Secondary spectrometer
Sample environment
Strategy:
Full scope primary spectrometer
Do not sacrifice resolution
Reduce tank size as full scope tank will cost 1.5 M€
Sacrifice detection area and sample environment

## Slide 7: Guide: Design after shielding requirements
E02 - Long section –
Need low price pr m
D03 – short section with fast neutron
expensive

## Slide 8: Guide: Design after shielding requirements
E02 - Long section –
Need low price pr m
D03 – short section with fast neutron
expensive
Guide boundary conditions:
1) Metallic substrates before bunker wall
2) Borcron after bunker wall
3) Loose LoS to avoid too heavy a shutter and for background
4) Use m = 1 where possible due to gamma radiation

## Slide 9: Guide proposals
A straight double ellipse with static prompt pulse suppression
Kinked double ellipse, with kink before bunker wall
A S-bend m = 3 guide, loosing LoS twice – once before bunker wall ending up in m = 1 straight section
Long m = 1 section for all options

## Slide 10: Comparison

## Slide 11: Optimizing the feeder for early procurement
Benchmark with flat moderator and pinhole for a chopper
Source
Sample
Pinhole
Feeder
Elliptical guide

## Slide 12: Robust guide price
Feeder
Two ellipses – 2x5 cm -> 8x8 cm
Metallic substrate < 30 m and Borcron elsewhere
Price for the guide (Swiss neutronics quote): 2.160 k€
Price for vacuum housing: 630 k€
Monolith insert infrastructure: 100 k€
Special housing near monolith (steel, sapphire filters): 70 k€
Price for guide installation (Included in quote): 100 k€
Monolith gamma shutter customization: 80 k€
Total price for the system: 3167 k€
Almost independent of our choice of guide – m = 1 requires expansion of guide in all cases, and substrates are the cost drivers for that part….

## Slide 13: Kinked guide in D03: More steel
Lose LoS
outside bunker,
fast neutron beam
stop 4 m3 steel
Assume addtional steel
Before LoS is lost at 50 m:
an additional 10 m3
We still end up below ODIN
values due to our double slit before bunker wall
Old budget: 30 m3
New budget 45 m3 steel
to the price of 500 k€
Increased by 150 k€

## Slide 14: Steel considerations in D03
Fast neutron shielding relevant < 50 m from moderator…
Stuarts work on ODIN (straight):
Between 25-35 cm steel, 30 cm of concrete
Pending calculations, we estimate a little less for our guide up to lossing LoS 40 m3
Add 5 cm of steel for gammas: An additional 5 m3 of steel
Fair kinked guide estimate: 45 m3 of steel
Increased by 150 k€ compare

## Slide 15: D03 guide summary
Kinked guide is the better performer, but expensive to shield and we might need metal substrate outside the bunker
PPS-guide performs 40 % worse than kinked but is cheaper to shield, possibly 3-400 k€ in total. Optimisation pending.
S-bend is requires little fast neutron sheilding, but does require steel in D03 to shield for hard gammas. Possibly the cheapest option: chosen if performance is acceptable (Brilliance transfer, divergence, beam profile, expansion into m = 1)

## Slide 16: Concrete
Chart: 
Length [m]
Thickness [m]
Few fast neutrons
Low gamma-flux
We hope to get away with
using only concrete
50 cm thickness for 100 m E02 hall: 200 m2
Extra concrete near D03 hall
And cave: 30 m2
Floor and shielding in D03: 70 m2
Due to craning limitations we will end up with complex elements to avoid streaming and limit
the number of elements needed:
300 m3 x 1000 €/ m3 = 300 k€

## Slide 17: Lead near cave
Chart: 
Length [m]
Thickness [m]
Normal steel magnetic, so:
Lead versus stainless steel?
Lead: 90 k€/m3
Stainless: 40 k€/m3
Similar shielding power pr Euro
We don’t want activation,
and ferromagnetism!
Stainless steel not cost
effective and lead is
10 times better than
concrete for hard gammas!
We chosea lead+concrete
solution:
5 cm lead on the last
15 m’s
Price: 250 k€

## Slide 18: Labor for shielding
1 crane
1 crane
1 crane
Mobile gantry
cranes
We estimate
3 people take
3 hours on average
Pr element, from
The load-off point
to secure bolted
position,
Two FTE’s in total for shielding

## Slide 19: The cave
We aim for a large cave of reinforced concrete. This can shield for primary beam scattered in 4pi or converted into 2 meV gammas
1 m thick concrete: 450 k€
3 mm of Boralcan plates: 200 K€
Manpower: 100 k€

## Slide 20: Total shielding costs
Item | Price [k€]
Design – 1.75 FTE | 190
Bunker feedthrough | 30
Mirrobor in bunker | 8
Steel – D03 | 500
Concrete – beamline | 300
Lead near cave | 250
Chopper pit Concrete+Boralcan) | 40
Cave (Boralcan + concrete) | 600
Beamstop + vacuum tube | 250
Tungsten | 80
Installation – 3 FTE | 260
Total | 2508
Material costs
for beamline:
1050 k€
Cave costs is
excluding cabling

## Slide 21: The chopper system
1 fast double-disc PSC chopper
3 slow frame overlap choppers
Total price: 910 k€ including MP, Vac, cooling, Racks, comm
Option: A fast order sorting chopper

## Slide 22: Pulse shaping chopper
Item | Cost [k€]
Double disc chopper, face-to-face, 700 mm, speeds above 196 Hz | 280
Chopper integration module | 33
Vacuum | 5
Chopper cooling | 10
Master rack | 15
Chopper drive | 40
Installation | 17
Integration | 15
Commissioning | 10
Power system | 5
Sum | 405
ESS estimate:
405 K€
LLB Estimate: 400 k€

## Slide 23: The three slow choppers
Item | Cost [k€]
3 single disc choppers 700 mm, max 70 Hz | 221
Chopper integration modules | 100
Vacuum | 15
Chopper cooling | 20
Master rack x 2 | 30
Slave rack | 7
Drives + power | 65
Installation | 17
Integration | 15
Commissioning | 10
Power system | 5
Sum | 505
Total for all choppers:
910 k€
Workforce:
85 k€
Second slow chopper double disc?
Add 120k€

## Slide 24: Detector system
Our costing of the secondary spectrometer tank
uses the PSI quotes as a starting point.
We are considering separate detectors
to allow for rowland geometry and more flexibility

## Slide 25: Option 1 – sacrifices
No sample environment
To afford detectors at all, we need to build a small tank – 22 degrees
Tank unevacuated
We save money on filter and collimators but costs are not linear
We develop and buy equipment to be thrown away upon upgrades
Detector voltage needs to be ramped doen before moving the tank – substantial dead time

## Slide 26: Cost Category B detector system
For cost category B, we can afford a simple tank in air covering 21.5 degrees,
possibly more if we redesign everything after CC-B
Item | Price [k€]
Detectors | 110
PG | 100
Mounts | 50
Electronics | 180
Cross talk | 70
Detector tank | 120
Frame+rail+motor | 100
Detector safety | 50
Beam monitor/Camera/preamps | 120
Total | 900
Labor for design, installation and commissioning: 330 k€

## Slide 27: Instrument specific cost category
Item | Price [k€]
Management labor incl travel | 1260
Be-filter | 100
Radial collimator | 70
Light shutter | 150
Divergence jaws | 100
Installation + commissioning | 50
Total | 1730

## Slide 28: Sample environment, motion control, instrument infrastructure and PSS
Sample table: 108 k€
Cryofurnace: 50 k€ Activation sticks: 60 k€
Total: 220 k€
16 slit motors (divergence jaws)
2 sample rotation motors
2 translation stages
2 oscillation motors
1 tank motor
Total:180 k€, with
the tank motor included in spectrometer budget
Design of infrastructure: 27 k€
Cranes: 40 k€
Media distribution: 80 k€
False floor + He recovery: 40 k€
Control Hutch: 150 k€
Installation: 173 k€
Total: 500 k€
Total PSS system (NMX):
120 k€
Sample environment in CC-B budget:
Infrastructure in all budgets
Motion control, all budgets

## Slide 29: WBS breakdown
 | Project Management and integration | Design | Procurement and fabrication | Install-ation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 2.078.750 | € 259.200 | € 0 | € 2.078.750 | € 448.200 | € 2.526.950
Neutron Optics | € 0 | € 27.000 | € 3.040.000 | € 100.000 | € 0 | € 2.960.000 | € 207.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 0 | € 210.000 | € 3.840 | € 3.840 | € 210.000 | € 7.680 | € 217.680
Detector and beam monitor | € 0 | € 175.000 | € 880.000 | € 54.000 | € 108.000 | € 880.000 | € 337.000 | € 1.217.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion contro | € 17.400 | € 60.000 | € 112.000 | € 6.000 | € 34.000 | € 77.000 | € 152.400 | € 229.400
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 420.000 | € 21.600 | € 27.000 | € 480.000 | € 1.248.600 | € 1.728.600
Instrument Infrastructure | € 0 | € 27.000 | € 300.000 | € 172.800 | € 0 | € 300.000 | € 199.800 | € 499.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.169.281
Phase I |  |  |  |  |  |  |  | € 250.000
TOTAL | € 1.277.400 | € 498.000 | € 7.931.750 | € 709.840 | € 212.840 | € 7.916.750 | € 2.713.080 | € 12.049.111
Man year | 12 | 5 | 0 | 8 | 2 |  |  | 27

## Slide 30: WBS breakdown
 | Project Management and integration | Design | Procurement and fabrication | Install-ation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 1.905.500 | € 259.200 | € 0 | € 1.905.500 | € 448.200 | € 2.353.700
Neutron Optics | € 0 | € 27.000 | € 3.040.000 | € 100.000 | € 0 | € 2.970.000 | € 197.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 0 | € 150.000 | € 3.840 | € 3.840 | € 150.000 | € 7.680 | € 157.680
Detector and beam monitor | € 0 | € 200.000 | € 1.010.000 | € 54.000 | € 108.000 | € 1.010.000 | € 362.000 | € 1.387.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion contro |  | € 0 | € 116.000 | € 15.000 | € 50.000 | € 116.000 | € 65.000 | € 181.000
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 420.000 | € 21.600 | € 27.000 | € 480.000 | € 1.248.600 | € 1.728.600
Instrument Infrastructure | € 0 | € 27.000 | € 300.000 | € 172.800 | € 0 | € 300.000 | € 199.800 | € 499.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.157.000
Phase I |  |  |  |  |  |  |  | € 280.000
TOTAL | € 1.260.000 | € 463.000 | € 7.832.500 | € 718.840 | € 228.840 | € 7.862.500 | € 2.640.680 | € 11.955.180
Man year | 12 | 5 | 0 | 8 | 2 |  |  | 27
Chart: Cost category B budget
| Shielding | Neutron Optics | Choppers | Sample Environment | Detector and beam monitor | DMSC | Motion control | Instrument specific Tech equipment | Instrument Infrastructure | Vacuum | PSS | Contingency | Phase I
 | 2353700 | 3167000 | 923400 | 157680 | 1387000 | 0 | 229400 | 1728600 | 499800 | 0 | 120000 | 1162323.8 | 280000

## Slide 31: WBS breakdown
 | Project Management and integration | Design | Procurement and fabrication | Install-ation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 1.905.500 | € 259.200 | € 0 | € 1.905.500 | € 448.200 | € 2.353.700
Neutron Optics | € 0 | € 27.000 | € 3.040.000 | € 100.000 | € 0 | € 2.970.000 | € 197.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 0 | € 150.000 | € 3.840 | € 3.840 | € 150.000 | € 7.680 | € 157.680
Detector and beam monitor | € 0 | € 200.000 | € 1.010.000 | € 54.000 | € 108.000 | € 1.010.000 | € 362.000 | € 1.387.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion contro |  | € 0 | € 116.000 | € 15.000 | € 50.000 | € 116.000 | € 65.000 | € 181.000
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 420.000 | € 21.600 | € 27.000 | € 480.000 | € 1.248.600 | € 1.728.600
Instrument Infrastructure | € 0 | € 27.000 | € 300.000 | € 172.800 | € 0 | € 300.000 | € 199.800 | € 499.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.157.000
Phase I |  |  |  |  |  |  |  | € 280.000
TOTAL | € 1.260.000 | € 463.000 | € 7.832.500 | € 718.840 | € 228.840 | € 7.862.500 | € 2.640.680 | € 11.955.180
Man year | 12 | 5 | 0 | 8 | 2 |  |  | 27
Chart: Labor - 2.625 M€ - 25 %
| Project Management and integration | Design | Installation | Cold commissioning
 | 1277400 | 498000 | 709840 | 212840

## Slide 32: Evaluation of option 1 from Bifrost team
Overview studies is essential for Bifrost with its polychromatic beam. The flux allows you to find the unexpected but only if you are looking!
Option 1 wastes money and manpower on an tank that will be replaced as soon as possible
We compete not only on flux, but also on sample environment, simplicity and software. Option 1 will not be world leading…

## Slide 33: Cost category worst case
ThALES monochromatic flux: 3*108 n/s/cm2
BIFROST 3MW polychromatic flux: 1010 n/s/cm2
similar flux at specific wavelengths
ThALES multiplexing option, assume the worst: 0.75 m2 of PG.
Gain factor reduced to 8. Assume BIFROST background could be higher, and reduce gain to 2-4. Worst case: ThALES: 2 meV dynamic range at three Ei settings. BIFROST: Four A4 settings 13.8 M€ BIFROST excluding phase I:
Gain factor 12+ With sample environment limitations as they are, this is what we need to call ourselves world leader!

## Slide 34: Option 2
Option 2 covers 90 degrees of scattering angle, and is able to cover a Brillouin Zone in most systems using a cold beam.
Option 2 includes a dilution stick, which allows us partly to address the quantum magnetism and unconventional SC communities
The upgrade path of option 2 is straightfoward and cost effective in terms of scope/price ratio. We will aim to leave space for a single analyzer E > 5 meV for order sorting

## Slide 35: Option 2 – sacrifices
Limited magnet availability – we do not directly address the magnetism and SC communities
Bifrost can do small samples – options for very high magnetic field not there.
No polarization analysis – definite loss of scope and excellence
V. O. Garlea, et. al,* PRL 98, 167202 (2007)
P. Dalla Piazza, Nature Physics
(2014)

## Slide 36: Detector system for option 2
Here, we cover 90 degrees as recommended by the STAP council.
Item | Price [k€]
Detectors | 450
PG | 400
Mounts | 120
Electronics | 400
Cross talk | 100
Detector tank | 500
Frame+rail+motor | 100
Detector safety | 50
Beam monitor/Camera/preamps | 120
Total | 2240
Labor for design, installation and commissioning: 360 k€

## Slide 37: Instrument specific + SA for option 2
Item | Price [k€]
Management labor incl travel | 1260
Be-filter | 250
Radial collimator | 100
Light shutter | 150
Divergence jaws | 100
Installation + commissioning | 50
Total | 1910
For the sample environment, we include a dilution stick
For 150 k€ and 100 k€ for development labor costs (SAD estimate)

## Slide 38: WBS breakdown
 | Project Management and integration | Design | Procure-ment and fabrication | Installation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 2.078.750 | € 259.200 | € 0 | € 2.078.750 | € 448.200 | € 2.526.950
Neutron Optics | € 0 | € 27.000 | € 3.040.000 | € 100.000 | € 0 | € 2.960.000 | € 207.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 108.000 | € 360.000 | € 3.840 | € 3.840 | € 360.000 | € 115.680 | € 475.680
Detector and beam monitor | € 0 | € 207.000 | € 2.240.000 | € 54.000 | € 108.000 | € 2.240.000 | € 369.000 | € 2.609.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion control | € 17.400 | € 60.000 | € 112.000 | € 6.000 | € 34.000 | € 77.000 | € 152.400 | € 194.990
Instrument specific Tech equipment | € 1.356.000 | € 0 | € 600.000 | € 21.600 | € 0 | € 660.000 | € 1.317.600 | € 1.977.600
Instrument Infrastructure | € 0 | € 27.000 | € 310.000 | € 172.800 | € 0 | € 310.000 | € 199.800 | € 509.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.375.486
Phase I |  |  |  |  |  |  |  | € 250.000
TOTAL | € 1.373.400 | € 638.000 | € 9.631.750 | € 709.840 | € 185.840 | € 9.616.750 | € 2.922.080 | € 14.129.906
Man year | 12 | 7,4 | 0 | 8,2 | 2,23 |  |  | 29,83

## Slide 39: WBS breakdown
 | Project Management and integration | Design | Procure-ment and fabrication | Installation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 1.905.500 | € 259.200 | € 0 | € 1.905.500 | € 448.200 | € 2.353.700
Neutron Optics | € 0 | € 27.000 | € 3.020.000 | € 100.000 | € 0 | € 2.950.000 | € 197.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 100.000 | € 300.000 | € 3.840 | € 3.840 | € 300.000 | € 107.680 | € 407.680
Detector and beam monitor | € 0 | € 200.000 | € 2.390.000 | € 54.000 | € 108.000 | € 2.390.000 | € 362.000 | € 2.752.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion control |  | € 0 | € 116.000 | € 15.000 | € 50.000 | € 116.000 | € 65.000 | € 181.000
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 600.000 | € 21.600 | € 27.000 | € 660.000 | € 1.248.600 | € 1.908.600
Instrument Infrastructure | € 0 | € 27.000 | € 310.000 | € 172.800 | € 0 | € 310.000 | € 199.800 | € 509.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.355.550
Phase I |  |  |  |  |  |  |  | € 300.000
TOTAL | € 1.260.000 | € 563.000 | € 9.532.500 | € 718.840 | € 228.840 | € 9.562.500 | € 2.740.680 | € 13.978.730
Man year | 12 | 7 | 0 | 9 | 2 |  |  | 30
Chart: Bifrost budget - WBS breakdown
| Shielding | Neutron Optics | Choppers | Sample Environment | Detector and beam monitor | DMSC | Motion control | Instrument specific Tech equipment | Instrument Infrastructure | Vacuum | PSS | Contingency | Phase I
 | 2353700 | 3167000 | 923400 | 407680 | 2752000 | 0 | 229400 | 1908600 | 509800 | 0 | 120000 | 1360873.8 | 300000

## Slide 40: WBS breakdown
 | Project Management and integration | Design | Procure-ment and fabrication | Installation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 1.905.500 | € 259.200 | € 0 | € 1.905.500 | € 448.200 | € 2.353.700
Neutron Optics | € 0 | € 27.000 | € 3.020.000 | € 100.000 | € 0 | € 2.950.000 | € 197.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 100.000 | € 300.000 | € 3.840 | € 3.840 | € 300.000 | € 107.680 | € 407.680
Detector and beam monitor | € 0 | € 200.000 | € 2.390.000 | € 54.000 | € 108.000 | € 2.390.000 | € 362.000 | € 2.752.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion control |  | € 0 | € 116.000 | € 15.000 | € 50.000 | € 116.000 | € 65.000 | € 181.000
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 600.000 | € 21.600 | € 27.000 | € 660.000 | € 1.248.600 | € 1.908.600
Instrument Infrastructure | € 0 | € 27.000 | € 310.000 | € 172.800 | € 0 | € 310.000 | € 199.800 | € 509.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.355.550
Phase I |  |  |  |  |  |  |  | € 300.000
TOTAL | € 1.260.000 | € 563.000 | € 9.532.500 | € 718.840 | € 228.840 | € 9.562.500 | € 2.740.680 | € 13.978.730
Man year | 12 | 7 | 0 | 9 | 2 |  |  | 30
Chart: Labor costs - 2,75 M€ - 22,3 % (excl contingency)
| Project Management and integration | Design | Installation | Cold commissioning
 | 12 | 7 | 9 | 2

## Slide 41: Option 3
Option 3 includes a high field magnet to fully address the magnetism and SC communities
Option 3 includes polarization analysis. We use an S-bender before the sample and either a wide angle analyser or a set of Heusler analysers
after the sample. Bifrost would have the worlds most intense polarized beam and offer completely new possibilities.
The upgrade path of option 3 is to install a fast OS chopper before the sample.

## Slide 42: Option 3 – sacrifices
We do not implement order sorting – a unique option is postponed
We do not cover the full range of scattering angle – a loss of efficiency.
Y. Li, et. al,
Nature, 468
(2010)
Jan Zaanen
Nature
(2011)

## Slide 43: Detector system for option 3
Item | Price [k€]
Detectors | 450
PG | 400
Mounts | 120
Electronics | 400
Cross talk | 100
Detector tank | 500
Frame+rail+motor | 100
Detector safety | 50
Beam monitor/Camera/preamps | 120
S-bender | 200
Heussler analyzers | 1000
Flipper | 100
Guide fields | 200
Total | 3890
Labor for design, installation and commissioning: 510 k€

## Slide 44: Sample environment option 3
Possibility 1: A 14 T large bore magnet – 620 k€
(Dy-booster could get us to 16.5 T)
Possbility 2: Design a magnet. No heat shield, built-in radial collimation large bore – targeted specifications - 1000 k€

## Slide 45: Polarization option 3
Item | Price [k€]
S-bender polarizer | 350
Heussler analysers | 850
Flipper | 100
Guide fields | 200
Wide angle PA instead of Heusler | 1500
Total | 1500-2000
Wide angle analyser: Focus on PA for all analysers, more efficient
Heusler: Spin-flip and total cross sections simultaneously, more efficient on the specific analyser

## Slide 46: WBS breakdown
 | Project Management and integration | Design | Procurement and fabrication | Installation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 2.078.750 | € 259.200 | € 0 | € 2.078.750 | € 448.200 | € 2.526.950
Neutron Optics | € 0 | € 27.000 | € 3.040.000 | € 100.000 | € 0 | € 2.960.000 | € 207.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 100.000 | € 1.000.000 | € 3.840 | € 3.840 | € 1.000.000 | € 107.680 | € 1.107.680
Detector and beam monitor | € 0 | € 350.000 | € 3.740.000 | € 54.000 | € 108.000 | € 3.740.000 | € 512.000 | € 4.252.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion control | € 17.400 | € 60.000 | € 112.000 | € 6.000 | € 34.000 | € 77.000 | € 152.400 | € 229.400
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 600.000 | € 21.600 | € 27.000 | € 660.000 | € 1.248.600 | € 1.908.600
Instrument Infrastructure | € 0 | € 27.000 | € 310.000 | € 172.800 | € 0 | € 310.000 | € 199.800 | € 509.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.621.931
Phase I |  |  |  |  |  |  |  | € 250.000
TOTAL | € 1.277.400 | € 773.000 | € 11.771.750 | € 709.840 | € 212.840 | € 11.756.750 | € 2.988.080 | € 16.616.761
Man year | 12 | 8 | 0 | 9 | 2 |  |  | 31

## Slide 47: WBS breakdown
 | Project Management and integration | Design | Procurement and fabrication | Installation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 1.905.500 | € 259.200 | € 0 | € 1.905.500 | € 448.200 | € 2.353.700
Neutron Optics | € 0 | € 27.000 | € 3.020.000 | € 100.000 | € 0 | € 2.950.000 | € 197.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 100.000 | € 800.000 | € 3.840 | € 3.840 | € 800.000 | € 107.680 | € 907.680
Detector and beam monitor | € 0 | € 350.000 | € 3.890.000 | € 54.000 | € 108.000 | € 3.890.000 | € 512.000 | € 4.402.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion control |  | € 0 | € 116.000 | € 15.000 | € 50.000 | € 116.000 | € 65.000 | € 181.000
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 600.000 | € 21.600 | € 27.000 | € 660.000 | € 1.248.600 | € 1.908.600
Instrument Infrastructure | € 0 | € 27.000 | € 310.000 | € 172.800 | € 0 | € 310.000 | € 199.800 | € 509.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.592.050
Phase I |  |  |  |  |  |  |  | € 300.000
TOTAL | € 1.260.000 | € 713.000 | € 11.532.500 | € 718.840 | € 228.840 | € 11.562.500 | € 2.890.680 | € 16.365.230
Man year | 12 | 8 | 0 | 9 | 2 |  |  | 31
Chart: Option 3 budget - Primary scope - 16.4 M€
| Shielding | Neutron Optics | Choppers | Sample Environment | Detector and beam monitor | DMSC | Motion control | Instrument specific Tech equipment | Instrument Infrastructure | Vacuum | PSS | Contingency | Phase I
 | 2353700 | 3167000 | 923400 | 907680 | 4402000 | 0 | 229400 | 1908600 | 509800 | 0 | 120000 | 1597373.8 | 300000

## Slide 48: WBS breakdown
 | Project Management and integration | Design | Procurement and fabrication | Installation | Cold commissioning | Non labour | Labour | Total
Shielding | € 0 | € 189.000 | € 1.905.500 | € 259.200 | € 0 | € 1.905.500 | € 448.200 | € 2.353.700
Neutron Optics | € 0 | € 27.000 | € 3.020.000 | € 100.000 | € 0 | € 2.950.000 | € 197.000 | € 3.167.000
Choppers | € 0 | € 0 | € 841.000 | € 62.400 | € 20.000 | € 861.000 | € 62.400 | € 923.400
Sample Environment | € 0 | € 100.000 | € 800.000 | € 3.840 | € 3.840 | € 800.000 | € 107.680 | € 907.680
Detector and beam monitor | € 0 | € 350.000 | € 3.890.000 | € 54.000 | € 108.000 | € 3.890.000 | € 512.000 | € 4.402.000
DMSC | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
Motion control |  | € 0 | € 116.000 | € 15.000 | € 50.000 | € 116.000 | € 65.000 | € 181.000
Instrument specific Tech equipment | € 1.260.000 | € 0 | € 600.000 | € 21.600 | € 27.000 | € 660.000 | € 1.248.600 | € 1.908.600
Instrument Infrastructure | € 0 | € 27.000 | € 310.000 | € 172.800 | € 0 | € 310.000 | € 199.800 | € 509.800
Vacuum | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0 | € 0
PSS | € 0 | € 20.000 | € 50.000 | € 30.000 | € 20.000 | € 70.000 | € 50.000 | € 120.000
Contingency |  |  |  |  |  |  |  | € 1.592.050
Phase I |  |  |  |  |  |  |  | € 300.000
TOTAL | € 1.260.000 | € 713.000 | € 11.532.500 | € 718.840 | € 228.840 | € 11.562.500 | € 2.890.680 | € 16.365.230
Man year | 12 | 8 | 0 | 9 | 2 |  |  | 31
Chart: Labor, primary scope - 2.9 M€ - 20 %
| Project Management and integration | Design | Installation | Cold commissioning
 | 1277400 | 773000 | 709840 | 212840

## Slide 49: Option 2+
Sharing a magnet with another instrument – compromise specifications, but focus on high field.

## Slide 50: Backup slides

## Slide 51: Thank you for your attention