## BUDGET
Timestamp | Untitled Question
k€
OPTION 3 (17.7 M€)
|  | Phase 1 | Design | Procurement & Fabrication | Installation | 06 Cold Commissioning | LABORAL BUDGET | NON LABORAL BUDGET | Total k€
| Primary Shielding | 0 | 200 | 809 | 100 | 0 | 300 | 809 | 1109
| Secondary shielding (cave + vacuum tank shielding) | 0 | 200 | 860 | 20 | 0 | 220 | 860 | 1080
| Neutron Optics without vacuum pumps | 0 | 0 | 3558 | 100 | 0 | 100 | 3558 | 3658
| Choppers | 0 | 0 | 1100.75 | 65.56 | 19.14 | 65.56 | 1119.89 | 1185.45
| Sample Environment | 0 | 0 | 0 | 300 | 808 | 300 | 808 | 1108
| Detector and Beam Monitors | 0 | 0 | 1024.95 | 76.65 | 0 | 76.65 | 1024.95 | 1101.6
| Motion Control and Automation | 0 | 62.1 | 39.494 | 5.995 | 33.984 | 68.095 | 73.478 | 141.573
Specific equipment | Analyser | 0 | 300 | 890 | 100 | 0 | 400 | 890 | 1290
| Radial collimator | 0 | 50 | 105 | 20 | 0 | 70 | 105 | 175
| Berylium filter | 0 | 0 | 450 | 20 | 0 | 20 | 450 | 470
| Beam Stop | 0 | 50 | 100 | 20 | 0 | 70 | 100 | 170
| Slit | 0 | 0 | 60 | 10 | 0 | 10 | 60 | 70
| Vacuum tank | 0 | 300 | 940 | 100 | 0 | 400 | 940 | 1340
| Instrument Infrastructure | 0 | 0 | 637 | 0 | 0 | 0 | 637 | 637
| Management & Science | 480 | 600 | 480 | 480 | 480 | 2520 |  | 2520
| Contingency | 0 | 116.21 | 1057.4194 | 93.8205 | 86.1124 | 210.0305 | 1143.5318 | 1605.5623
| Total | 480 | 1878.31 | 12111.6134 | 1512.0255 | 1427.2364 | 3870.3355 | 12578.8498 | 17661.1853
-----o-----
| Labour included in above (Person-Years) | 4.1 | 15.65 | 7.6 | 6.1 | 2.7 | 36.15 |  | 36,15
OPTION 2 (15.9 M€)
|  | Phase 1 | Design | Procurement & Fabrication | Installation | 06 Cold Commissioning | LABORAL BUDGET | NON LABORAL BUDGET | Total k€
| Primary Shielding | 0 | 200 | 809 | 100 | 0 | 300 | 809 | 1109
| Secondary shielding (cave + vacuum tank shielding) | 0 | 200 | 860 | 20 | 0 | 220 | 860 | 1080
| Neutron Optics without vacuum pumps | 0 | 0 | 3558 | 100 | 0 | 100 | 3558 | 3658
| Choppers | 0 | 0 | 1100.75 | 65.56 | 19.14 | 65.56 | 1119.89 | 1185.45
| Sample Environment | 0 | 0 | 0 | 180 | 431 | 180 | 431 | 611
| Detector and Beam Monitors | 0 | 0 | 758 | 76.65 | 0 | 76.65 | 758 | 834.65
| Motion Control and Automation | 0 | 53.712 | 29.373 | 4.65 | 28.578 | 58.362 | 57.951 | 116.313
Specific equipment | Analyser | 0 | 215 | 712.16 | 70 | 0 | 285 | 712.16 | 997.16
| Radial collimator | 0 | 50 | 105 | 20 | 0 | 70 | 105 | 175
| Berylium filter | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Beam Stop | 0 | 50 | 100 | 20 | 0 | 70 | 100 | 170
| Slit | 0 | 0 | 60 | 10 | 0 | 10 | 60 | 70
| Vacuum tank | 0 | 300 | 940 | 100 | 0 | 400 | 940 | 1340
| Instrument Infrastructure | 0 | 0 | 637 | 0 | 0 | 0 | 637 | 637
| Management & Science | 480 | 600 | 480 | 480 | 480 | 2520 |  | 2520
| Contingency | 0 | 106.8712 | 966.9283 | 76.686 | 47.8718 | 183.5572 | 1014.8001 | 1450.3573
| Total | 480 | 1775.5832 | 11116.2113 | 1323.546 | 1006.5898 | 3579.1292 | 11162.8011 | 15953.9303
-----o-----
| Labour included in above (Person-Years) | 4.1 | 14.65 | 7.6 | 6.1 | 2.7 | 35.15 |  | 35.15
OPTION 1 (12.9 M€)
|  | Phase 1 | Design | Procurement & Fabrication | Installation | 06 Cold Commissioning | LABORAL BUDGET | NON LABORAL BUDGET | Total k€
| Primary Shielding | 0 | 200 | 809 | 100 | 0 | 300 | 809 | 1109
| Secondary shielding (cave + vacuum tank shielding) | 0 | 200 | 600 | 20 | 0 | 220 | 600 | 820
| Neutron Optics without vacuum pumps | 0 | 0 | 3558 | 100 | 0 | 100 | 3558 | 3658
| Choppers | 0 | 0 | 1100.75 | 65.56 | 19.14 | 65.56 | 1119.89 | 1185.45
| Sample Environment | 0 | 0 | 0 | 0 | 76.5 | 0 | 76.5 | 76.5
| Detector and Beam Monitors | 0 | 0 | 310.69 | 76.65 | 0 | 76.65 | 310.69 | 387.34
| Motion Control and Automation | 0 | 48.757 | 27.424 | 3.413 | 23.599 | 52.17 | 51.023 | 103.193
Specific equipment | Analyser | 0 | 60 | 178.04 | 20 | 0 | 80 | 178.04 | 258.04
| Radial collimator | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Berylium filter | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Beam Stop | 0 | 50 | 100 | 20 | 0 | 70 | 100 | 170
| Slit | 0 | 0 | 60 | 10 | 0 | 10 | 60 | 70
| Vacuum tank | 0 | 200 | 600 | 70 | 0 | 270 | 600 | 870
| Instrument Infrastructure | 0 | 0 | 637 | 0 | 0 | 0 | 637 | 637
| Management & Science | 480 | 600 | 480 | 480 | 360 | 2400 |  | 2400
| Contingency | 0 | 75.8757 | 798.0904 | 48.5623 | 11.9239 | 124.438 | 810.0143 | 1174.4523
| Total | 480 | 1434.6327 | 9258.9944 | 1014.1853 | 491.1629 | 2928.818 | 8910.1573 | 12918.9753
-----o-----
| Labour included in above (Person-Years) | 4.1 | 13.65 | 7.6 | 6.1 | 2.7 | 34.15 |  | 34.15

## Primary Guides
| Guide Section ID | Length [m] | Dimensions [cm] | m | Substrate | Curvature [m] | €/m | Total price € | Total price k€ | Reference
CHF to € | Section 1 | 5 | from 6 x 7,5 to 4 x 12 | 3 | Superpolished copper |  |  | 144060.822 | 144.0608 | Swiss Neutronics: 146929 CHF
0.918 | Section 2 | 0.5 | from 4 x 12 to 4,3 x 12 | 3 | aluminium |  |  | 13939.83 | 13.9398 | Swiss Neutronics: 15185 CHF
| Section 3 | 9 | from 4,3 x 12 to 11 x 12 | 3 | aluminium |  |  | 283809.798 | 283.8098 | Swiss Neutronics: 309161 CHF
| Section 4 | 6 | from 11 x 12 to 11 x 12 | 1,5 | aluminium |  |  | 130392.72 | 130.3927 | Swiss Neutronics: 142040 CHF
| Section 5 | 32.5 | from 11 x 12 to 11 x 12 | 2 | borkron glass (N-BK7 or S-BSL7) | 5000 |  | 542186.406 | 542.1864 | Swiss Neutronics: 590617 CHF
| Section 6 | 1.15 | from 11 x 12 to 11 x 12 | 1,5 | borkron glass (N-BK7 or S-BSL7) |  |  | 14681.574 | 14.6816 | Swiss Neutronics: 15993 CHF
| Section 7 | 22.5 | from 11 x 12 to 11 x 12 | 2 | borkron glass (N-BK7 or S-BSL7) | 5000 |  | 375359.184 | 375.3592 | Swiss Neutronics: 408888 CHF
| Section 8 | 75 | from 11 x 12 to 11 x 12 | 1,5 | borkron glass (N-BK7 or S-BSL7) |  |  | 957484.098 | 957.4841 | Swiss Neutronics: 1043011 CHF
| Section 9 | 7 | from 11 x 12 to 3.2 x 3.2 | 4 | borkron glass (N-BK7 or S-BSL7) |  |  | 214239.168 | 214.2392 | Swiss Neutronics: 233376 CHF
| vacuum system | 160 |  |  |  |  | 4400 | 704000 | 704 | SD-H email: The price for a 2 m long vacuum system made from Aluminium with 3 mm integrated B4C shielding is about 4.4 k€/m for cross section 12x12 cm². Included are the rubbers sleeves and clamps to connect the 2 m long vacuum housings.
| steel support structure | 160 |  |  |  |  | 600 | 96000 | 96 | SD-H email: A steel support structure will costs about 1.2 k€ every 2 meter or 0.6 k€/m.
| installation, alignment and documentation | 160 |  |  |  |  | 1200 | 192000 | 192 | SD-H email: The price for the installation, alignment and documentation is about 1.2 k€/m (for both guide and support)
[8 empty] | 3668153.6 | 3668.1536
SHIELDING
| Guide shielding materials | m³ | €/m³ | €
| Steel | 50.55 | 12000 | 606600
| Concrete | 254 | 800 | 203200
| Total |  |  | 809800
| Reference | "NOSG Cost Calculation Process" + "MIRACLES. Preliminary shielding budget estimation for the scope setting meeting"

## Primary slit
| k€ | k€
Pre-sample slits | 60 | 40 | Source: Mirrotron
| With motors | without motors (20-40% less)
motorized slit 4 motors, absolut rotary encoder, ball screw drive, vacuum feedthroughs, safety limit switchers, CF/ISO flange vacuum window
· operation under vaccuum of 10-3 mbar
· absorbing material 10B4C in resin on Cd plate
· control cabinet control software for Windows system
· interface for ESS requirement (strongly depending on specification!)
· the slit size and the beam center axis position can be adjusted over the whole cross-section of 50 x 50 mm

## Detectors
CONFIGURATION 3
| Cost/unit (k eur) | Units | Total Cost (k eur) | Laboral cost per component (keur)
1. Spectrometer detector |  |  | 326.76 | 35.529
3He tubes | 1.51 | 150 | 226.5
Power supplies + cables | 50.26 | 1 | 50.26
Mounting frames | 50 | 1 | 50
2. Diffractometer backscattering detector |  |  | 147.96 | 16.0879
3He tubes | 3 | 30 | 90
Power supplies + cables | 7.96 | 1 | 7.96
Mounting frames | 50 | 1 | 50
3. Detector electronics |  |  | 160
Back end electronics | 160 | 1 | 160
4. Beam Monitors |  |  | 230.23 | 25.0332
Position sensitive | 30 | 7 | 210
Power supplies + cables | 2.89 | 7 | 20.23
5. Labour Cost |  |  | 76.65
Specs & Procurement (1person/2 months) | 0.06 | 350 | 21
Tests in ESS Bilbao (1 person/2 months) | 0.06 | 175 | 10.5
Installation (1 person / 1 months) | 0.078 | 175 | 13.65
Comissioning (1 person /2 months) | 0.09 | 350 | 31.5
TOTALS non laboral |  |  | 864.95
TOTALS detectors & beam monitors |  |  | 704.95
CONFIGURATION 2
| Cost/unit (k eur) | Qtty | Total Cost (k eur) | Laboral cost per component (keur)
1. Spectrometer detector |  |  | 275.2 | 40.4233
3He tubes & Power supplies | 1.54 | 120 | 184.8
Power Supplies + Cables | 40.4 | 1 | 40.4
Mounting frames | 50 | 1 | 50
2. Diffractometer backscattering detector |  |  | 147.96 | 21.7334
3He tubes & Power supplies | 3 | 30 | 90
Power Supplies + Cables | 7.96 | 1 | 7.96
Mounting frames | 50 | 1 | 50
3. Detector electronics |  |  | 160
Back end electronics | 160 | 1 | 160
4. Beam Monitors |  |  | 98.67 | 14.4933
Position sensitive | 30 | 3 | 90
Power Supplies + Cables | 2.89 | 3 | 8.67
5. Labour Cost |  |  | 76.65
Specs & Procurement (1person/2 months) | 0.06 | 350 | 21
Test in ESS Bilbao (1person/1month) | 0.06 | 175 | 10.5
Installation (1 person / 1 months) | 0.078 | 175 | 13.65
Comissioning (1 person /2 months) | 0.09 | 350 | 31.5
TOTALS non laboral |  |  | 681.83
TOTALS detectors & beam monitors |  |  | 521.83
CONFIGURATION 1
| Cost/unit (k eur) | Qtty | Total Cost (k eur) | Electronics cost (keur)
1. Spectrometer detector |  |  | 117.8 | 59.9202
3He tubes | 1.8 | 30 | 54
Power supplies + cables | 13.8 | 1 | 13.8
Mounting frames | 50 | 1 | 50
2. Diffractometer backscattering detector |  |  | 0 | 0
3He tubes | 3 | 0 | 0
Power supplies + cables | 7.96 | 0 | 0
Mounting frames | 50 | 0 | 0
3. Detector electronics |  |  | 160
Back end electronics | 160 | 1 | 160
4. Beam Monitors |  |  | 32.89 | 16.7298
Position sensitive | 30 | 1 | 30
Power supplies + cables | 2.89 | 1 | 2.89
5. Labour Cost |  |  | 76.65
Specs & Procurement (1person/2 months) | 0.06 | 350 | 21
Test in ESS Bilbao (1person/1month) | 0.06 | 175 | 10.5
Installation (1 person / 1 months) | 0.078 | 175 | 13.65
Comissioning (1 person /2 months) | 0.09 | 350 | 31.5
TOTALS non laboral |  |  | 310.69
TOTALS detectors & beam monitors |  |  | 150.69

## Analyzer
Si-111 ANALYZER | 150º horizontal coverage | 120º horizontal coverage
| Cost [k€] | Cost [k€] | Sources &Comments:
Si-111 wafers: | 340.2 | 272.16 | BASIS
B4C Coating: | 0 | 0 | IN16B
Cast Aluminum Support Structure: | 0 | 0 | IN16B graphite composite (120º horizontal, 38º vertical analyser)
IN16B: HDPE + B10 sheet + support structure (Area of IN16B)
TOTAL COST [k€] = | 340.2 | 272.16
Shielding (HDPE + B10): | 660
| Bending radius crystal analyzer (p) | p = L / cos (90 - THbragg)
Si perfect crystals of hexagonal shape that will be bent and glued to a spherical backing plate | sample - analyzer distance, R [m] = | 2.5
Si-111 hexagonal crystal wafers, dimensions: 2mm thick, 12cm diagonal. | THBragg [deg] = | 88 | 1.5359
The surface area of the analyzer increases by 50% with a radial increase from 2.5 to 3 m |  | 90 | 1.5708
| p [m] = | 2.5015
SILICON CRYSTALS | [5 empty] | Scattering Characterization System (SCS)
Source: BASIS | [5 empty] | - Beam monitors
| Conversion $-eur |  |  |  |  | - Crystal analyzer
| $ | € |  |  |  | - Shielding / background
| 1$ | 0.89645 |  |  |  | - Beamstop
Si Crystals Price | 150 $/wafer | 135
| 200 $/wafer | 180
Silicon density (kg/m3) | 2330
Volume per wafer (m3) | 1.9e-05
Weight per wafer (kg) | 0.04427
44.27 g
Single Analyzer PANEL (16.66666667º horizontal coverage, vertical coverage 22º)
horizontal coverage per panel | wafers per panel | Total cost (k€) | - One shell (panel) covers a solid angle of about 16.7º x 22º approx. 0.112 rad²
16.6667 | 84 | 15.12
weight of the crystals in a panel (kg) | 3.7187 | horizontal coverage 1 panel [deg] | 16.6667 | 0.29089 | rad
0.29089 | rad | vertical coverage 1 panel [deg] | 22 | 0.38397
0.72722 | m | solid angle = area / radius² |  | 6980.8159
6.545 | m | Area of 1 panel of the analyzer [m²] = | 0.69808 | m²
For 150º | 9 | panels per set
For 120º | 7.2 | panels per set
SCENARIO 1: 150º horizontal coverage | 150
MIRACLES analyzer 150º horizontal coverage = Full coverage (Span is 150º = coverage of nearly 21% of 4*Pi steradians) |  | Total Cost (k€) Si analyzer 12.6 m²
Sets of panels (1 upper and 1 lower) | 2 |  | Analyser coverage | Horizontal coverage | Vertical coverage
Panels per set (for 150º coverage) | 9 |  | [deg] | 150 | 44
Total number of PANELS | 18 |  | [rad] | 2.618 | 0.76794
Total number WAFERS | 1512 | 272.16 | [m] | 6.545 | 1.9199
Set aside +25% for breaking wafers, Total number of wafers | 1890 | 340.2 | solid angle [sr] | 2.0105
|  |  | Area of the analyzer for 150º [m²] = | 12.5655
SCENARIO 2: 120º horizontal coverage | 120
MIRACLES analyzer 120º horizontal coverage (Span is 116º = coverage of nearly % of 4*Pi steradians) |  | Total Cost (k€) Si analyzer 9.7 m²
Sets of panels (1 upper and 1 lower) | 2 |  | Analyser coverage | Horizontal coverage | Vertical Coverage
Panels per set (for 120º coverage) | 7.2 |  | [deg] | 120 | 44
Total number of PANELS | 14.4 |  | [rad] | 2.0944 | 0.76794
Total number WAFERS | 1209.6 | 217.728 | [m] | 5.236 | 1.9199
Set aside +25% for breaking wafers, Total number of wafers | 1512 | 272.16
|  |  | Area of the analyzer for 67º [m²] = | 10.0524
COMPLETE ANALYZER
List of components
- Hexagonal crystal wafers of Si, 111 orientation
B4C Coating backside of the Silicon wafers
Aluminum support structure
Shielding of the analyser (HDPE + B10 + support structure)
Manpower for gluing and alignment
A B4C sheet is mounted between the Silicon wafers and the aluminum support frame
Note on A11: Reference: Study of the analyzer crystals for use in the near-backscattering spectrometer DNA at J-PARC
Si-111 wafers
B4C,
neutron absorber
Aluminum plate
Analyzer layout

## Primary Choppers
| Chopper type | Chopper number | Distance from moderator [m] | Grounding Floor |  | MIRACLES Chopper Cascade | TOTAL Cost (k€)
| PWD | Ch1 | 7 | Bunker (R 28m) |  |  |  |  | cost keur
|  | Ch2 | 7.01 |  |  |  | 1185.45 | Chopper discs | 682.55
| PS | Ch3 | 7.5 |  |  |  |  | CHIM | 199
|  | Ch4 | 7.51 |  |  |  |  | Vacuum | 8
| WBD | Ch5 | 22.5 |  |  |  |  | Chopper cooling | 10
|  | Ch6 | 56.25 | E02/Guide Hall (R 154.5m) |  |  |  | Chopper drive | 113.4
|  | Ch7 | 80.5 |  |  |  |  | Master rack | 48.6
[8 empty] | Slave rack | 6.9
[8 empty] | Condition monitoring | 18
[8 empty] | External Power | 14.3
[8 empty] | Integration | 29.92
[8 empty] | Installation | 35.64
[8 empty] | Commissioning | 19.14
[8 empty] | TOTAL | 1185.45
FINAL VERSION (3 Master Rack)
| No |  | no. units | Price each | Cost (keur)
1 Inside Bunker | 1 | Single Disc Chopper Assembly, Small rotor, Slow speed (x3) | 3.1 | 71.5 | 221.65 | One chopper is inside the bunker, = +10%
Inside Bunker+Safety Class | 3 | Over/Under disc chopper assembly, Small rotor, Slow speed (PS) | 1.1 | 139 | 152.9 | Inside the bunker = +10%
Inside Bunker+Safety Class | 9 | Over/Under disc chopper assembly, Small rotor, High speed (PWD) | 1.1 | 280 | 308 | Inside the bunker = +10%
| 14 | CHIM, Horizontal Split, small rotor (x3) | 3 | 33 | 99 | Reduce cost by 20% per unit if five or more units are used.
| 15 | CHIM, Over/Under, small rotor (x3) | 2 | 50 | 100 | Reduce cost by 20% per unit if five or more units are used.
| 21 | Control Rack Master (x3) | 3 | 16.2 | 48.6 | 3 units. At least one per grounding zone, but we only have 2 zones (Bunker + E02/Guide Hall). |  |  | One master rack can host up to four drives and one slave rack.
| 22 | Control Rack Slave (x1) | 1 | 6.9 | 6.9 | One slave control rack can host up to eight drives.
| 23 | Chopper Drive (x7) | 6.3 | 18 | 113.4 | One drive per axis. Reduce cost by 10% per unit if five or more units are used.
| 24 | Condition Monitoring (x2) | 2 | 9 | 18 | For choppers primarily in the bunker. One per master rack inside the bunker.
| 25 | External Power (x7) | 2.2 | 6.5 | 14.3 | One per axis. UPS and cables. +20% per additional chopper axis after the first.
| 26 | Vacuum System (x5) | 1.6 | 5 | 8 | Including flanges and tubes. +15% per additional achopper assembly after the first.
| 27 | Cooling System (x1, for the high speed chopper assembly) | 1 | 10 | 10 | 1 unit (for the fast PWD chopper assembly). |  |  | Maximum one chopper cooling system per assembly. Only the intermediate and fast speed choppers need cooling. Add 15% per additional chopper assembly after the first.
| 28 | Installation tooling | 1 | 0 | 0 | Cost is zero if CHIM installation variants are used. Otherwise cost for lifting and extraction tooling is added.
| 29 | Maintenance tooling | 1 | 0 | 0 | Cost is zero if CHIM installation variants are used. Otherwise cost for lifting and extraction tooling is added.
| 30 | Insfraestructure | 1 | 0 | 0 | Cost is zero if CHIM installation variants are used. Otherwise cost for lifting and extraction tooling is added.
| 31 | Integration of the complete chopper system (x5) | 2.2 | 13.6 | 29.92 | One per assembly. +30% per additional chopper assembly.
| 32 | Installation of the complete chopper system (x5) | 2.2 | 16.2 | 35.64 | One per assembly. +30% per additional chopper assembly.
| 33 | Commissioning of the complete chopper system (x5) | 2.2 | 8.7 | 19.14 | One per assembly. +30% per additional chopper assembly.
| TOTAL COST : |  |  |  | 1185.45
|  |  |  | Laboral budget | 84.7
|  |  |  | Non laboral | 1100.75
Note on D26: 2 Master rack for choppers placed inside the bunker (1 condition monitoring each), the remainder Master rack is placed outside the bunker.
Note on C32: Only intermediate and fast choppers need cooling

## Sample environment
|  |  | Device | Laboral
|  |  | k€ | k€ | Source
SAD (Harald Schneider) | Total Day 1 (option 2) | 611.5 | 431.5 | 180
| Total Day 1 + Upgrade (option 3) | 1108 | 808 | 300
Quotations and extra info | Total Day 1 (option 2) | 451.6 | 356.6 | 95
| Total Day 1 + Upgrade (option 3) | 701.6 | 506.6 | 195
SAD (Harald Schneider)
Option 2 (Day 1) | Orange Cryofurnace (SAD-TEFI; 2-600K; plug&play) |  | 50
| Top loading CCR, with customer High Temperature option |  | 80
| High temperature furnace |  | 75
| Humidity chamber |  | 100
| Gas process handling, special dedicated cells |  | 200
| High pressure cell |  | 50
| Electromagnet |  | 50
| Laser pump probe setup |  | 200
| Orange Cryofurnace (SAD-TEFI; 2-600K; plug&play; with sample changer 10-20 samples) |  | 250
| 6 kV HV supply (for e.g. electric-field control cells) |  | 25 |  | Option MINIPOOL with SKADI??
| Commissioning & first real tests |  |  | 180
| ESS support |  | 76.5 |  | ESS-0063538 - ESS contribution for instrument construction and infrastructure (Harald Schneider confirmed)
Option 3 (or Option 2 upgrade) | Dilution refrigerator insert / He3 insert |  | 250
| Clamp cells |  | 50
| In-situ thermal analysis measurements |  | 250
| In-situ gas adsorption measurements |  | 250
| Commissioning & first tests & Continuous user operation |  |  | 300
Quotations and extra info
Option 2 (Day 1) | Orange Cryofurnace (ILL; 50 mm; 2-600K) |  | 32.4 |  | AS-Scientific and ECNS_2011/Poster_Dry_Cryostat.pdf
| Installation & accesories O-CF (dewar, electronics,...) |  | 35 |  | AS-Scientific and ECNS_2011/Poster_Dry_Cryostat.pdf
Top-loading type closed-cycle refrigerator CCR (cryogen-free): 2-750 K
| Standard Furnace -(Vanadium): < 1600ºC (~1900 K) |  | 38.2 |  | AS-Scientific | 0
| Accesories Furnace (Vacuum, electronics, spares....) |  | 40.5 |  | AS-Scientific
| Gas Pressure cell / Chemical reactor (design) |  |  | 30 | Heloisa | 0
| Pressure rigs for gas and liquid handling (in cell loading) |  |  | 30 | Heloisa | 0
| Gas processing / manifold and accesories |  |  |  |  | 0
| Humidity chamber | ESS SE Pool |  | 25 | Heloisa
| Pump-probe (Light-excited) experiments | ESS SE Pool
In-situ gas adsorption measurements
| Sample holders, different geometries: | To share it with the pool | 10 | 10 | Flat & Annular aluminium alloy cans for powders, liquids or films, with varying thicknesses (0.1, 0.2, 0.25, 0.3, 0.4, 0.5, 1, 2 mm)
| Dry roughing and turbomolecular pumps | 3.5 k€ rough (x2); 5 k€ turbo 50 l; 12 k€ turbo 500 l | 24 |  | Agilent (TeraTorr) and Varian
| Sample table (positioning stage) |  | 100 |  | BIFROST
| ESS support |  | 76.5 |  | ESS-0063538 - ESS contribution for instrument construction and infrastructure (Harald Schneider confirmed) | 0
Option 3 (or Option 2 upgrade) | Dilution stick |  | 150 |  | BIFROST | 0
| Dilution stick (design) |  |  | 100 | BIFROST (1 man year) | 0
| Cryomagnet 9T | ESS SE Pool |  |  |  | 0
Clamp-pressure cell: < 2-3 GPa (20-30 kbar)
| Corrosive gas processing |  |  |  |  | Source
| In-situ electric field control for membrane protein studies |  |  | 0 |  | Xtreme D
In-situ thermal analysis measurements
| In-situ Raman system |  |  |  | In-situ Raman + Catalysis cell: 250 k€ (Heloisa)
Automatic sample changer for catalysis studies
Facility to enable optical access to sample via fibre optics
| Total Day 1 |  | 356.6 | 95
| Total Day 1 + Phase 2 |  | 506.6 | 195
| Pool equipment (provided by ESS SAD) 0.00 |  | k€
Note on C22: Jones: 700 / m2 plus 500 transport per tank
Note on C44: Jones: 700 / m2 plus 500 transport per tank
Note on C45: Jones: 700 / m2 plus 500 transport per tank

## Secondary
Vacuum Tank |  | CAMEA: Aluminum vaccum tank, 5cm thick to sustain the outside pressure. 6 tons total weight of the tank (para 130º horizontal coverage y 2.8º vertical)
BIFROST: Vacuum tank 120º coverage and 2.8m = 1.5 Meur
Instrument cave |  | CAMEA REPORT: Instrument Cave and beam stop = 1M€
BIFROST: CAVE 750 keur (1m thick concrete 450 keur, 3mm boralcan plates = 200keur, manpower = 100 keur.)
Shielding of the tank |  | CAMEA: The tank is covered with a 30 cm layer of B4C on the outside and a thin layer of Cadmium on the inside (to reduce penetration of fast, epithermal, and thermal neutrons)
Holders/supports and cadmium shielding for detectors
Mechanical holders for analyzer
Beam stop
Mechanics needed to rotate all analyzers around sample position
Radial Collimator |  | 105 k€ (JJXRAY)
Beryllium Filter | The beryllium filter supresses higher order reflections to provide access to high energy transfers
| Instruments | TOTAL Cost k€: | MIRACLES
| OSIRIS [k€] | 390.696 | 451.8169
MIRACLES
[7 empty] | Be-filter cost (kpounds) | Cost (k eur)
Be-Filter MIRACLES: | Notes: | A retractable filter is required if some other reflections (apart from Si-111) are planned to be used. |  |  |  |  | 390.696 | 451.8169
Beryllium Cost: | For 150º of horizontal coverage:
| Cost per block [€] | 382.75
| Number of Be blocks (0.5º each) | 300
| Beryllium cost [k€] | 114.825
Housing and shielding of the filter: [k€] |  | 155
Electronics, Cooling & Installation: [k€] |  | 181.9919
TOTAL Cost [k€] |  | 451.8169 |  | Conversion pound-euro:
|  |  |  | 1 | 1.1627 | eur
Note on D31: Fuente: OSIRIS
Note on D32: Sin cooling!