## Slide 1
MIRACLES SCOPE (2/2)
Secondary Spectrometer,
Motion Control, Infrastructure

## Slide 2
Index
Secondary Spectrometer
  Analyzer
  Radial Collimator
  Beryllium filter
  Secondary cave, vacuum tank and beam stop
  Detectors and beam monitors
Motion Control
Infrastructure

## Slide 3
Analyzer
Si-111 hexagonal crystal wafers: dimensions: 2mm thick, 12 cm diagonal.
Single analyzer panel (aluminum): 15 º horizontal coverage, 22º vertical coverage.
B4C coating backside of the Silicon wafers.
Si-111 wafers
B4C,
neutron absorber
Aluminum plate
Analyzer layout
Components
Configurations
3: 150 º horizontal coverage and 44 º vertical coverage.
  20 panels, 1890 wafer. Analyzer area = 12.60 m².
2: 120 º horizontal coverage and 44 º vertical coverage.
  16 panels, 1512 wafer. Analyzer area = 10.08 m².
1: 30 º horizontal coverage and 44 º vertical coverange.
  4 panels, 378 wafer. Analyzer area = 2.52 m².

## Slide 4
Analyzer
Components

## Slide 5
Analyzer
Waffers: 200 $/waffer. Source: BASIS.
  340 k€ configuration 3.
  272 k€ configuration 2.
  68 k€ configuration 1.
QUOTATION AND SOURCES OF INFORMATION
B4C coating. 7.93 k€/m². Source: IN16B.
  100 k€ configuration 3.
  80 k€ configuration 2.
  20 k€ configuration 1.
Aluminum support structure 22.5 k€/panel. Source: IN16B.
  450 k€ configuration 3.
  360 k€ configuration 2.
  90 k€ configuration 1.

## Slide 6
Radial collimator
Detailed design of our radial collimator not done yet.
Preliminary concept:
  Foils: Mylar 0.036mm.
  Absorber: Gd2O3.
  Coated foil thickness:0.085mm/0.090mm.
  Foil Height: 600mm.
  Inside Radius: 348mm.
  Outside Radius: 780mm.
  Angular Spacing: 2°.
  Bottom Aluminium Support Plate: 40mm.
  Top Aluminium Support Plate: 40mm.
  Radial Beam Scan:160°.
Quotation: 105k€. Reference: Source: JJ-Xray. Responsible of the radial collimator of ARCS (Spectrometer at the SNS).

## Slide 7
Beryllium filter
Detailed design of our Be filter not done yet. Our quotation is 450 k€:
  300 Be blocks. 383 €/block: Be cost: 114.8K€.
  Housing and shielding of the filter: 155 k€.
  Electronics, cooling and installation: 182 k€.
We can find an example of BSS using Be filter in OSIRIS (ISIS).
During a conversation with the OSIRIS team, the suggested us a full economic cost of £400k, including overhead, electronics, installation…

## Slide 8
Secondary cave, vacuum tank & shielding, beam stop
Vaccum tank
  Estimation based on IN16B vacuum tank, made by Cadinox.
  Extrapolation with volume difference.
  Configurations 2 and 3 include vacuum tank with full covering (940 k€).
  Configuration 1 includes vacuum tank with half covering (600 k€).

## Slide 9
Secondary cave, vacuum tank & shielding, beam stop
Vacuum tank shielding:
  Estimation based on IN16B, considering the area difference.
  HDPE + ¹⁰B rubber + structure.
  660 k€ in configurations 3 and 2. 400 k€ in configuration 1.
Secondary cave:
  Concrete volume: 10x10 m inner surface, 6 m height, 0.5 m thickness.
  This means an estimation of ca 186 m³ of concrete.
  Assuming 800 €/m³ this would mean ca 150 k€.
  In order to include labyrinth and door, we assume 200 k€.
Beam stop:
  Design not ready yet.
  Assumed 150 k€, based on BIFROST estimation.
The sum of these 3 items gives us 1.08 M€, whereas P. Benley suggested us 1.5 M€ for the secondary shielding. Necessary to clarify if something is missing.

## Slide 10
Detectors and beam monitors
Spectrometer detectors:
  ³He tubes. Source: Toshiba.
    326.7 k€ in configuration 3 (150 tubes).
    275.2 k€ in configuration 2 (120 tubes).
    117.8 k€ in configuration 1 (30 tubes).
  Power supplies and cables. Source: CAEN quotation.
    50.26 k€ in configuration 3.
    40.4 k€ in configuration 2.
    13.8 k€ in configuration 1.
Difractometer detectors.
  90 k€ for 30 ³He tubes. Source: Toshiba.
  7.96 k€ for power supplies and cables. Source: CAEN quotation + cables.
Detector electronics. Source: ESS.
  160 k€ for back end electronics. Assumed a fixed cost for the 3 configurations.
  Front end electronics not considered in our budget.
  Necessary to discuss what is missing regarding detector electronics (the quotation for BASIS spectrometer detectors is 273.2k€).

## Slide 11
Detectors and beam monitors
PSD Beam monitors. 7 in configuration 3. Quotation: 30 k€ each. Source: ESS
  After the PWD chopper pair.
  After the PS chopper pair.
  After Chopper 5.
  After Chopper 6.
  After Chopper 7.
  Between the guide exit and the sample.
  After the sample.

## Slide 12
Motion Control
Axis Number | Device Description | WBS Element | Custom / Off The Shelf | Axis description | Motion Type | Actuator Type | Range (mm/°) | Accuracy (mm/°) | Precision / Repeatability (mm/°) | Encoder Type | Synch' with other axis | Synch' with signals / sensors | Load (Kg) | Force / Torque | Speed | Acceleration | Brake Required | Fixed or Removable
1 | Slit | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 50 | 0,1 | 0,1 | Absolute |  |  | 5 |  | Low | Low |  | Fixed
2 | Slit | 13.6.16.X.X.X- |  | Z2 | Linear | Electrical | 50 | 0,1 | 0,1 | Absolute |  |  | 5 |  | Low | Low |  | Fixed
3 | Slit | 13.6.16.X.X.X- |  | Y1 | Linear | Electrical | 50 | 0,1 | 0,1 | Absolute |  |  | 5 |  | Low | Low |  | Fixed
4 | Slit | 13.6.16.X.X.X- |  | Y2 | Linear | Electrical | 50 | 0,1 | 0,1 | Absolute |  |  | 5 |  | Low | Low |  | Fixed
5 | Sample Positioner | 13.6.16.X.X.X- |  | X1 | Linear | Electrical | 50 | 0,1 | 0,1 | Absolute |  |  | 50 |  | Medium | Low |  | Removable
6 | Sample Positioner | 13.6.16.X.X.X- |  | Y1 | Linear | Electrical | 50 | 0,1 | 0,1 | Absolute |  |  | 50 |  | Medium | Low |  | Removable
7 | Sample Positioner | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 50 | 0,1 | 0,1 | Absolute |  |  | 50 |  | Medium | Low |  | Removable
8 | Sample Positioner | 13.6.16.X.X.X- |  | X1 | Rotary | Electrical | 90 | 0,1 | 0,1 | Absolute |  |  | 50 |  | Medium | Low |  | Removable
9 | Sample Positioner | 13.6.16.X.X.X- |  | Y1 | Rotary | Electrical | 90 | 0,1 | 0,1 | Absolute |  |  | 50 |  | Medium | Low |  | Removable
10 | Sample Positioner | 13.6.16.X.X.X- |  | Z1 | Rotary | Electrical | 360 | 0,1 | 0,1 | Absolute |  |  | 50 |  | Medium | Low |  | Removable
11 | Upper detector array | 13.6.16.X.X.X- |  | X1 | Rotary | Electrical | 30 | 0,001 | 0,001 | Absolute |  |  | 100 |  | Low | Low |  | Fixed
12 | Lower detector array | 13.6.16.X.X.X- |  | X1 | Rotary | Electrical | 30 | 0,001 | 0,001 | Absolute |  |  | 100 |  | Low | Low |  | Fixed
13 | Be filter | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 3000 | 1 | 1 | Absolute |  |  | 100 |  | Low | Low |  | Removable
14 | Radial Collimator | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 3000 | 1 | 1 | Absolute |  |  | 300 |  | Low | Low |  | Removable
15 | Beam monitor 1 | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 1000 | 0,001 | 0,001 | Absolute |  |  | 50 |  | Low | Low |  | Removable
16 | Beam monitor 2 | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 1000 | 0,001 | 0,001 | Absolute |  |  | 50 |  | Low | Low |  | Removable
17 | Beam monitor 3 | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 1000 | 0,001 | 0,001 | Absolute |  |  | 50 |  | Low | Low |  | Removable
18 | Beam monitor 4 | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 1000 | 0,001 | 0,001 | Absolute |  |  | 50 |  | Low | Low |  | Removable
19 | Beam monitor 5 | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 1000 | 0,001 | 0,001 | Absolute |  |  | 50 |  | Low | Low |  | Removable
20 | Beam monitor 6 | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 1000 | 0,001 | 0,001 | Absolute |  |  | 50 |  | Low | Low |  | Removable
21 | Beam monitor 7 | 13.6.16.X.X.X- |  | Z1 | Linear | Electrical | 1000 | 0,001 | 0,001 | Absolute |  |  | 50 |  | Low | Low |  | Removable
Preliminary elaboration of Table of Motion.
21 axis in configuration 3; 16 axis in configuration 2; 12 axis in configuration 1.
Sent to Federico Rojas for quotation.
Added 1k€/axis for motor and encoder (not MC budget).
Modifications came after the scope setting report was uploaded.

## Slide 13
Motion Control
  Configuration 3 = 138.573 k€ according to Feredico Rojas. Currently 209.158 k€.
    17.4 k€ of Project Management and Integration.
    41.7 k€ of Design. Currently 59.952 k€.
    39.494 k€ of Procurement and Fabrication. Currently 90.409 k€.
    5.995 k€ of Installation. Currently 6.263 k€.
    33.984 k€ of Cold Commissioning. Currently 35.134 k€.
  Configuration 2 = 116.313 k€ according to Feredico Rojas. Currently 185.279 k€.
    17.4 k€ of Project Management and Integration.
    36.312 k€ of Design. Currently 53.95 k€.
    29.373 k€ of Procurement and Fabrication. Currently 80.046 k€.
    4.65 k€ of Installation. Currently 4.768 k€.
    28.578 k€ of Cold Commissioning. Currently 29.115 k€.
  Configuration 1 = 103.193 k€ according to Feredico Rojas. Currently 125.696 k€.
    17.4 k€ of Project Management and Integration.
    31.357 k€ of Design. Currently 37.057 k€.
    27.424 k€ of Procurement and Fabrication. Currently 44.244 k€.
    3.413 k€ of Installation. Currently 3.396 k€.
    23.599 k€ of Cold Commissioning.

## Slide 14
Infrastructure
  Already estimated items:
    Cabling: 5 k€. Source: Gábor László.
    Personal Safety System Design: 120 k€. From BIFROST.
    Crane outside cave: 30 k€. From BIFROST.
    Crane in cave: 10 k€. From BIFROST.
    Media distribution: 80 k€. From BIFROST.
    False floor: 20 k€. From BIFROST.
    Helium recovery. 20 k€. From BIFROST.
    Control hutch. 150 k€. From BIFROST.
    Infrastructure installation. 172 k€. From BIFROST.
    Instrument cabin. 30 k€. From TREX.
  Items not estimated yet:
    Cooling installation. ESS-0063538: is provided by ESS and paid by the instrument.
    Fiber connection from the instrument to the galleries: ESS-0063538: is provided by ESS and paid by the instrument.
    Domestic and waste water: ESS-0063538: is provided by ESS and paid by the instrument.