## Slide 1
ESS 2nd BI Forum Bilbao October 3-4, 2016
Laura Segui & Philippe Legou
for Saclay group
A. Delbart, D. Desforge, F. Gougnaud, C. Lahonde-Hamdoun,
P. Legou, J. Marroncle,T. Papaevangelou, L. Segui, G.Tsiledakis
DSM/Irfu, CEA Saclay
nBLM project at ESS
A New Beam Loss Monitor Concept Based on Fast Neutrons Detection using Micromegas detector

## Slide 2
 Why new BLM?
 Micromegas in few words
nBLM simulations
Examples of Micromegas detectors setup
 Summary
Overview

## Slide 3
Why new BLM?
Two objectives: monitoring and safety
Challenges
 RF cavities emit g's → background for ionization chambers used as BLMs
 Continuous monitoring of small losses is needed
  Measure the Beam Losses quite accurately and their locations
Requirements
1 - Blind for γ’s and X-rays contributions from cavity emissions
2 - directionality  good correlation beam loss location / detection (blind for thermal neutrons)
3 - good efficiency
4 - Sensitive enough to monitor small losses
5 - Appropriate for high rates, radiation hardness
6 - Sensitive and fast enough to react on “catastrophic event”
 Proposal:
Neutron BLM (nBLM), based on Micromegas detector equipped with neutron convertors and moderators
A NEW BEAM LOSS MONITOR CONCEPT BASED ON FAST NEUTRON DETECTION AND VERY LOW PHOTON SENSITIVITY, Jacques Marroncle et al, IBIC 2016

## Slide 4
Lot of improvements, evolutions can be done on Micromegas and by changing their parameters (gaps, gas, electric potential, read-out…). It can achieved:
  high fluxes greater than 108 counts/cm2/s (KABES, DEMIN)
  spatial resolutions down to 15 µm
  time resolution down to 30 ps
  cylindrical shape are now working routinely (ASACUSA, Clas12)
large surface area (>1 m2) can be covered (COMPASS, ATLAS)
Resistive bulk technologies, reduce spark effects, reducing dead time  BLM
Micromegas: Multi-Pattern Gaseous Detector, invented in 1995 at CEA Saclay1
Parallel plate detector with a strengthened thin mesh dividing the gas volume in 2 parts:
  drift region (1 to 10 mm)  E ≈ 100 V/mm
  amplification region (30 to 100 µm)  E ≈ 10000 V/mm
Micromegas Detector
1 Y. Giomataris, P. Rebourgeard, J.P. Robert and G. Charpak, “Micromegas: A high-granularity position sensitive gaseous detector for high particle-flux environments”, Nuc. Instrum. Meth. A 376 (1996) 29.

## Slide 5
Cadmium envelop
  to absorb the incident thermal neutrons
Polyethylene moderator
  to thermalize the incident fast-neutrons  varying thickness allows to adjust the energy threshold
  to absorb the remaining incident thermal neutrons
Double Micromegas
  to increase the neutron detection efficiency with B4C thin films (1.5 - 2 µm)
  gas: He (≈ 1.1 bar) or N2, Ne…
    - He is better for photon discrimination
This geometry was simulated using FLUKA1 and GEANT 42 codes to check the compliance with the requirements
nBLM geometry
1 G. Battistoni et al., The FLUKA code: Description and Benchmarking, in Proc. AIP Conf. Proc. 03, vol. 896, 2007, pp.31. http://dx.doi.org/10.1063/1.2720459.
2 GEANT Collaboration, S. Agostinelli et al., GEANT4-a sim-ulation toolkit, NIM A 509 (2003) 250.
(zoom)

## Slide 6
Thermal neutrons with respect to detection thresholds:
  10 keV  Eff. < 0.007 %
  30 keV  nBLM is blind to external thermal neutrons
  Background: γ contributions coming from 114Cd and 10B neutron are taken into account, but almost completely removed with low detection thresholds.
nBLM response to external “thermal” neutrons
0.01 < Eneutron (eV) < 1
113Cd (n, γ) 114Cd capture
1
0.1
0.01
0.001
Incident Eneutron (eV)
10
0.01
0.1
1
0.001
(eV)

## Slide 7
Photons with respect to detection thresholds:
  10 keV  Eff. < 0.0062 %
  20 keV  nBLM almost blind to photons
  Note: Micromegas is a thin gas chamber with small amount of material, explaining their transparency to photons (low RL)
nBLM response to X-rays and γ
0.01 < Ephoton (MeV) < 100
10-2
10-1
1
10-1
102
Incident Ephoton (MeV)
10
14
18
22
26
ΔEphoton (keV)

## Slide 8
Moderator thickness
  2, 4 and 6 cm
  Threshold energy = 10 keV
  Overall efficiency for 4 cm  3.8%
  Moderator thickness can be changed to tune slightly the neutron energy threshold.
Contribution of all neutrons under 0.2 eV is suppressed
  Thermal neutrons are almost removed
nBLM efficiency to fast neutrons

## Slide 9
nBLM time response:
Neutrons 0.1 eV < E < 100 MeV
Isotropically distributed in Cd surface, pointing towards the detector, ~15x104 n/cm2
Time since particle creation until first energy deposit (Geant4 simulation).
nBLM MonteCarlo simulations
 Only 10% of events detected in the first 10 μs
 95% after 150μs (time due to neutron moderation time)
Could be too slow for safety reasons → Add a fast stage nBLM
Deposited
Eth = 10 keV
Deposited
Eth = 10 keV

## Slide 10
Principle of operation: detection of recoil protons produced in polypropylene
Fast nBLM geometry
  1 mm Al + 2 mm polypropylene will be enough to be quite insensitive to thermal neutron
  thin Al (50 nm) coating on polypropylene to polarize the Micromegas and to insure a high transparent to recoil protons.
Fast nBLM
Double MM
neutrons
Polyethylene
moderator
fast MM
or γ

## Slide 11
Time response and efficiency
Fast response of the fast nBLM MonteCarlo studies
 Neutrons between 0.1 – 100 MeV / Isotropically distributed in external Al surface, pointing towards the detector
 ~95% of events interacts in the volume in < 4 ns
Time distribution
Initial n energy
vs. time
L. Segui ESS BI Forum, Oct. 2016 1 1
Fast neutron (Th=10 keV)  low efficiency
  for Eneut. = 1 MeV  Eff. = 3 10-4
  for Eneut. = 10 MeV  Eff. = 8 10-4
Thermal neutrons: E=0.025 eV
  Thres. = 1 keV  Eff. = 6.6 10-5
  Thres. = 5 keV  Eff. = 5.8 10-6
  Thres. = 10 keV  Eff. < 5 10-7
Note : Fast-BLM’s mission is dealing with critical events  huge neutrons emission

## Slide 12
Fast nBLM MonteCarlo studies
Blind to gammas
Very good photon/neutron discrimination
With an energy cut >20 keV → detector blind to gammas
Gammas and neutrons
Geant4, isotropic beam
L. Segui ESS BI Forum, Oct. 2016 12

## Slide 13
DEMIN
AND ANCILLARY RACK LOCATIONS
DEMIN
ANCILLARY RACK
NY- USA
Example of experimental conditions in neutrons detection
(Inertial Confinement Fusion)

## Slide 14
70cm=27.5’’
180cm=71’’
   Digitizers for pramp cards
   GPIB/Ethernet interface
   Power supplies for digitizers
   Gas flow regulator
   PC (drives the gas regulator)
   NIM crate :
     High volyage unit
    (cathodes & Meshes)
     Power supplies for preamps
14
Philippe Legou ESS Meeting 2016, July 25th
Ancillary rack

## Slide 15
Coaxial connectors for analog
signals
preamplifiers
Reading strips
The 4-layers PCB
Strips, spacers and the mesh are sticked
on the PCB by a single process.
Cables for :
 Signals
 Low voltage / High Voltage
 Temperature sensor
Strips, preamplifiers and conectors : good signals integrity

## Slide 16
Size: 22 x 20 x 5 mm
weight : 4g
Power supply : + 5V -5V
Consumption  50 mW
Noise: 600 µV rms
In few figures …
Input: positive or negative
Rise time : < 1ns
The new front-end : FAMMAS
( Fast Amplifier Module for Micromegas ApplicationS)
Example of Front end electronics used in neutron detection
16

## Slide 17
Test conditions :  visualisation at about 30 meters coaxial cable.
 Gaz : COMPASS.
 Beam.
Mesh Voltage : 350 V
Mesh Voltage : 370 V
Beam test
No crosstalk between neighbour channels
Good signal / noise ratio

## Slide 18
18
plegou@cea.fr
Size: 60 x 22 x 5 mm
weight : 8g
Power supply : + 5V -5.2V
Noise: 700 µV rms
In few figures …
Input: positive or negative
Output : ECL differential
Buffer area
Front-end module with differential output

## Slide 19
KABES on the SPS at CERN
NA 48 experiment in operation at CERN
from summer 2003 to end of 2005.
Flux 3x106/cm2/s –
Principle TPC + Micromegas
KABES & SEXTANT
Beam Spectrometers using Micromegas in Time Projection Chamber mode,
Ph.Legou et al , 2007. NSS '07. IEEE (Volume:1 )
KABES
TPC – CERN
2003 / 2005
ASACUSA
Curved tracker
2014 / today
The ASACUSA Micromegas Tracker: A cylindrical, bulk Micromegas
detector for antimatter research
REVIEW OF SCIENTIFIC INSTRUMENTS 86, 083304 (2015)
Examples of Micromegas applications

## Slide 20
 Dec. 2016 : PRD1, first design of the prototype
  On going design and simulations to optimize:
  Drift, polyethylene thickness, convertor, …
 June 2017: PDR2 and prototype
Test the detector at:
     LICORNE (fast neutrons)
     COCASE (high activity 60Co irradiator)
     ORPHEE (high thermal neutron flux / aging)
     SEDI laboratory (long term stability)
 Nov. 2017: CDR1
  Initial prototype tests report
 April 2018: CDR2
  Final prototype tests
  Documentation of different systems:
  Slow control, electronics, etc...
 Dec. 2018 : System delivered to ESS
 ESS hand over: begin of 2019
ESS operation 2019
Dec.2016
PDR1
April 2018
Final design
Dec. 2018
System
delivered
nBLM project for ESS
CEA/Saclay in-kind contract with ESS
Expected time for optimization, design, construction, characterization, installation & commissioning of 42 modules: 2.5 year
L. Segui ESS BI Forum, Oct. 2016 15
June.2017
PDR2 prototype

## Slide 21
Simulation of a new nBLM based on Micromegas detector has been performed
    good sensitivity to fast neutrons  overall efficiency ≈ 4%
    “blind” to thermal neutrons  directionality
    “blind” to X-rays and γ’s  to be safe from cavity photon emissions
    fast BLM component  t < 5 ns
    big neutron signal deposit (due to moderation) allows to count neutrons individually
    Reliable Micromegas solution
Future: design prototypes
  robustness, reliability, radiation hardness…
  gas choice
  fast and low noise FEE
  sealed mode… already tested but need to be checked
… and test with real neutrons
Conclusion / Summary

## Slide 22
Gracias …

## Slide 24
Fast nBLM MonteCarlo studies
Gammas and neutrons: drift and gas studies
As an example of studies on-going
For HE gammas, 0 events detected below 8 mm drift →
<2.0 x 10-7
Equilibrium between efficiency and gamma rejection
  Below 8 mm drift, “blind” for HE gammas with Eth at 20 keV
  Efficiency still of 3% (0.3%) for 10 (1) MeV neutrons at 5 mm
Repeat study with Ne → higher efficiency but also for gammas → smaller rejection

## Slide 25
 | Efficiency (x10-4) |  | 
 | 1 keV, min Edep | 10 keV cut | 20 keV cut
n 1 MeV | 3.56 | 3.41 | 3.27
n 10 MeV | 34.60 | 13.5 | 5.18
g 25 keV | 20.0 | 1.78 | 0.008
g 100 keV | 7.01 | 0.13 | 0.014
g 2 MeV | 129.00 | 0.07 | 0.040
Fast nBLM MonteCarlo studies
Blind to gammas
With an energy cut >20 keV → detector blind to gammas
Gammas and neutrons

## Slide 26
nBLM angular response
  quite low effect due to neutron slowing down inside moderator
  this behavior let us expect a nBLM efficiency greater than the active surface of Micromegas
angular and time responses of nBLM
Neutron energy (MeV)
Efficiency (%)
nBLM time response
  only 17% of events are detected during the 10 first µs, while they are all after 300 µs!
  due to neutron moderation time
  might be too slow for safety purposes
   Proposition to add a fast stage of BLM
Ratio =

## Slide 27
Hypothesis for FLUKA & GEANT 4 codes
  Neutrons: double exponential distribution for energy ranging from 0.1 eV to 100 MeV
  Photons: double exponential distribution from 10 keV to 100 MeV
  Withdrawing is done upstream and transversely to the nBLM entrance window in a volume filled with air
  angular divergence of 10 mrad for incident neutrons
Codes: calculate the energy deposition in the gas
Checked: results obtained with both codes are similar!
Event simulation