## Slide 1: LEBT EMU and DOPPLER STATUS
BI FORUM
Franck Senée presented by Jacques Marroncle
Bilbao 2016, Oct. 3rd 2016

## Slide 2: ESS LEBT
ESS ion source parameters:
Proton beam
75 keV
Current max = 100 mA
Pulse length = 6 ms
Repetition rate = 14 Hz
Duty cycle 8%
BEAM
EMU V
EMU H
CEA: 2 emu (H - V) & 1 doppler unit
PERMANENT TANK

## Slide 3: Beam profil and Commissioning phases at catania
EMU design
(Ø min = 6 mm)
Doppler design
(Ø max = 80 mm)
C TANK
P TANK
PERMANENT TANK
P TANK
PRELIMINARY TANK
C TANK
COMMISSIONING TANK
PRELIMINARY TANK
installed at Catania
LEBT Beam profilE and Commissioning phases at catania
1
1
2
2
3
3
C TANK

## Slide 4: EMU: Allison scanners type EMU @ IRFU: philippe abbon - philippe daniel-thomas - jean-françois denis -françoise gougnaud - tom joannem - nicolas misiara - FRANCK SENÉE - VICTOR SILVA - Olivier tuske

## Slide 5: Working Principle
y3
y1
y2
y’3
y’1
y’2
z
y
x
y1
y2
y3
Beam Axis
y’1
y’3
intensity
Angle y’
y’1
y’2
y’3
intensity
Angle y’
intensity
Angle y’
y’2

## Slide 6: Allison scanner
thermal
screen
U: kinetic energy (eV) of the beam particles (H+)
I (µA)
0.39 π.mm.mrad
with =

## Slide 7: Design overview
7
Beam-Stopper
FC
Deviation plates
100 µm Slit for beamlet selection
Actuator
repeller

## Slide 8: Thermo-mechanical calculations
Beam inputs:
Description | Value | Unit
Beam energy | 75 | kV
Beam intensity at peak | 100 | mA
Repetition rate | 14 | Hz
Pulse length | 6 | ms
  Gaussian power deposition for a 3 mm RMS beam: 165 W/mm²
Duty cycle | 8 | %
MINIMAL radius beam size (rms) | 3 | mm
MAXIMAL beam size diameter | 80 | mm
Average power absorbed | 630 | W
Peak power | 7.5 | kW
.357 Magnum (9×33mm) = 646 J
7,63 Mauser = 510 J
Beam energy deposition during 1 second = 630 J

## Slide 9: Tmc: TZM entrance slit
1
2
3
TZM
selected
Cu
1
2
3
Elasticity limit [MPa] for Copper (left) and for TZM (right) versus temperature [k]
TZM alloy: Titanium – Molybdenum – Zirconia
ANSYS simulation: Von Mises (MPa) stress
BEAM IMPACT:
6 mm diameter

## Slide 10: TZM entrance slit : BRAZING TESTS
Result of several brazing tests : TZM cracking (July 2016)
Brazing tests on :
damaged ESS slits
and other TZM samples
OK with Au-Ni as brazing filler metal: Leakage rate = 2,7.10-10 mbar.l.s-1
NOT OK with Au-Cu as brazing filler metal
After this failure, the new brazing solution (Au-Ni) works successfully, with very low leakage
ESS
ESS
OTHER
OTHER

## Slide 11: Tm calculations: BEAM STOPPER
4 bars
Thickness
W= 5 mm
Cu= 20 mm

## Slide 12: Tm calculations: BEAM STOPPER
ANSYS simulation: Von Mises [Mpa] stress
BEAM IMPACT:
6 mm diameter

## Slide 13: Tm calculations: BEAM STOPPER
1
2
4
3
5
4
W+Cu
selected
W+Cu
Cu

## Slide 14: PARTS Assembly before the HIP cycle (950°C/1400bars: 2H)
Goal improving the mechanic contact
Works before assembly:
- New design for assembly whole parts
- Machining W parts and Copper
- Nickel-W face being in contact with the Copper
- Manufacturing of the water box
- Machining of the Carbon inserts and stainless steel container
Water box
Copper parts
W parts
W parts
Copper parts
Water box
HIP: Hot Isostatic Pressure

## Slide 15: PArts Assembly before the HIP cycle (950°C/1400bars: 2H)
Setting up of the parts in the container
Setting up completed
Container ready for the HIP cycle

## Slide 16: PArts Assembly before the HIP cycle (950°C/1400bars: 2H)
Machining after HIP (Hot Isostatic pressing) :
Remove the container (expected on September 30th, 2016)
Final machining and water-cooling stainless pipes brazing on the BS  expected at Saclay by mid October 2016.
HIP cycle performed on September 23rd 2016

## Slide 17: BACKUP SOLUTION
In case of the HIP process will not be ready on time, a whole Beam-Stopper made of Copper has been performed
Delivery foreseen on October 5th 2016
It will allow to get at least one complete EMU for the acceptance test date but at lower power density than the nominal one.

## Slide 18: Status and schedule
all mechanical EMU parts are already at Saclay, but the slits and the BS scheduled by mid-October 2016.
we will do our very best to install that in Catania during the week 46 (Nov. 14th-18th)
acceptance test at Saclay with a single EMU and with beam expected the week before.
cabinet
patch box
actuators
mast
deviation plate mechanics

## Slide 19: THE DOPPLER UNIT DOppler @ IRFU: françoise gougnaud - pierre mattei - yves lussignol - FRANCK SENÉE- Olivier tuske

## Slide 20: THE DOPPLER UNIT: principLe
DlH3+
DlH2+
DlH+
Hydrogen beam
H0(H+) H0(H2+) H0(H3+)
Plasma H0 (light emission excited residual gas)
656.3 nm
θ
θ
Doppler shift of the Hα ray
expected  versus angle, for 75 keV beam
measured by a spectrometer which is able to separate the species
The intensity of the spectroscopic rays is proportional to the quantity of the produced species, but the production rate of the Hα ray by the residuel gas interaction is different for each species

implying that corrections using recombined cross sections must be applied
Spectrometer

## Slide 21: ACCEptance test approved on the silhi source on end May 2016
A mirror located inside the VC at
179mm from the viewport and tilted by 35° wrt the
perpendicular allow to
observe the beam under 20° incident angle.
beam direction
DN100CF viewport
20°
beam axis

## Slide 22: ACCEptance test approved on the silhi source
30% more light is expected with the ESS optical fiber coupling and the whole beam will be observed (100 mm diameter).
86
61
ESS environment
SILHI environment
The measurement of the specie fractions works in real time mode and for a single pulse of 6 ms length (world first!)
Apilux simulation obtained with both environments:

## Slide 23: ACCEptance test approved on the silhi source
Operator Control sytem interface made from a single GUI screen
A special thanks to Hinko and Cyrille for their help
The essential
Parameters can be control:
Exposure time
Binning Y
Slit width
Grating….

## Slide 24: Status and schedule
Interface Viewport / Optical system
Doppler unit delivered at Catania on July 7th 2016.
Installation and test with beam of the Doppler unit is scheduled on week 42 (Oct. 17-21)

## Slide 25: SUmmary
Interface Viewport / Optical system
EMU
  brazing and HIP:
      done on Sept. 23rd 2016
      waiting for final conclusion, but optimistic at first glance
  backup BD in cooper should be ready by Oct. 5th
  all components foreseen to be at Saclay by mid-Oct.
  Acceptance test at Saclay: week 45 (Nov. 7th-11th)
  installation at Catania foreseen for week 46
Doppler Unit
  acceptance test successfully done at Saclay on May 2016
  delivered at Catania on July 7th 2016