## Slide 1: Instrument maintenance with in the Common Shielding bunker
28-9-2016
Iain Sutton
On behalf of the Neutron Technologies Division

## Slide 2: Intro
Why is having a coherent maintenance strategy important ?
Because it contributes to;
  Instrument availability for science
  Operating costs
  Minimizing exposure of personnel
  The facilities overall reputation

## Slide 3: Context
ESS specific context
  Long source instruments
    Components@ R6
  Shielding common volume
    Vertical access
HPower source issues
  Neutron activation
  Neutron damage
Project constraints
  Minimise initial investment

## Slide 4: Equipment within Common shielding bunker
Component type Numbers
Choppers 98 axis
Optics 400 m (200 sections)
Shutters 30
Collimators 10
Monitors 30
Supports 200
Service supply
22 instruments assumed (13L-9S)

## Slide 5: Operational context
Schedule
2 long shut downs >20days
2 short shut downs ~5days
18 very short shut downs <3days
Notes: Early years
  instrument availability
  rapid ramp up
  ongoing build

## Slide 6: Workload
Type 1: Scheduled activities
Type 1A: Long periodicity
  Installation / Decommissioning
Type 1B: Short periodicity
  Inspection/re-alignment/overhaul
Type 2: Unscheduled activities
Breakdown
  Repair / Replacement
INTERVENTIONS
in CSB pa
SCHEDULED
1(A)
PERIODIC
1(B)
REGULAR
UNSCHEDULED
2
BREAKDOWN
10-20
6
60
(All estimates are steady state – post burn-in)
Notes: Interventions may be considered as two types;
Type 1: Scheduled activities
  Type 1A: Long periodicity events
    Installation / Decommissioning
  Type 1B: Short periodicity events
    Inspection/ Adjustment /Replacement
Type 2: Unscheduled activities
(‘random’ events with 5-20 incidents p.a.)
  Breakdown repair
    Inspection / Adjustment / Replacement
  Accident scenario

## Slide 7: Activation
Activation
Integrated beam
Beam power
Beam days
Full remote handling ?
Long tools & local shielding / hot spot removal
Restricted floor access with local shielding
Floor level access prohibited
Vertical access only
MIXED REGIME

## Slide 8: Strategy
Is to have a strategy !
Be prepared ….with a minimal upfront investment
Classification of access zones
  Graded requirements
  Equipment validation
Established Interfaces
  Fixation
  Support & alignment
  Service connections
Standardised methods & tooling
Pre-equip units for RH strategically
Notes: don't reinvent the wheel SNS
ISIS
Ensure future compatibility
upgrade path

## Slide 9: Operational constraints
Hotspots
Shutters
Collimators
PPS choppers (Tzero)
Beam stop
Monolith face (?)

## Slide 10
Logistics
Hall Crane
Bunker Crane
Hand over zone
Bunker crane
20t winch
2t winch (requesting)
Coverage Acceptable
Hall Crane
30t winch
5t winch
Coverage Acceptable
Issues
Bunker area limited lift height ( modifications under study)
Handover zone is limited
System will be slow !

## Slide 11: Maintenance
Access on temporary platform on top of bunker pillars.
Main strategy is long tools and other “simple” solutions.
Highly activated components will require hot-swap.

## Slide 12: Extraction at monolith face
Straight lift of optics component section
Straight vertical extraction
using crane
Horizontal translation on rails
1.
1.
2.
2.

## Slide 13: Risks to operational performance
Bottle necks
  Number of interventions
  Saturation of crane capacity or tooling
  Cranes are single point failures
Ensuring implementation of standards
  Poor reliability
  Incompatible equipment
  Material activation
Aggressive build schedule
  Getting out of burn in before activity build up

## Slide 14: Vibration mitigation
Reduce source terms
  Balance
  Magnetic bearings
Avoid coupling
  Rigid structures
  Rigid fixation
Isolate sources
  Support – Base
  Separate base plates (PPSc)
Attenuate
  Bolted & grouted base plate
Bolted & grouted baseplate
Spring loaded fixation
Reservation for Thin isolator
Rigid support
Notes: don't reinvent the wheel SNS
ISIS
Ensure future compatibility
upgrade path

## Slide 15: Target Safety related instrument failure
Current strategy “a layered and graded approach.”
Beam port a ‘thick’ window (Rated for normal safety case)
BP shutter a ’thick-ish’ window (v50)
Upstream side of high speed choppers a ’complimentary’ window
@ 11.5m a ’thick-ish’ window ‘fire break’
Metallic guides only upstream of ‘fire break’
BEAM PORT
GUIDE
GUIDE
Common shielding
SHUTTER
2
1
3
4
Safety related equipment
Normal equipment
Fire break

## Slide 16: Remote handling
Current thinking
On which equipment this should apply to and why

## Slide 17: Which equipment should be equipped to be remote handled ?
Equipment which represents a significant hazard to personnel during interventions (due to activation).
PPS choppers
Shutters & collimators
Some supports (?)
Equipment which requires regular intervention.
Overhaul or inspect period < 10 years.
Removed frequently i.e. is installed between R6 – R11.5
Equipment presenting a significant risk of breakdown. Reliability < 98%

## Slide 18: Which operation should be carried out using remote handling ?
Operations
Disconnection
  Beamline components, hold down and Services
Removal
  Fitment of cask if required
(Re)insertion
Repositioning or alignment (if required)
Reconnection

## Slide 19: Status
Where are we now
A plan for the way forward
outstanding issues

## Slide 20: Thanks you for your attention !
Many thanks for contribution from the Bunker group and the Neutron Chopper Group

## Slide 21: Interventions - type 1A - Scheduled, PERIODIC (?)

## Slide 22: Type 1A Once in an (instrument) life time (10yrs) aka bunker annual events !
These activities are expected to be infrequent (periodicity >10yrs) when considered on a beam line by beam line basis.
Note when each bunker is considered as a system of 8-12 beam lines interventions become annual.
Interventions
Beam port activation
Instrument installation
Beam port deactivation
Instrument removal
Installation of neighboring instrument

## Slide 23: Current situation ‘Vertical shutter removal’
N
Monolith
Key points
Components situated in bunker and in basement
Access to mobile components via basement
Access to some components requires access via bunker
Mobile assembly accessed via basement
Static assembly access via bunker

## Slide 24: Interventions -type 1(b) schedule Regular (every shut-down

## Slide 25: Type 1B ‘the regulars’
These activities will be regular occurrences
These events will predominate work load of tech support during operation
Systems must be in place to enable these activities to be carried out minimising exposure to personnel
– ie minimising hands on.

## Slide 26: Scheduled & Regular
Activity Incidence over 10yr period
Beam port systems inspection 40
Neutron chopper maintenance 100
Optics realignment 20
Instrument shutter maintenance 20
Other component maintenance 60
Removal neighboring & Hot 80
Total 320

## Slide 27: Interventions - Type II unscheduled (BREAKDOWNS)
Breakdown, Screw ups, and other stuff that just happens

## Slide 28: Sub system level breakdown data
Source: ISIS
Note: excluding - Sample environments, Computing & Motion control
Causes often requiring access within heavy shielding for repair
Causes not requiring access within heavy shielding for repair
Chart: 
 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 1 | 1 | 1
 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 1
 | 3 | 1 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 2 | 3 | 1 | 0 | 1 | 1 | 1 | 2 | 3
Beam interruptions
Cycles (Normalized to 35 day)
CAUSE OF STOPPAGE
BLUE Chopper system
RED Other component
GREEN At request of I.S.
~ 3 ‘significant’ interruption per 100 beam days
+3.5 uses of shutters to effect ‘running repairs’
Notes:   Calculated failure rates for existing instruments and equipment
  And from that
  Factoring in the increase comlexity of ESS instruments tried to come up with some estimates for our own instruments
  Failure of a specific instruments every 3 or 4 years
  Failure of ‘An’ instrument each 35day cycle

## Slide 29: Unscheduled (breakdowns & unforseen)
Activity Incidence over 10yr period
Beam-port systems breakdown 4
Neutron chopper failure 40
Optics breakdown 20
Instrument shutter breakdown 20
Other component breakdown 60
Removal of hot components (neighboring work) 30
Accident scenarios (H2) 2
NOTE : Indicative figures – work in progress

## Slide 30: Availability & Source Schedule
Lengthening short breaks 3 > 5 days enables them to be used for repair.
Results in increased overall availability
Chart: 
| option A /B | Option C/D | Option E
225 BD | 66.4 | 93.4 | 81
215 BD + 4 Short breaks | 79.8 | 93.4 | 89
Notes: As a general rule unscheduled shut downs for repair are restricted to GREEN zones
Access to components in other zones must await scheduled ‘short shutdown’ – to cope with more complex interventions.
Longer interventions will be dealt with in two steps, with removal of equipment during one S-SD, repair off line and replacement at the next S-SD.
In the case of interventions requiring extended SD then shutters can be held close after start up of the facility.

## Slide 31: DC-LR installation
1.4m
.2.1m

## Slide 32
Worst case
Disc chopper Large rotor
Odin, Freia, (4-6 units)
N
2100
300
Assuming a rotor of 1800mm
The height of the unit is 2100mm
Assuming a lifting tool of 300mm
Gap to floor of 200mm
Free height require > 2600mm
200 gap

## Slide 33: Hook height
~1800 (700mm disc)
300mm (20t hook)
~2300
N BEAM
Gap 200mm

## Slide 34
Types of equipment in the bunker
Over/Under Choppers
Guide modules (2m each)
Face-to-Face Choppers
and Single Choppers
Other components are not considered especially collimators
Large Rotor choppers
(1m and above)
PPS Choppers

## Slide 35
Components and their placement
(with respect to the instrument in the beam line allocation document)
Components | Total number | In bunker | In bunker and serviced by monolith crane | Serviced by monolith crane but not in bunker
Over/Under Chopper | 23 | 16 | 16 | 0
Single Chopper | 33 | 19 | 16 | 2
Face-to-Face chopper | 26 | 20 | 19 | 0
Large rotor choppers (1-1.4m) | 7 | 7 | 5 | 0
Large rotor choppers (>1.4m) | 3 | 2 | 1 | 0
PPS | 6 | 4 | 4 | 0
Guide module | 218* | 201 | 165 | 10
Other | 12 | 3 | 2 | 2
Total | 391 | 272 | 228 | 14
*Every instrument is considered to have a single guide module from m30 onwards

## Slide 36: Installed systems What
Insert schematic of bunker components here

## Slide 37: Installed systems Neutron choppers
High speed double choppers 8-10 assemblies
PSS = 5 assemblies
High or Hot equipment

## Slide 38: Component map

## Slide 39: DCSR (2PA-1-CRD) Conclusions
A hook height of 2.5m is a minimum for current equipment
2.7m prefered
2t – 5t lifting equipment would be strongly preferred.
Special mountings will be required to ensure installation compatibility

## Slide 40: 20t lifting kit
20t equipment is big !
Hook assembly can be 1.3m tall
Spreader bar is 500mm deep
Hook eye is 500mm high

## Slide 41: CHIM – Umbilical and routing
Standard ESS package for:
  utilities
  cables
Supplier specific cables
Connection plate is now simplified
Routing is work in progress
WIP

## Slide 42: Bunker remote handling plate (Horizontal) solution
Rack
4 connectors per cable

## Slide 43: ChIMF (Chopper Integration and Maintenance Facility)
ChIMF-1 in hall 2 (Partially moved to Pre-OPS)
ChIMF-2 in hall 1 (Pre-OPS)
Principal activities,
  Systems Integration
    Assembly of new systems
    Minor modification of components
  Commissioning
    Setup
    Electrical tests
    Mechanical tests (low energy)
    Pressure tests
    Cooling tests
  Maintenance
  Repair
Area defined in ESS-0047351
Area flexible to allow for site changes
  Example, one area more focused on T0
150sqm broken down into
Basic machine shop (15sqm)
Mechanical assembly (25sqm)
Metrology (10sqm)
Electrical assembly (15sqm)
General assembly (25sqm)
System test area (40sqm)
Parts storage (15sqm)
QA (10sqm)
Not including active storage area (20+25sqm)